## Sheet1
Method name | N (no. labs) | Mean | Bias | CV | Sample | Reference Value | Year | Source (CAP/EurA1c) | Type (Fresh/frozen) | Manufacturer | Total Error
(Ortho Clin Diag) Vitros 5.1 FS Chem System | 183 | 6.14 | -0.16 | 2.3 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Vitros | #VALUE!
(Ortho Clin Diag) Vitros 5.1 FS Chem System | 183 | 7.59 | -0.01 | 2.5 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Vitros | 2.49
(Ortho Clin Diag) Vitros 5.1 FS Chem System | 183 | 9.42 | 0.22 | 3 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Vitros | 3.22
(Ortho Clin Diag) Vitros 5.1 FS Chem System | 210 | 5.42 | -0.18 | 2.3 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Vitros | 2.12
(Ortho Clin Diag) Vitros 5.1 FS Chem System | 210 | 9.34 | -0.06 | 2.6 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Vitros | 2.54
(Ortho Clin Diag) Vitros 5.1 FS Chem System | 210 | 7.01 | -0.19 | 2.4 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Vitros | 2.21
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 151 | 8.33 | -0.3 | 2.5 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Vitros | 2.2
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 151 | 5.17 | -0.15 | 2.8 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Vitros | 2.65
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 209 | 6.41 | -0.17 | 2.2 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Vitros | 2.03
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 209 | 8.32 | -0.07 | 2.3 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Vitros | 2.23
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 209 | 5.51 | -0.14 | 2.4 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Vitros | 2.26
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 208 | 6.37 | -0.12 | 2.7 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Vitros | 2.58
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 208 | 6.85 | -0.12 | 2.6 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Vitros | 2.48
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 208 | 9.54 | -0.11 | 2.8 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Vitros | 2.69
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 191 | 6.46 | -0.33 | 2.5 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Vitros | 2.17
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 191 | 9.9 | -0.38 | 2.7 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Vitros | 2.32
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 191 | 6.47 | -0.35 | 2.5 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Vitros | 2.15
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 154 | 5.57 | -0.27 | 3.5 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Vitros | 3.23
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 154 | 11.31 | -0.4 | 2.8 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Vitros | 2.4
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 154 | 9.23 | -0.3 | 2.7 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Vitros | 2.4
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 154 | 4.95 | -0.09 | 3.2 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Vitros | 3.11
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 154 | 7.01 | -0.37 | 3.1 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Vitros | 2.73
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 156-192 | 11.64 | -0.05 | 3.7 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Vitros | 3.65
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 156-192 | 5.82 | -0.11 | 4 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Vitros | 3.89
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 156-192 | 5.14 | -0.03 | 3 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Vitros | 2.97
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 156-192 | 8.02 | -0.12 | 3.4 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Vitros | 3.28
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 156-192 | 9.27 | -0.03 | 3.2 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Vitros | 3.17
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 157-187 | 5.35 | 0.03 | 2.1 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Vitros | 2.13
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 157-187 | 9.37 | 0.21 | 3.3 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Vitros | 3.51
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 157-187 | 5.34 | 0.03 | 2.2 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Vitros | 2.23
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 157-187 | 12.35 | 0.32 | 4.1 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Vitros | 4.42
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 157-187 | 5.89 | -0.05 | 2.6 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Vitros | 2.55
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 161 | 5.28 | 0.01 | 2.1 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Vitros | 2.11
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 161 | 10.87 | 0.28 | 3.3 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Vitros | 3.58
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 161 | 6.06 | -0.14 | 2.2 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Vitros | 2.06
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 161 | 12.58 | 0.35 | 3.2 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Vitros | 3.55
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 161 | 7.43 | -0.08 | 2.3 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Vitros | 2.22
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 155-190 | 9.25 | 0.14 | 2.7 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Vitros | 2.84
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 155-190 | 5.94 | -0.07 | 2.3 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Vitros | 2.23
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 155-190 | 12.11 | 0.4 | 2.9 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Vitros | 3.3
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 155-190 | 5.04 | 0.02 | 2.4 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Vitros | 2.42
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 155-190 | 7.56 | -0.02 | 2.8 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Vitros | 2.78
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 170-191 | 6.5 | 0.09 | 2.7 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Vitros | 2.79
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 170-191 | 9.71 | 0.18 | 3 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Vitros | 0.18
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 170-191 | 5.4 | 0.06 | 2.1 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Vitros | 2.16
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 170-191 | 8.67 | 0.16 | 3 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Vitros | 3.16
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 170-191 | 7.42 | 0.17 | 3.1 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Vitros | 3.27
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 167 | 7.47 | 0.05 | 2.6 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Vitros | 2.65
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 167 | 5.21 | 0.01 | 2 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Vitros | 2.01
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 167 | 8.34 | 0.03 | 2.8 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Vitros | 2.83
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 167 | 10.51 | 0.18 | 2.9 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Vitros | 3.08
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 167 | 5.92 | 0.05 | 2.4 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Vitros | 2.45
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 194 | 7.28 | 0.03 | 2.6 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Vitros | 2.63
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 194 | 6.83 | 0.02 | 2.4 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Vitros | 2.42
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 194 | 8.66 | 0.01 | 2.7 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Vitros | 2.71
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 194 | 9.60 | 0.1 | 3 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Vitros | 3.1
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 194 | 5.42 | -0.03 | 2.1 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Vitros | 2.07
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 190 | 7.16 | 0.01 | 3 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Vitros | 3.01
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 190 | 5.25 | 0.06 | 2.4 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Vitros | 2.46
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 190 | 8.40 | -0.02 | 2.8 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Vitros | 2.78
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 190 | 9.71 | -0.08 | 3.1 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Vitros | 3.02
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 190 | 6.10 | -0.02 | 2.7 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Vitros | 2.68
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 171 | 6.28 | -0.03 | 2.5 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Vitros | 2.47
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 171 | 9.06 | -0.05 | 3 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Vitros | 2.95
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 171 | 5.27 | -0.04 | 2.2 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Vitros | 2.16
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 171 | 8.40 | 0.05 | 2.7 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Vitros | 2.75
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 | 171 | 7.60 | -0.02 | 2.3 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Vitros | 2.28
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 Chem System | 176 | 6.92 | -0.19 | 2.4 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Vitros | 2.21
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 Chem System | 176 | 9.22 | -0.1 | 2.3 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Vitros | 2.2
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 Chem System | 176 | 5.83 | -0.24 | 2.6 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Vitros | 2.36
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 Chem System | 193 | 7.69 | -0.36 | 2.4 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Vitros | 2.04
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 Chem System | 193 | 5.08 | -0.22 | 3.3 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Vitros | 3.08
(Ortho Clin Diag) Vitros 5.1 FS. 4600. 5600 Chem System | 193 | 5.81 | -0.33 | 3.1 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Vitros | 2.77
Abbott Alinity ci series | 18 | 5.13 | -0.12 | 1.7 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Abbott | 1.58
Abbott Alinity ci series | 18 | 6.34 | -0.07 | 1.5 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Abbott | 1.43
Abbott Alinity ci series | 18 | 8.32 | 0.11 | 0.8 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Abbott | 0.91
Abbott Alinity ci series | 18 | 5.41 | -0.06 | 1.4 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Abbott | 1.34
Abbott Alinity ci series | 18 | 9.62 | 0.14 | 0.9 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Abbott | 1.04
Abbott Alinity ci series | 19 | 8.03 | 0.06 | 1.2 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Abbott | 1.26
Abbott Alinity ci series | 19 | 5.84 | -0.05 | 1.2 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Abbott | 1.15
Abbott Alinity ci series | 19 | 5.05 | -0.08 | 1.4 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Abbott | 1.32
Abbott Alinity ci series | 19 | 7.39 | -0.01 | 1 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Abbott | -0.01
Abbott Alinity ci series | 19 | 9.28 | 0.11 | 1 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Abbott | 0.11
Abbott Alinity ci series | 36 | 7.79 | -0.02 | 1.4 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Abbott | 1.38
Abbott Alinity ci series | 36 | 5.78 | -0.13 | 1.6 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Abbott | 1.47
Abbott Alinity ci series | 36 | 8.63 | -0.06 | 1.7 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Abbott | 1.64
Abbott Alinity ci series | 36 | 5.76 | -0.1 | 1.4 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Abbott | 1.3
Abbott Alinity ci series | 36 | 9.53 | -0.03 | 1.6 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Abbott | 1.57
Abbott Alinity ci series | 53 | 5.34 | -0.14 | 3.1 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Abbott | 2.96
Abbott Alinity ci series | 53 | 8.28 | 0.02 | 2.2 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Abbott | 2.22
Abbott Alinity ci series | 53 | 5.01 | -0.14 | 3.3 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Abbott | 3.16
Abbott Alinity ci series | 53 | 10.17 | 0.12 | 1 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Abbott | 0.12
Abbott Alinity ci series | 53 | 5.87 | -0.12 | 2.2 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Abbott | 2.08
Abbott Alinity ci series | 59 | 8.93 | 0.07 | 1.2 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Abbott | 1.27
Abbott Alinity ci series | 59 | 6.77 | -0.07 | 1.9 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Abbott | 1.83
Abbott Alinity ci series | 59 | 9.49 | 0.1 | 1.3 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Abbott | 1.4
Abbott Alinity ci series | 59 | 6.02 | -0.11 | 1.8 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Abbott | 1.69
Abbott Alinity ci series | 59 | 7.51 | -0.01 | 1.6 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Abbott | 1.59
Abbott Alinity ci series | 67 | 5.62 | -0.18 | 1.7 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Abbott | 1.52
Abbott Alinity ci series | 67 | 9.21 | 0.09 | 1.6 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Abbott | 1.69
Abbott Alinity ci series | 67 | 5.35 | -0.12 | 1.4 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Abbott | 1.28
Abbott Alinity ci series | 67 | 7.29 | 0.04 | 1.4 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Abbott | 1.44
Abbott Alinity ci series | 67 | 9.32 | 0.06 | 1.2 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Abbott | 1.26
Abbott Alinity ci series | 87 | 5.1 | -0.11 | 2 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Abbott | -0.11
Abbott Alinity ci series | 87 | 6.59 | -0.09 | 1.6 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Abbott | 1.51
Abbott Alinity ci series | 87 | 8.27 |  | 1.8 | GH5-13 |  | 2021 | CAP | Fresh | Abbott | 1.8
Abbott Alinity ci series | 87 | 5.92 | -0.06 | 1.9 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Abbott | 1.84
Abbott Alinity ci series | 87 | 10.4 | 0.28 | 1.7 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Abbott | 1.98
Abbott Architect | 24 | 10.9 |  | 6.4 | GH2-01 | 10.7 | 2006 | CAP | Fresh | Abbott | 6.4
Abbott Architect | 24 | 4.8 |  | 8.8 | GH2-02 | 5.30 | 2006 | CAP | Fresh | Abbott | 8.8
Abbott Architect | 24 | 8.3 |  | 5.3 | GH2-03 | 8.40 | 2006 | CAP | Fresh | Abbott | 5.3
Abbott Architect | 28 | 6.5 |  | 7.4 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Abbott | 7.4
Abbott Architect | 28 | 6.6 |  | 7.7 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Abbott | 7.7
Abbott Architect | 28 | 11.6 |  | 5 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Abbott | 5
Abbott Architect | 28 | 6.5 |  | 7.4 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Abbott | 7.4
Abbott Architect | 28 | 6.6 |  | 7.7 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Abbott | 7.7
Abbott Architect | 28 | 11.6 |  | 5 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Abbott | 5
Abbott Architect | 47 | 5 |  | 5 | GH2-01 | 5.4 | 2007 | CAP | Fresh | Abbott | 5
Abbott Architect | 47 | 11.2 |  | 3.9 | GH2-02 | 11.6 | 2007 | CAP | Fresh | Abbott | 3.9
Abbott Architect | 47 | 7.4 |  | 4.1 | GH2-03 | 7.6 | 2007 | CAP | Fresh | Abbott | 4.1
Abbott Architect | 46 | 8.9 |  | 5.7 | GH2-04 | 9.2 | 2007 | CAP | Fresh | Abbott | 5.7
Abbott Architect | 46 | 10.9 |  | 4.1 | GH2-05 | 11.1 | 2007 | CAP | Fresh | Abbott | 4.1
Abbott Architect | 46 | 6 |  | 4.2 | GH2-06 | 6.2 | 2007 | CAP | Fresh | Abbott | 4.2
Abbott Architect | 50 | 6.73 | 0.13 | 4.5 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Abbott | 4.63
Abbott Architect | 50 | 9.62 | 0.12 | 4.2 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Abbott | 4.32
Abbott Architect | 50 | 7.7 | 0.3 | 5.4 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Abbott | 5.7
Abbott Architect | 52 | 5.94 | 0.04 | 5 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Abbott | 5.04
Abbott Architect | 52 | 9.98 | 0.18 | 3.9 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Abbott | 4.08
Abbott Architect | 52 | 7.56 | 0.16 | 5.2 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Abbott | 5.36
Abbott Architect | 53 | 5.02 | -0.18 | 5.9 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Abbott | 5.72
Abbott Architect | 53 | 8.74 | 0.04 | 3.5 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Abbott | 3.54
Abbott Architect | 53 | 6.22 | -0.08 | 4.2 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Abbott | 4.12
Abbott Architect | 59 | 8.73 | 0.23 | 3.5 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Abbott | 3.73
Abbott Architect | 59 | 5.38 | -0.02 | 3.3 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Abbott | 3.28
Abbott Architect | 59 | 6.56 | 0.16 | 3 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Abbott | 3.16
Abbott Architect c | 66 | 6.14 | -0.16 | 3.9 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Abbott | 3.74
Abbott Architect c | 66 | 7.45 | -0.15 | 4 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Abbott | 3.85
Abbott Architect c | 66 | 8.97 | -0.23 | 5.3 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Abbott | 5.07
Abbott Architect c | 74 | 5.46 | -0.14 | 3.6 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Abbott | 3.46
Abbott Architect c | 74 | 9.36 | -0.04 | 4.3 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Abbott | 4.26
Abbott Architect c | 74 | 7.1 | -0.1 | 3.1 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Abbott | 3
Abbott Architect c | 83 | 5.4 | 0 | 4.4 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Abbott | 4.4
Abbott Architect c | 83 | 8.28 | -0.02 | 3.8 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Abbott | 3.78
Abbott Architect c | 83 | 6.8 | 1.15 | 3.4 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Abbott | 4.55
Abbott Architect c | 47 | 6.76 | 0.18 | 3.8 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Abbott | 3.98
Abbott Architect c | 47 | 8.68 | 0.29 | 4.4 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Abbott | 4.69
Abbott Architect c | 47 | 5.7 | 0.05 | 4.1 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Abbott | 4.15
Abbott Architect c | 58 | 6.69 | 0.2 | 4.7 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Abbott | 4.9
Abbott Architect c | 58 | 7.2 | 0.23 | 3.7 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Abbott | 3.93
Abbott Architect c | 58 | 9.96 | 0.31 | 3.9 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Abbott | 4.21
Abbott Architect c | 42 | 10.35 | 0.07 | 3.8 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Abbott | 3.87
Abbott Architect c | 42 | 6.92 | 0.1 | 2.9 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Abbott | 3
Abbott Architect c | 26 | 8.74 | 0.11 | 2.7 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Abbott | 2.81
Abbott Architect c | 26 | 5.41 | 0.09 | 3 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Abbott | 3.09
Abbott Architect c | 24 | 5.93 | 0.09 | 2.6 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Abbott | 2.69
Abbott Architect c | 24 | 12.27 | 0.56 | 3.8 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Abbott | 4.36
Abbott Architect c | 24 | 9.77 | 0.24 | 4.3 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Abbott | 4.54
Abbott Architect c | 24 | 5.1 | 0.06 | 3.9 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Abbott | 3.96
Abbott Architect c | 24 | 7.49 | 0.11 | 2.8 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Abbott | 2.91
Abbott Architect c | 21-32 | 12.2 | 0.51 | 3 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Abbott | 3.51
Abbott Architect c | 21-32 | 6.22 | 0.29 | 3.3 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Abbott | 3.59
Abbott Architect c | 21-32 | 5.45 | 0.28 | 3.3 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Abbott | 3.58
Abbott Architect c | 21-32 | 8.43 | 0.29 | 1.6 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Abbott | 1.89
Abbott Architect c | 21-32 | 9.7 | 0.4 | 2.7 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Abbott | 3.1
Abbott Architect c | 5-14 | 5.56 | 0.24 | 5.3 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Abbott | 5.54
Abbott Architect c | 5-14 | 9.36 | 0.2 | 2.2 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Abbott | 2.4
Abbott Architect c | 5-14 | 5.51 | 0.2 | 4.1 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Abbott | 4.3
Abbott Architect c | 5-14 |  |  |  | GH5-04 | 12.03 | 2016 | CAP | Fresh | Abbott | 0
Abbott Architect c | 5-14 |  |  |  | GH5-05 | 5.94 | 2016 | CAP | Fresh | Abbott | 0
Abbott Architect c (enzymatic) | 26 | 6.53 | -0.05 | 1.6 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Abbott | 1.55
Abbott Architect c (enzymatic) | 26 | 8.39 | 0 | 1.7 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Abbott | 1.7
Abbott Architect c (enzymatic) | 26 | 5.53 | -0.12 | 2 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Abbott | 1.88
Abbott Architect c (enzymatic) | 53 | 6.44 | -0.05 | 4.7 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Abbott | 4.65
Abbott Architect c (enzymatic) | 53 | 6.98 | 0.01 | 4.9 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Abbott | 4.91
Abbott Architect c (enzymatic) | 53 | 9.39 | -0.26 | 5.6 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Abbott | 5.34
Abbott Architect c (enzymatic) | 42 | 10.15 | -0.13 | 1.3 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Abbott | 1.17
Abbott Architect c (enzymatic) | 42 | 6.66 | -0.16 | 1.4 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Abbott | 1.24
Abbott Architect c (enzymatic) | 32 | 8.55 | -0.08 | 1.1 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Abbott | 1.02
Abbott Architect c (enzymatic) | 32 | 5.18 | -0.14 | 1.7 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Abbott | 1.56
Abbott Architect c (enzymatic) | 47 | 5.67 | -0.17 | 2 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Abbott | 1.83
Abbott Architect c (enzymatic) | 47 | 11.76 | 0.05 | 2.1 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Abbott | 2.15
Abbott Architect c (enzymatic) | 47 | 9.52 | -0.01 | 2 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Abbott | 1.99
Abbott Architect c (enzymatic) | 47 | 4.9 | -0.14 | 2.4 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Abbott | 2.26
Abbott Architect c (enzymatic) | 47 | 7.22 | -0.16 | 1.7 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Abbott | 1.54
Abbott Architect c (enzymatic) | 59-80 | 11.89 | 0.2 | 1.7 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Abbott | 1.9
Abbott Architect c (enzymatic) | 59-80 | 5.78 | -0.15 | 1.4 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Abbott | 1.25
Abbott Architect c (enzymatic) | 59-80 | 5.07 | -0.1 | 1.7 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Abbott | 1.6
Abbott Architect c (enzymatic) | 59-80 | 8.1 | -0.04 | 1.4 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Abbott | 1.36
Abbott Architect c (enzymatic) | 59-80 | 9.3 | 0 | 1.1 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Abbott | 1.1
Abbott Architect c (enzymatic) | 90-103 | 5.22 | -0.1 | 1.5 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Abbott | 1.4
Abbott Architect c (enzymatic) | 90-103 | 9.24 | 0.08 | 1.3 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Abbott | 1.38
Abbott Architect c (enzymatic) | 90-103 | 5.24 | -0.07 | 1.5 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Abbott | 1.43
Abbott Architect c (enzymatic) | 90-103 | 12.24 | 0.21 | 1.3 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Abbott | 1.51
Abbott Architect c (enzymatic) | 90-103 | 5.83 | -0.11 | 1.6 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Abbott | 1.49
Abbott Architect c (enzymatic) | 97 | 5.17 | -0.1 | 2.1 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Abbott | 2
Abbott Architect c (enzymatic) | 97 | 10.60 | 0.01 | 2.5 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Abbott | 2.51
Abbott Architect c (enzymatic) | 97 | 6.04 | -0.16 | 2 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Abbott | 1.84
Abbott Architect c (enzymatic) | 97 | 12.31 | 0.08 | 2.1 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Abbott | 2.18
Abbott Architect c (enzymatic) | 97 | 7.42 | -0.09 | 1.9 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Abbott | 1.81
Abbott Architect c (enzymatic) | 109-132 | 9.04 | -0.07 | 1.6 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Abbott | 1.53
Abbott Architect c (enzymatic) | 109-132 | 5.88 | -0.13 | 1.6 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Abbott | 1.47
Abbott Architect c (enzymatic) | 109-132 | 11.75 | 0.04 | 1.3 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Abbott | 1.34
Abbott Architect c (enzymatic) | 109-132 | 4.91 | -0.11 | 1.5 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Abbott | 1.39
Abbott Architect c (enzymatic) | 109-132 | 7.48 | -0.1 | 1.4 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Abbott | 1.3
Abbott Architect c (enzymatic) | 146-170 | 6.31 | -0.1 | 1.6 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Abbott | 1.5
Abbott Architect c (enzymatic) | 146-170 | 9.43 | -0.1 | 1.6 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Abbott | 1.5
Abbott Architect c (enzymatic) | 146-170 | 5.25 | -0.09 | 1.7 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Abbott | 1.61
Abbott Architect c (enzymatic) | 146-170 | 8.42 | -0.09 | 1.5 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Abbott | 1.41
Abbott Architect c (enzymatic) | 146-170 | 7.23 | -0.02 | 1.5 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Abbott | 1.48
Abbott Architect c (enzymatic) | 158 | 7.32 | -0.1 | 1.3 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Abbott | 1.2
Abbott Architect c (enzymatic) | 158 | 5.05 | -0.15 | 1.7 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Abbott | 1.55
Abbott Architect c (enzymatic) | 158 | 8.27 | -0.04 | 1.1 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Abbott | 1.06
Abbott Architect c (enzymatic) | 158 | 10.41 | 0.08 | 1.1 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Abbott | 1.18
Abbott Architect c (enzymatic) | 158 | 5.77 | -0.1 | 1.3 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Abbott | 1.2
Abbott Architect c (enzymatic) | 189 | 7.14 | -0.11 | 1.3 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Abbott | 1.19
Abbott Architect c (enzymatic) | 189 | 6.67 | -0.14 | 1.4 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Abbott | 1.26
Abbott Architect c (enzymatic) | 189 | 8.60 | -0.05 | 1.2 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Abbott | 1.15
Abbott Architect c (enzymatic) | 189 | 9.47 | -0.03 | 1.2 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Abbott | 1.17
Abbott Architect c (enzymatic) | 189 | 5.29 | -0.16 | 1.3 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Abbott | 1.14
Abbott Architect c (enzymatic) | 200 | 6.99 | -0.16 | 1.3 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Abbott | 1.14
Abbott Architect c (enzymatic) | 200 | 5.05 | -0.14 | 1.5 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Abbott | 1.36
Abbott Architect c (enzymatic) | 200 | 8.21 | -0.21 | 1.3 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Abbott | 1.09
Abbott Architect c (enzymatic) | 200 | 9.70 | -0.09 | 1.2 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Abbott | 1.11
Abbott Architect c (enzymatic) | 200 | 5.91 | -0.21 | 1.4 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Abbott | 1.19
Abbott Architect c (enzymatic) | 174 | 6.18 | -0.13 | 1.1 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Abbott | 0.97
Abbott Architect c (enzymatic) | 174 | 9.09 | -0.02 | 0.9 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Abbott | 0.88
Abbott Architect c (enzymatic) | 174 | 5.17 | -0.14 | 1.1 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Abbott | 0.96
Abbott Architect c (enzymatic) | 174 | 8.30 | -0.05 | 1 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Abbott | 0.95
Abbott Architect c (enzymatic) | 174 | 7.53 | -0.09 | 1 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Abbott | 0.91
Abbott Architect c System | 205 | 6.39 | -0.02 | 1.3 | GH5-06 | 6.41 | 2019 | CAP | Fresh | Abbott | 1.28
Abbott Architect c System | 205 | 8.70 | 0.1 | 1.1 | GH5-07 | 8.6 | 2019 | CAP | Fresh | Abbott | 1.2
Abbott Architect c System | 205 | 5.38 | -0.04 | 1.4 | GH5-08 | 5.42 | 2019 | CAP | Fresh | Abbott | 1.36
Abbott Architect c System | 205 | 7.41 | 0.03 | 1.2 | GH5-09 | 7.38 | 2019 | CAP | Fresh | Abbott | 1.23
Abbott Architect c System | 205 | 9.93 | 0.18 | 1.2 | GH5-10 | 9.75 | 2019 | CAP | Fresh | Abbott | 1.38
Abbott Architect c System | 234 | 5.14 | -0.11 | 1.6 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Abbott | 1.49
Abbott Architect c System | 234 | 6.37 | -0.04 | 1.6 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Abbott | 1.56
Abbott Architect c System | 234 | 8.35 | 0.14 | 1.3 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Abbott | 1.44
Abbott Architect c System | 234 | 5.40 | -0.07 | 1.4 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Abbott | 1.33
Abbott Architect c System | 234 | 9.66 | 0.18 | 1.2 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Abbott | 1.38
Abbott Architect c System | 235 | 8.1 | 0.13 | 1 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Abbott | 0.13
Abbott Architect c System | 235 | 5.9 | 0.01 | 1 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Abbott | 0.01
Abbott Architect c System | 235 | 5.1 | -0.03 | 1.1 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Abbott | 1.07
Abbott Architect c System | 235 | 7.46 | 0.06 | 1.1 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Abbott | 1.16
Abbott Architect c System | 235 | 9.35 | 0.18 | 1.2 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Abbott | 1.38
Abbott Architect c System | 214 | 7.86 | 0.05 | 1.1 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Abbott | 1.15
Abbott Architect c System | 214 | 5.82 | -0.09 | 1.1 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Abbott | 1.01
Abbott Architect c System | 214 | 8.76 | 0.07 | 0.9 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Abbott | 0.97
Abbott Architect c System | 214 | 5.79 | -0.07 | 1.3 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Abbott | 1.23
Abbott Architect c System | 214 | 9.6 | 0.04 | 1 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Abbott | 0.04
Abbott Architect c System | 236 | 5.36 | -0.12 | 1.7 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Abbott | 1.58
Abbott Architect c System | 236 | 8.31 | 0.05 | 1.3 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Abbott | 1.35
Abbott Architect c System | 236 | 5.04 | -0.11 | 1.7 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Abbott | 1.59
Abbott Architect c System | 236 | 10.22 | 0.17 | 0.9 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Abbott | 1.07
Abbott Architect c System | 236 | 5.91 | -0.08 | 1.3 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Abbott | 1.22
Abbott Architect c System | 78 | 7.21 | 0.1 | 4 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Abbott | 4.1
Abbott Architect c System | 78 | 9.53 | 0.21 | 3.5 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Abbott | 3.71
Abbott Architect c System | 78 | 6.1 | 0.03 | 3.2 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Abbott | 3.23
Abbott Architect c System | 57 | 8.06 | 0.01 | 3.2 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Abbott | 3.21
Abbott Architect c System | 57 | 5.28 | -0.02 | 3.5 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Abbott | 3.48
Abbott Architect c System | 57 | 6.15 | 0.01 | 3.1 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Abbott | 3.11
Abbott Architect c System | 210 | 9.28 | -0.09 | 1.5 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Abbott | 1.41
Abbott Architect c System | 210 | 4.87 | -0.05 | 1.8 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Abbott | 1.75
Abbott Architect c System | 210 | 5.96 | -0.08 | 1.6 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Abbott | 1.52
Abbott Architect c System | 210 | 5.14 | -0.14 | 1.6 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Abbott | 1.46
Abbott Architect c System | 210 | 7.95 | -0.13 | 1.1 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Abbott | 0.97
Abbott Architect c System | 233 | 5.42 | -0.04 | 1.6 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Abbott | 1.56
Abbott Architect c System | 233 | 5.62 | -0.04 | 1.5 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Abbott | 1.46
Abbott Architect c System | 233 | 9.39 | 0.08 | 1.8 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Abbott | 1.88
Abbott Architect c System | 233 | 5.22 | -0.06 | 1.6 | GH5-04 | 5.28 | 2019 | CAP | Fresh | Abbott | 1.54
Abbott Architect c System | 233 | 7.40 | -0.01 | 1.6 | GH5-05 | 7.41 | 2019 | CAP | Fresh | Abbott | 1.59
Abbott Architect c System | 240 | 8.94 | 0.08 | 1.4 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Abbott | 1.48
Abbott Architect c System | 240 | 6.74 | -0.1 | 1.5 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Abbott | 1.4
Abbott Architect c System | 240 | 9.49 | 0.1 | 1.4 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Abbott | 1.5
Abbott Architect c System | 240 | 5.98 | -0.15 | 1.6 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Abbott | 1.45
Abbott Architect c System | 240 | 7.49 | -0.03 | 1.6 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Abbott | 1.57
Abbott Architect c System | 217 | 5.63 | -0.17 | 1.8 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Abbott | 1.63
Abbott Architect c System | 217 | 9.16 | 0.04 | 1.5 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Abbott | 1.54
Abbott Architect c System | 217 | 5.34 | -0.13 | 1.6 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Abbott | 1.47
Abbott Architect c System | 217 | 7.25 | 0 | 1.5 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Abbott | 1.5
Abbott Architect c System | 217 | 9.31 | 0.05 | 1.5 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Abbott | 1.55
Abbott Architect c System | 225 | 5.1 | -0.11 | 1.5 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Abbott | 1.39
Abbott Architect c System | 225 | 6.6 | -0.08 | 1.3 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Abbott | 1.22
Abbott Architect c System | 225 | 8.26 |  | 1.7 | GH5-13 |  | 2021 | CAP | Fresh | Abbott | 1.7
Abbott Architect c System | 225 | 5.93 | -0.05 | 1.2 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Abbott | 1.15
Abbott Architect c System | 225 | 10.35 | 0.23 | 1 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Abbott | 0.23
Abbott Architect I | 53 | 7.07 | 0.49 | 5.6 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Abbott | 6.09
Abbott Architect I | 53 | 9.26 | 0.87 | 6.1 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Abbott | 6.97
Abbott Architect I | 53 | 6.02 | 0.37 | 4.2 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Abbott | 4.57
Abbott Architect I | 58 | 6.47 | -0.02 | 3.7 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Abbott | 3.68
Abbott Architect I | 58 | 6.85 | -0.12 | 3.2 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Abbott | 3.08
Abbott Architect I | 58 | 9.42 | -0.23 | 3.2 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Abbott | 2.97
Abbott Architect I | 35 | 10.28 | 0 | 8.4 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Abbott | 8.4
Abbott Architect I | 35 | 6.86 | 0.04 | 8 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Abbott | 8.04
Abbott Architect I | 28 | 8.79 | 0.16 | 8.5 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Abbott | 8.66
Abbott Architect I | 28 | 5.47 | 0.15 | 8.1 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Abbott | 8.25
Abbott Architect I | 21 | 5.87 | 0.03 | 6.2 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Abbott | 6.23
Abbott Architect I | 21 | 11.5 | -0.21 | 6.5 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Abbott | 6.29
Abbott Architect I | 21 | 9.65 | 0.12 | 7 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Abbott | 7.12
Abbott Architect I | 21 | 5.2 | 0.16 | 8.6 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Abbott | 8.76
Abbott Architect I | 21 | 7.49 | 0.11 | 6.6 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Abbott | 6.71
Abbott Architect I | 16-23 | 11.33 | -0.36 | 4.8 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Abbott | 4.44
Abbott Architect I | 16-23 | 5.78 | -0.15 | 3.9 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Abbott | 3.75
Abbott Architect I | 16-23 | 5.09 | -0.08 | 5.6 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Abbott | 5.52
Abbott Architect I | 16-23 | 8.01 | -0.13 | 5.1 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Abbott | 4.97
Abbott Architect I | 16-23 | 9.17 | -0.13 | 5.5 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Abbott | 5.37
Abbott Architect I | 12-17 | 5.24 | -0.08 | 4.3 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Abbott | 4.22
Abbott Architect I | 12-17 | 8.91 | -0.25 | 5.4 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Abbott | 5.15
Abbott Architect I | 12-17 | 5.26 | -0.05 | 5.9 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Abbott | 5.85
Abbott Architect I | 12-17 | 11.36 | -0.67 | 4.7 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Abbott | 4.03
Abbott Architect I | 12-17 | 5.74 | -0.2 | 4.5 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Abbott | 4.3
Abbott Architect I | 11 | 5.32 | 0.05 | 5.4 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Abbott | 5.45
Abbott Architect I | 11 | 10.28 | -0.31 | 5.2 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Abbott | 4.89
Abbott Architect I | 11 | 5.99 | -0.21 | 4.9 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Abbott | 4.69
Abbott Architect I | 11 | 11.73 | -0.5 | 5.9 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Abbott | 5.4
Abbott Architect I | 11 | 7.29 | -0.22 | 5.6 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Abbott | 5.38
Abbott Architect I | 8-11 | 8.93 | -0.18 | 5.4 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Abbott | 5.22
Abbott Architect I | 8-11 | 5.87 | -0.14 | 5.5 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Abbott | 5.36
Abbott Architect I | 8-11 | 11.09 | -0.62 | 6.3 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Abbott | 5.68
Abbott Architect I | 8-11 |  |  |  | GH5-14 | 5.02 | 2016 | CAP | Fresh | Abbott | 0
Abbott Architect I | 8-11 |  |  |  | GH5-15 | 7.58 | 2016 | CAP | Fresh | Abbott | 0
Abbott Architect I | 0-12 | 6.34 | -0.07 | 5.4 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Abbott | 5.33
Abbott Architect I | 0-12 | 9.36 | -0.17 | 5.7 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Abbott | 5.53
Abbott Architect I | 0-12 | 5.42 | 0.08 | 2.7 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Abbott | 2.78
Abbott Architect I | 0-12 |  |  |  | GH5-04 | 8.51 | 2017 | CAP | Fresh | Abbott | 0
Abbott Architect I | 0-12 |  |  |  | GH5-05 | 7.25 | 2017 | CAP | Fresh | Abbott | 0
Abbott Architect I System | 31 | 8.02 | -0.03 | 4.6 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Abbott | 4.57
Abbott Architect I System | 31 | 5.37 | 0.07 | 5.2 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Abbott | 5.27
Abbott Architect I System | 31 | 6.14 | 0 | 5.5 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Abbott | 5.5
Abbott Architect/Aeroset | 49 | 12.5 |  | 5.1 | GH2-01 | 13 | 2008 | CAP | Fresh | Abbott | 5.1
Abbott Architect/Aeroset | 49 | 5.1 |  | 4.2 | GH2-02 | 5.3 | 2008 | CAP | Fresh | Abbott | 4.2
Abbott Architect/Aeroset | 49 | 9.2 |  | 5.1 | GH2-03 | 8.8 | 2008 | CAP | Fresh | Abbott | 5.1
Abbott Architect/Aeroset | 47 | 11.7 |  | 4.8 | GH2-04 | 11.9 | 2008 | CAP | Fresh | Abbott | 4.8
Abbott Architect/Aeroset | 47 | 6.9 |  | 4.3 | GH2-05 | 6.8 | 2008 | CAP | Fresh | Abbott | 4.3
Abbott Architect/Aeroset | 47 | 5 |  | 5.3 | GH2-06 | 5 | 2008 | CAP | Fresh | Abbott | 5.3
Abbott Architect/Aeroset | 47 | 4.9 |  | 5.8 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Abbott | 5.8
Abbott Architect/Aeroset | 47 | 8.2 |  | 4.3 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Abbott | 4.3
Abbott Architect/Aeroset | 47 | 5.8 |  | 8 | GH2-03 | 6 | 2009 | CAP | Fresh | Abbott | 8
Alere Afinion | 32-119 | 6.3 | -0.11 | 1.9 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Alere | 1.79
Alere Afinion | 32-119 | 9.49 | -0.04 | 1.8 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Alere | 1.76
Alere Afinion | 32-119 | 5.26 | -0.08 | 1.9 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Alere | 1.82
Alere Afinion | 32-119 | 8.49 | -0.02 | 1.7 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Alere | 1.68
Alere Afinion | 32-119 | 7.35 | 0.1 | 1.7 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Alere | 1.8
Alere Afinion | 22 | 7.18 | -0.24 | 1.9 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Alere | 1.66
Alere Afinion | 22 | 5.14 | -0.06 | 1.7 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Alere | 1.64
Alere Afinion | 22 | 8.14 | -0.17 | 1.7 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Alere | 1.53
Alere Afinion | 22 | 10.25 | -0.08 | 1.9 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Alere | 1.82
Alere Afinion | 22 | 5.79 | -0.08 | 1.7 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Alere | 1.62
Alere Afinion | 113 | 7.15 | -0.1 | 2 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Alere | -0.1
Alere Afinion | 113 | 6.74 | -0.07 | 2 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Alere | 1.93
Alere Afinion | 113 | 8.78 | 0.13 | 2.1 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Alere | 2.23
Alere Afinion | 113 | 9.23 | -0.27 | 2.2 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Alere | 1.93
Alere Afinion | 113 | 5.34 | -0.11 | 1.5 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Alere | 1.39
Alere Afinion | 153 | 7.07 | -0.08 | 1.6 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Alere | 1.52
Alere Afinion | 153 | 5.24 | 0.05 | 2.5 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Alere | 2.55
Alere Afinion | 153 | 8.28 | -0.14 | 1.6 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Alere | 1.46
Alere Afinion | 153 | 9.73 | -0.06 | 1.8 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Alere | 1.74
Alere Afinion | 153 | 6.10 | -0.02 | 2.3 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Alere | 2.28
Alere Afinion | 37 | 6.13 | -0.18 | 2.1 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Alere | 1.92
Alere Afinion | 37 | 8.98 | -0.13 | 1.4 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Alere | 1.27
Alere Afinion | 37 | 5.30 | -0.01 | 2.5 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Alere | 2.49
Alere Afinion | 37 | 8.21 | -0.14 | 1.8 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Alere | 1.66
Alere Afinion | 37 | 7.48 | -0.14 | 1.5 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Alere | 1.36
Alere Afinion 2 | 12 |  |  |  | GH5-01 | 5.46 | 2019 | CAP | Fresh | Alere | 0
Alere Afinion 2 | 12 |  |  |  | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Alere | 0
Alere Afinion 2 | 12 | 9.15 | -0.16 | 2 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Alere | 1.84
Alere Afinion 2 | 12 |  |  |  | GH5-04 | 5.28 | 2019 | CAP | Fresh | Alere | 0
Alere Afinion 2 | 12 |  |  |  | GH5-05 | 7.41 | 2019 | CAP | Fresh | Alere | 0
Alere Afinion 2 | 22 | 5.20 | -0.05 | 2.6 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Alere | 2.55
Alere Afinion 2 | 22 | 6.38 | -0.03 | 2.7 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Alere | 2.67
Alere Afinion 2 | 22 | 8.15 | -0.06 | 2.1 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Alere | 2.04
Alere Afinion 2 | 22 |  |  |  | GH5-14 | 5.47 | 2019 | CAP | Fresh | Alere | 0
Alere Afinion 2 | 22 |  |  |  | GH5-15 | 9.48 | 2019 | CAP | Fresh | Alere | 0
Alere Afinion 2 | 24 | 7.85 | -0.12 | 1.9 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Alere | 1.78
Alere Afinion 2 | 24 | 5.88 | -0.01 | 2.1 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Alere | 2.09
Alere Afinion 2 | 24 | 5.11 | -0.02 | 2.4 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Alere | 2.38
Alere Afinion 2 | 24 | 7.4 | 0 | 2.1 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Alere | 2.1
Alere Afinion 2 | 24 | 8.96 | -0.21 | 1.8 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Alere | 1.59
Alere Afinion 2 | 29 | 7.7 | -0.11 | 1.4 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Alere | 1.29
Alere Afinion 2 | 29 | 5.83 | -0.08 | 1.7 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Alere | 1.62
Alere Afinion 2 | 29 | 8.64 | -0.05 | 1.5 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Alere | 1.45
Alere Afinion 2 | 29 | 5.79 | -0.07 | 1.8 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Alere | 1.73
Alere Afinion 2 | 29 | 9.62 | 0.06 | 1.7 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Alere | 1.76
Alere Afinion 2 | 56 | 5.42 | -0.06 | 2.1 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Alere | 2.04
Alere Afinion 2 | 56 | 8.09 | -0.17 | 1.6 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Alere | 1.43
Alere Afinion 2 | 56 | 5.11 | -0.04 | 1.9 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Alere | 1.86
Alere Afinion 2 | 56 | 9.83 | -0.22 | 1.5 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Alere | 1.28
Alere Afinion 2 | 56 | 5.86 | -0.13 | 1.8 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Alere | 1.67
Alere Afinion 2 | 100 | 8.81 | -0.05 | 1.6 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Alere | 1.55
Alere Afinion 2 | 100 | 6.85 | 0.01 | 2.1 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Alere | 2.11
Alere Afinion 2 | 100 | 9.3 | -0.09 | 2 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Alere | -0.09
Alere Afinion 2 | 100 | 6.1 | -0.03 | 2.1 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Alere | 2.07
Alere Afinion 2 | 100 | 7.61 | 0.09 | 1.4 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Alere | 1.49
Alere Afinion 2 | 38 | 5.74 | -0.06 | 2.2 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Alere | 2.14
Alere Afinion 2 | 38 | 9.19 | 0.07 | 1.9 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Alere | 1.97
Alere Afinion 2 | 38 | 5.44 | -0.03 | 1.8 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Alere | 1.77
Alere Afinion 2 | 38 | 7.34 | 0.09 | 1.8 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Alere | 1.89
Alere Afinion 2 | 38 | 9.28 | 0.02 | 2 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Alere | 0.02
Alere Afinion 2 | 114 | 5.18 | -0.03 | 2.2 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Alere | 2.17
Alere Afinion 2 | 114 | 6.59 | -0.09 | 1.8 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Alere | 1.71
Alere Afinion 2 | 114 | 9.12 |  | 2.5 | GH5-13 |  | 2021 | CAP | Fresh | Alere | 2.5
Alere Afinion 2 | 114 | 5.84 | -0.14 | 2.3 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Alere | 2.16
Alere Afinion 2 | 114 | 10.11 | -0.01 | 2.3 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Alere | 2.29
Alere Afinion AS100 | 97 |  |  |  | GH5-01 | 5.46 | 2019 | CAP | Fresh | Alere | 0
Alere Afinion AS100 | 97 |  |  |  | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Alere | 0
Alere Afinion AS100 | 97 | 9.24 | -0.07 | 1.3 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Alere | 1.23
Alere Afinion AS100 | 97 | 5.28 | 0 | 2.2 | GH5-04 | 5.28 | 2019 | CAP | Fresh | Alere | 2.2
Alere Afinion AS100 | 97 | 7.28 | -0.13 | 1.5 | GH5-05 | 7.41 | 2019 | CAP | Fresh | Alere | 1.37
Alere Afinion AS100 | 150 | 9.22 | -0.15 | 1.8 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Alere | 1.65
Alere Afinion AS100 | 150 | 4.96 | 0.04 | 3.1 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Alere | 3.14
Alere Afinion AS100 | 150 | 5.96 | -0.08 | 2.4 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Alere | 2.32
Alere Afinion AS100 | 150 | 5.23 | -0.05 | 2.7 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Alere | 2.65
Alere Afinion AS100 | 150 | 8.02 | -0.06 | 1.6 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Alere | 1.54
Alere Afinion AS100 | 28 | 6.16 | -0.25 | 1.8 | GH5-06 | 6.41 | 2019 | CAP | Fresh | Alere | 1.55
Alere Afinion AS100 | 28 | 8.54 | -0.06 | 1.9 | GH5-07 | 8.6 | 2019 | CAP | Fresh | Alere | 1.84
Alere Afinion AS100 | 28 | 5.36 | -0.06 | 1.7 | GH5-08 | 5.42 | 2019 | CAP | Fresh | Alere | 1.64
Alere Afinion AS100 | 28 | 7.30 | -0.08 | 1.8 | GH5-09 | 7.38 | 2019 | CAP | Fresh | Alere | 1.72
Alere Afinion AS100 | 28 | 9.93 | 0.18 | 1.3 | GH5-10 | 9.75 | 2019 | CAP | Fresh | Alere | 1.48
Alere Afinion AS100 | 136 | 5.11 | -0.14 | 1.7 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Alere | 1.56
Alere Afinion AS100 | 136 | 6.23 | -0.18 | 2 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Alere | 1.82
Alere Afinion AS100 | 136 | 8.14 | -0.07 | 1.7 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Alere | 1.63
Alere Afinion AS100 | 136 | 5.30 | -0.17 | 2.1 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Alere | 1.93
Alere Afinion AS100 | 136 | 9.38 | -0.1 | 1.6 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Alere | 1.5
Alere Afinion AS100 | 79 | 7.86 | -0.11 | 2.2 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Alere | 2.09
Alere Afinion AS100 | 79 | 5.85 | -0.04 | 2.7 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Alere | 2.66
Alere Afinion AS100 | 79 | 5.09 | -0.04 | 2.7 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Alere | 2.66
Alere Afinion AS100 | 79 | 7.39 | -0.01 | 3 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Alere | -0.01
Alere Afinion AS100 | 79 | 9.01 | -0.16 | 2.7 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Alere | 2.54
Alere Afinion AS100 | 26 | 7.65 | -0.16 | 1.7 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Alere | 1.54
Alere Afinion AS100 | 26 | 5.8 | -0.11 | 2.6 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Alere | 2.49
Alere Afinion AS100 | 26 | 8.6 | -0.09 | 2.1 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Alere | 2.01
Alere Afinion AS100 | 26 | 5.73 | -0.13 | 3.1 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Alere | 2.97
Alere Afinion AS100 | 26 | 9.6 | 0.04 | 2 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Alere | 0.04
Alere Afinion AS100 | 112 | 5.3 | -0.18 | 2.1 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Alere | 1.92
Alere Afinion AS100 | 112 | 8.05 | -0.21 | 1.9 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Alere | 1.69
Alere Afinion AS100 | 112 | 5.06 | -0.09 | 2 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Alere | -0.09
Alere Afinion AS100 | 112 | 9.85 | -0.2 | 1.7 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Alere | 1.5
Alere Afinion AS100 | 112 | 5.85 | -0.14 | 2.5 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Alere | 2.36
Alere Afinion AS100 | 105 | 8.64 | -0.22 | 1.8 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Alere | 1.58
Alere Afinion AS100 | 105 | 6.72 | -0.12 | 2.1 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Alere | 1.98
Alere Afinion AS100 | 105 | 9.21 | -0.18 | 1.9 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Alere | 1.72
Alere Afinion AS100 | 105 | 6 | -0.13 | 1.7 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Alere | 1.57
Alere Afinion AS100 | 105 | 7.59 | 0.07 | 1.9 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Alere | 1.97
Alere Afinion AS100 | 20 | 5.57 | -0.23 | 2.3 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Alere | 2.07
Alere Afinion AS100 | 20 | 9.04 | -0.08 | 1.8 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Alere | 1.72
Alere Afinion AS100 | 20 | 5.3 | -0.17 | 1.9 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Alere | 1.73
Alere Afinion AS100 | 20 | 7.22 | -0.03 | 1.7 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Alere | 1.67
Alere Afinion AS100 | 20 | 9.2 | -0.06 | 1.8 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Alere | 1.74
Alere Afinion AS100 | 88 | 5.09 | -0.12 | 2.2 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Alere | 2.08
Alere Afinion AS100 | 88 | 6.5 | -0.18 | 2.2 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Alere | 2.02
Alere Afinion AS100 | 88 | 9.07 |  | 3 | GH5-13 |  | 2021 | CAP | Fresh | Alere | 0
Alere Afinion AS100 | 88 | 5.74 | -0.24 | 1.8 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Alere | 1.56
Alere Afinion AS100 | 88 | 9.9 | -0.22 | 1.9 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Alere | 1.68
Arkray Adams HA-8180 | 16 | 5.39 | 0.07 | 1.1 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Arkray | 1.17
Arkray Adams HA-8180 | 16 | 9.25 | 0.09 | 0.6 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Arkray | 0.69
Arkray Adams HA-8180 | 16 | 5.37 | 0.06 | 1.3 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Arkray | 1.36
Arkray Adams HA-8180 | 16 |  |  |  | GH5-04 | 12.03 | 2016 | CAP | Fresh | Arkray
Arkray Adams HA-8180 | 16 |  |  |  | GH5-05 | 5.94 | 2016 | CAP | Fresh | Arkray
Arkray Adams HA-8180 | 0-17 | 9.18 | 0.07 | 0.5 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Arkray | 0.57
Arkray Adams HA-8180 | 0-17 | 6.03 | 0.02 | 1.5 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Arkray | 1.52
Arkray Adams HA-8180 | 0-17 | 11.81 | 0.1 | 1.1 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Arkray | 1.2
Arkray Adams HA-8180 | 0-17 |  |  |  | GH5-14 | 5.02 | 2016 | CAP | Fresh | Arkray
Arkray Adams HA-8180 | 0-17 |  |  |  | GH5-15 | 7.58 | 2016 | CAP | Fresh | Arkray
Arkray Adams HA-8180 series | 19 | 7.35 | 0.1 | 1.1 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Arkray | 1.2
Arkray Adams HA-8180 series | 19 | 6.82 | 0.01 | 1.6 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Arkray | 1.61
Arkray Adams HA-8180 series | 19 | 8.70 | 0.05 | 1.4 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Arkray | 1.45
Arkray Adams HA-8180 series | 19 |  |  |  | GH5-14 | 9.50 | 2017 | CAP | Fresh | Arkray
Arkray Adams HA-8180 series | 19 |  |  |  | GH5-15 | 5.45 | 2017 | CAP | Fresh | Arkray
Arkray Adams HA-8180 series | 17 | 7.16 | 0.01 | 1.4 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Arkray | 1.41
Arkray Adams HA-8180 series | 17 | 5.19 | 0 | 2 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Arkray | 2
Arkray Adams HA-8180 series | 17 | 8.46 | 0.04 | 1.4 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Arkray | 1.44
Arkray Adams HA-8180 series | 17 | 9.83 | 0.04 | 1 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Arkray | 1.04
Arkray Adams HA-8180 series | 17 | 6.09 | -0.03 | 1.5 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Arkray | 1.47
Arkray Adams HA-8180 series | 11 | 6.26 | -0.05 | 0.8 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Arkray | 0.75
Arkray Adams HA-8180 series | 11 | 9.09 | -0.02 | 1.8 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Arkray | 1.78
Arkray Adams HA-8180 series | 11 | 5.23 | -0.08 | 0.9 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Arkray | 0.82
Arkray Adams HA-8180 series | 11 | 8.38 | 0.03 | 1 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Arkray | 1.03
Arkray Adams HA-8180 series | 11 | 7.64 | 0.02 | 0.7 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Arkray | 0.72
Arkray Adams HA-8180 series | 21 | 9.34 | -0.03 | 1.5 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Arkray | 1.47
Arkray Adams HA-8180 series | 21 | 4.94 | 0.02 | 1.5 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Arkray | 1.52
Arkray Adams HA-8180 series | 21 | 6.07 | 0.03 | 1.4 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Arkray | 1.43
Arkray Adams HA-8180 series | 21 | 5.28 | 0 | 1.4 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Arkray | 1.4
Arkray Adams HA-8180 series | 21 | 8.13 | 0.05 | 1.3 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Arkray | 1.35
Arkray Adams HA-8180 series | 17 | 5.42 | -0.04 | 1.8 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Arkray | 1.76
Arkray Adams HA-8180 series | 17 | 5.54 | -0.12 | 2.9 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Arkray | 2.78
Arkray Adams HA-8180 series | 17 | 9.41 | 0.1 | 1.7 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Arkray | 1.8
Arkray Adams HA-8180 series | 17 | 5.20 | -0.08 | 1.2 | GH5-04 | 5.28 | 2019 | CAP | Fresh | Arkray | 1.12
Arkray Adams HA-8180 series | 17 | 7.43 | 0.02 | 0.9 | GH5-05 | 7.41 | 2019 | CAP | Fresh | Arkray | 0.92
Arkray Adams HA-8180 series | 22 | 8.87 | 0.01 | 1.4 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Arkray | 1.41
Arkray Adams HA-8180 series | 22 | 6.81 | -0.03 | 1.6 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Arkray | 1.57
Arkray Adams HA-8180 series | 22 | 9.36 | -0.03 | 1.5 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Arkray | 1.47
Arkray Adams HA-8180 series | 22 | 6.05 | -0.08 | 1.4 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Arkray | 1.32
Arkray Adams HA-8180 series | 22 | 7.42 | -0.1 | 1.1 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Arkray | 1
Arkray Adams HA-8180 series | 14 | 6.42 | 0.01 | 1.2 | GH5-06 | 6.41 | 2019 | CAP | Fresh | Arkray | 1.21
Arkray Adams HA-8180 series | 14 | 8.67 | 0.07 | 1.5 | GH5-07 | 8.6 | 2019 | CAP | Fresh | Arkray | 1.57
Arkray Adams HA-8180 series | 14 | 5.33 | -0.09 | 1.1 | GH5-08 | 5.42 | 2019 | CAP | Fresh | Arkray | 1.01
Arkray Adams HA-8180 series | 14 | 7.44 | 0.06 | 1.2 | GH5-09 | 7.38 | 2019 | CAP | Fresh | Arkray | 1.26
Arkray Adams HA-8180 series | 14 | 9.89 | 0.14 | 2 | GH5-10 | 9.75 | 2019 | CAP | Fresh | Arkray | 2.14
Arkray Adams HA-8180 series | 22 | 5.26 | 0.01 | 1.4 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Arkray | 1.41
Arkray Adams HA-8180 series | 22 | 6.47 | 0.06 | 1.2 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Arkray | 1.26
Arkray Adams HA-8180 series | 22 | 8.25 | 0.04 | 1.5 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Arkray | 1.54
Arkray Adams HA-8180 series | 22 | 5.53 | 0.06 | 1.3 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Arkray | 1.36
Arkray Adams HA-8180 series | 22 | 9.54 | 0.06 | 1.6 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Arkray | 1.66
Arkray Adams HA-8180 series | 16 | 8.03 | 0.06 | 1.6 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Arkray | 1.66
Arkray Adams HA-8180 series | 16 | 5.92 | 0.03 | 1.7 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Arkray | 1.73
Arkray Adams HA-8180 series | 16 | 5.14 | 0.01 | 1.2 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Arkray | 1.21
Arkray Adams HA-8180 series | 16 | 7.3 | -0.1 | 1.9 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Arkray | 1.8
Arkray Adams HA-8180 series | 16 | 9.18 | 0.01 | 1.3 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Arkray | 1.31
Arkray Adams HA-8180 series | 12 | 7.86 | 0.05 | 1.3 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Arkray | 1.35
Arkray Adams HA-8180 series | 12 | 5.9 | -0.01 | 0.7 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Arkray | 0.69
Arkray Adams HA-8180 series | 12 | 8.73 | 0.04 | 1.3 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Arkray | 1.34
Arkray Adams HA-8180 series | 12 | 5.93 | 0.07 | 1.3 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Arkray | 1.37
Arkray Adams HA-8180 series | 12 | 9.57 | 0.01 | 1 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Arkray | 0.01
Arkray Adams HA-8180 series | 24 | 5.54 | 0.06 | 1.8 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Arkray | 1.86
Arkray Adams HA-8180 series | 24 | 8.32 | 0.06 | 1.4 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Arkray | 1.46
Arkray Adams HA-8180 series | 24 | 5.18 | 0.03 | 1.7 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Arkray | 1.73
Arkray Adams HA-8180 series | 24 | 10.08 | 0.03 | 1.4 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Arkray | 1.43
Arkray Adams HA-8180 series | 24 | 6.01 | 0.02 | 1.7 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Arkray | 1.72
Arkray Adams HA-8180 series | 17 | 5.74 | -0.06 | 1.4 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Arkray | 1.34
Arkray Adams HA-8180 series | 17 | 9.04 | -0.08 | 1.1 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Arkray | 1.02
Arkray Adams HA-8180 series | 17 | 5.37 | -0.1 | 1.6 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Arkray | 1.5
Arkray Adams HA-8180 series | 17 | 7.15 | -0.1 | 1.3 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Arkray | 1.2
Arkray Adams HA-8180 series | 17 | 9.21 | -0.05 | 1.5 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Arkray | 1.45
Arkray Adams HA-8180 series | 26 | 5.2 | -0.01 | 1.4 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Arkray | 1.39
Arkray Adams HA-8180 series | 26 | 6.69 | 0.01 | 1.3 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Arkray | 1.31
Arkray Adams HA-8180 series | 26 | 8.28 |  | 1.4 | GH5-13 |  | 2021 | CAP | Fresh | Arkray | 1.4
Arkray Adams HA-8180 series | 26 | 5.96 | -0.02 | 0.9 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Arkray | 0.88
Arkray Adams HA-8180 series | 26 | 10.06 | -0.06 | 0.8 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Arkray | 0.74
Arkray Adams HA-8190V | 25 | 5.51 | 0.03 | 1.3 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Arkray | 1.33
Arkray Adams HA-8190V | 25 | 8.34 | 0.08 | 1.2 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Arkray | 1.28
Arkray Adams HA-8190V | 25 | 5.14 | -0.01 | 1.2 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Arkray | 1.19
Arkray Adams HA-8190V | 25 |  |  |  | GH5-14 | 10.05 | 2020 | CAP | Fresh | Arkray
Arkray Adams HA-8190V | 25 |  |  |  | GH5-15 | 5.99 | 2020 | CAP | Fresh | Arkray
Arkray Adams HA-8190V | 25 | 5.15 | -0.06 | 1.3 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Arkray | 1.24
Arkray Adams HA-8190V | 25 | 6.68 | 0 | 1.1 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Arkray | 1.1
Arkray Adams HA-8190V | 25 | 8.32 |  | 1.3 | GH5-13 |  | 2021 | CAP | Fresh | Arkray | 1.3
Arkray Adams HA-8190V | 25 |  |  |  | GH5-14 | 5.98 | 2021 | CAP | Fresh | Arkray
Arkray Adams HA-8190V | 25 |  |  |  | GH5-15 | 10.12 | 2021 | CAP | Fresh | Arkray
Arkray Adams The Lab 001 | 17 | 5.49 | 0.01 | 2.2 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Arkray | 2.21
Arkray Adams The Lab 001 | 17 | 8.15 | -0.11 | 1.4 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Arkray | 1.29
Arkray Adams The Lab 001 | 17 | 5.13 | -0.02 | 0.9 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Arkray | 0.88
Arkray Adams The Lab 001 | 17 |  |  |  | GH5-14 | 10.05 | 2020 | CAP | Fresh | Arkray
Arkray Adams The Lab 001 | 17 |  |  |  | GH5-15 | 5.99 | 2020 | CAP | Fresh | Arkray
Arkray The Lab 001 | 13 | 5.12 | -0.09 | 0.7 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Arkray | 0.61
Arkray The Lab 001 | 13 | 6.62 | -0.06 | 1.4 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Arkray | 1.34
Arkray The Lab 001 | 13 | 8.29 |  | 2.5 | GH5-13 |  | 2021 | CAP | Fresh | Arkray | 2.5
Arkray The Lab 001 | 13 |  |  |  | GH5-14 | 5.98 | 2021 | CAP | Fresh | Arkray
Arkray The Lab 001 | 13 |  |  |  | GH5-15 | 10.12 | 2021 | CAP | Fresh | Arkray
Axis-Shield Afinion | 16 | 6.19 | 0.29 | 3.6 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Axis-Shield | 3.89
Axis-Shield Afinion | 16 | 9.87 | 0.07 | 3.3 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Axis-Shield | 3.37
Axis-Shield Afinion | 16 | 7.49 | 0.09 | 3.5 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Axis-Shield | 3.59
Axis-Shield Afinion | 11 | 5.57 | 0.37 | 2.8 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Axis-Shield | 3.17
Axis-Shield Afinion | 11 | 8.75 | 0.05 | 1.9 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Axis-Shield | 1.95
Axis-Shield Afinion | 11 | 6.61 | 0.31 | 1.8 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Axis-Shield | 2.11
Axis-Shield Afinion | 18 | 8.54 | 0.04 | 2.2 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Axis-Shield | 2.24
Axis-Shield Afinion | 18 | 5.6 | 0.2 | 2.7 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Axis-Shield | 2.9
Axis-Shield Afinion | 18 | 6.56 | 0.16 | 3.3 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Axis-Shield | 3.46
Axis-Shield Afinion | 28 | 5.59 | -0.01 | 3.1 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Axis-Shield | 3.09
Axis-Shield Afinion | 28 | 8.98 | -0.42 | 2.9 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Axis-Shield | 2.48
Axis-Shield Afinion | 28 | 7.17 | -0.03 | 2.5 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Axis-Shield | 2.47
Axis-Shield Afinion | 23 | 5.54 | 0.14 | 4.1 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Axis-Shield | 4.24
Axis-Shield Afinion | 23 | 8.1 | -0.2 | 3.4 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Axis-Shield | 3.2
Axis-Shield Afinion | 23 | 5.65 | 0 | 3.5 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Axis-Shield | 3.5
Axis-Shield Afinion | 24 | 7.14 | 0.03 | 3.3 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Axis-Shield | 3.33
Axis-Shield Afinion | 24 | 9.02 | -0.3 | 3 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Axis-Shield | 2.7
Axis-Shield Afinion | 24 | 6.11 | 0.04 | 3.1 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Axis-Shield | 3.14
Axis-Shield Afinion | 34 | 7.64 | -0.41 | 2.4 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Axis-Shield | 1.99
Axis-Shield Afinion | 34 | 5.25 | -0.05 | 2.8 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Axis-Shield | 2.75
Axis-Shield Afinion | 34 | 6.04 | -0.1 | 3.1 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Axis-Shield | 3
Axis-Shield Afinion | 64 | 6.55 | -0.03 | 3.4 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Axis-Shield | 3.37
Axis-Shield Afinion | 64 | 8.37 | -0.02 | 3 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Axis-Shield | 2.98
Axis-Shield Afinion | 64 | 5.73 | 0.08 | 2.5 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Axis-Shield | 2.58
Axis-Shield Afinion | 55 | 6.28 | -0.21 | 3.8 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Axis-Shield | 3.59
Axis-Shield Afinion | 55 | 6.82 | -0.15 | 4.4 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Axis-Shield | 4.25
Axis-Shield Afinion | 55 | 9.45 | -0.2 | 4.7 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Axis-Shield | 4.5
Axis-Shield Afinion | 54 | 6.79 | 0 | 2.4 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Axis-Shield | 2.4
Axis-Shield Afinion | 54 | 10.16 | -0.12 | 2.8 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Axis-Shield | 2.68
Axis-Shield Afinion | 54 | 6.86 | 0.04 | 2.3 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Axis-Shield | 2.34
Axis-Shield Afinion | 15 | 8.43 | -0.2 | 2.3 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Axis-Shield | 2.1
Axis-Shield Afinion | 15 | 5.39 | 0.07 | 3.5 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Axis-Shield | 3.57
Axis-Shield Afinion | 19 | 5.95 | 0.11 | 2 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Axis-Shield | 2.11
Axis-Shield Afinion | 19 | 11.59 | -0.12 | 3.2 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Axis-Shield | 3.08
Axis-Shield Afinion | 19 | 9.44 | -0.09 | 1.9 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Axis-Shield | 1.81
Axis-Shield Afinion | 19 | 5.22 | 0.18 | 2.8 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Axis-Shield | 2.98
Axis-Shield Afinion | 19 | 7.47 | 0.09 | 2 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Axis-Shield | 2.09
Axis-Shield Afinion | 18-76 | 11.68 | -0.01 | 3.1 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Axis-Shield | 3.09
Axis-Shield Afinion | 18-76 | 5.95 | 0.02 | 2.7 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Axis-Shield | 2.72
Axis-Shield Afinion | 18-76 | 5.32 | 0.15 | 3.2 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Axis-Shield | 3.35
Axis-Shield Afinion | 18-76 | 8.13 | -0.01 | 3.5 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Axis-Shield | 3.49
Axis-Shield Afinion | 18-76 | 9.31 | 0.01 | 4 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Axis-Shield | 4.01
Axis-Shield Afinion | 13-57 | 5.26 | -0.06 | 3.1 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Axis-Shield | 3.04
Axis-Shield Afinion | 13-57 | 9.01 | -0.15 | 2.7 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Axis-Shield | 2.55
Axis-Shield Afinion | 13-57 | 5.28 | -0.03 | 3.6 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Axis-Shield | 3.57
Axis-Shield Afinion | 13-57 | 12.08 | 0.05 | 2.5 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Axis-Shield | 2.55
Axis-Shield Afinion | 13-57 | 6.05 | 0.11 | 4.5 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Axis-Shield | 4.61
Axis-Shield Afinion | 13 | 5.27 | 0 | 2.2 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Axis-Shield | 2.2
Axis-Shield Afinion | 13 | 10.47 | -0.12 | 2.1 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Axis-Shield | 1.98
Axis-Shield Afinion | 13 | 6.25 | 0.05 | 2 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Axis-Shield | 2.05
Axis-Shield Afinion | 13 | 12.32 | 0.09 | 3.5 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Axis-Shield | 3.59
Axis-Shield Afinion | 13 | 7.43 | -0.08 | 2.7 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Axis-Shield | 2.62
Axis-Shield Afinion | 15-82 | 8.97 | -0.14 | 1.4 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Axis-Shield | 1.26
Axis-Shield Afinion | 15-82 | 5.99 | -0.02 | 1.9 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Axis-Shield | 1.88
Axis-Shield Afinion | 15-82 | 11.65 | -0.06 | 2 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Axis-Shield | 1.94
Axis-Shield Afinion | 15-82 | 5.03 | 0.01 | 3.1 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Axis-Shield | 3.11
Axis-Shield Afinion | 15-82 | 7.50 | -0.08 | 2.2 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Axis-Shield | 2.12
Bayer (Metrika) A1cNOW | 23 | 5.05 | -0.35 | 4 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Bayer/Metrika | 3.65
Bayer (Metrika) A1cNOW | 23 | 7.49 | -0.81 | 6.2 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Bayer/Metrika | 5.39
Bayer (Metrika) A1cNOW | 23 | 6.31 | 0.66 | 4.9 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Bayer/Metrika | 5.56
Bayer (Metrika) A1cNOW | 19 | 7.96 | -0.54 | 5.5 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Bayer/Metrika | 4.96
Bayer (Metrika) A1cNOW | 19 | 5.14 | -0.26 | 4.9 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Bayer/Metrika | 4.64
Bayer (Metrika) A1cNOW | 19 | 6.02 | -0.38 | 5.2 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Bayer/Metrika | 4.82
Bayer (Metrika) A1cNOW | 19 | 6.34 | 0.04 | 3 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Bayer/Metrika | 3.04
Bayer (Metrika) A1cNOW | 19 | 7.58 | -0.02 | 2.6 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Bayer/Metrika | 2.58
Bayer (Metrika) A1cNOW | 19 | 9.21 | 0.01 | 4 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Bayer/Metrika | 4.01
Bayer (Metrika) A1cNOW | 25 | 4.86 | -0.34 | 4.7 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Bayer/Metrika | 4.36
Bayer (Metrika) A1cNOW | 25 | 8.02 | -0.68 | 6 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Bayer/Metrika | 5.32
Bayer (Metrika) A1cNOW | 25 | 5.85 | -0.45 | 4.9 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Bayer/Metrika | 4.45
Bayer (Metrika) A1cNOW | 28 | 5.14 | -0.46 | 3.8 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Bayer/Metrika | 3.34
Bayer (Metrika) A1cNOW | 28 | 8.49 | -0.91 | 6.3 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Bayer/Metrika | 5.39
Bayer (Metrika) A1cNOW | 28 | 6.59 | -0.61 | 5.4 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Bayer/Metrika | 4.79
Bayer A1cNOW | 16 | 6.37 | -0.74 | 5.3 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Bayer | 4.56
Bayer A1cNOW | 16 | 8.27 | -1.05 | 4.1 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Bayer | 3.05
Bayer A1cNOW | 16 | 5.4 | -0.67 | 7.3 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Bayer | 6.63
Bayer A1cNOW | 13 | 7.43 | -0.62 | 5.7 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Bayer | 5.08
Bayer A1cNOW | 13 | 4.95 | -0.35 | 4.6 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Bayer | 4.25
Bayer A1cNOW | 13 | 5.7 | -0.44 | 5.5 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Bayer | 5.06
Bayer Advia | 24 | 9.9 |  | 6.9 | GH2-01 | 10.7 | 2006 | CAP | Fresh | Bayer | 6.9
Bayer Advia | 24 | 5.2 |  | 3.5 | GH2-02 | 5.30 | 2006 | CAP | Fresh | Bayer | 3.5
Bayer Advia | 24 | 8 |  | 5.8 | GH2-03 | 8.40 | 2006 | CAP | Fresh | Bayer | 5.8
Bayer Advia | 30 | 6.9 |  | 5.7 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Bayer | 5.7
Bayer Advia | 30 | 6.9 |  | 5.6 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Bayer | 5.6
Bayer Advia | 30 | 11 |  | 6.7 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Bayer | 6.7
Bayer Advia | 30 | 6.9 |  | 5.7 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Bayer | 5.7
Bayer Advia | 30 | 6.9 |  | 5.6 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Bayer | 5.6
Bayer Advia | 30 | 11 |  | 6.7 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Bayer | 6.7
Bayer Advia | 33 | 5.5 |  | 4.2 | GH2-01 | 5.4 | 2007 | CAP | Fresh | Bayer | 4.2
Bayer Advia | 33 | 11.3 |  | 4.9 | GH2-02 | 11.6 | 2007 | CAP | Fresh | Bayer | 4.9
Bayer Advia | 33 | 7.7 |  | 5.2 | GH2-03 | 7.6 | 2007 | CAP | Fresh | Bayer | 5.2
Bayer Advia | 31 | 9 |  | 4.5 | GH2-04 | 9.2 | 2007 | CAP | Fresh | Bayer | 4.5
Bayer Advia | 31 | 10.9 |  | 8.1 | GH2-05 | 11.1 | 2007 | CAP | Fresh | Bayer | 8.1
Bayer Advia | 31 | 6.4 |  | 8 | GH2-06 | 6.2 | 2007 | CAP | Fresh | Bayer | 8
Bayer DCA 2000 | 156 | 10.4 |  | 3 | GH2-01 | 10.7 | 2006 | CAP | Fresh | Bayer | 3
Bayer DCA 2000 | 156 | 4.9 |  | 4.6 | GH2-02 | 5.30 | 2006 | CAP | Fresh | Bayer | 4.6
Bayer DCA 2000 | 156 | 8.1 |  | 3.2 | GH2-03 | 8.40 | 2006 | CAP | Fresh | Bayer | 3.2
Bayer DCA 2000 | 159 | 6.7 |  | 2.9 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Bayer | 2.9
Bayer DCA 2000 | 159 | 6.7 |  | 2.7 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Bayer | 2.7
Bayer DCA 2000 | 159 | 11.4 |  | 3.5 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Bayer | 3.5
Bayer DCA 2000 | 159 | 6.7 |  | 2.9 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Bayer | 2.9
Bayer DCA 2000 | 159 | 6.7 |  | 2.7 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Bayer | 2.7
Bayer DCA 2000 | 159 | 11.4 |  | 3.5 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Bayer | 3.5
Bayer DCA 2000 | 227 | 5.2 |  | 2.9 | GH2-01 | 5.4 | 2007 | CAP | Fresh | Bayer | 2.9
Bayer DCA 2000 | 227 | 11.5 |  | 4.8 | GH2-02 | 11.6 | 2007 | CAP | Fresh | Bayer | 4.8
Bayer DCA 2000 | 227 | 7.4 |  | 2.7 | GH2-03 | 7.6 | 2007 | CAP | Fresh | Bayer | 2.7
Bayer DCA 2000 | 171 | 8.8 |  | 3 | GH2-04 | 9.2 | 2007 | CAP | Fresh | Bayer | 3
Bayer DCA 2000 | 171 | 11 |  | 4.5 | GH2-05 | 11.1 | 2007 | CAP | Fresh | Bayer | 4.5
Bayer DCA 2000 | 171 | 6.1 |  | 2.9 | GH2-06 | 6.2 | 2007 | CAP | Fresh | Bayer | 2.9
Beckman AU | 67 | 6.45 | -0.13 | 5.4 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Beckman Coulter | 5.27
Beckman AU | 67 | 8.26 | -0.13 | 5.7 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Beckman Coulter | 5.57
Beckman AU | 67 | 5.5 | -0.15 | 3.7 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Beckman Coulter | 3.55
Beckman AU | 65 | 6.67 | -0.12 | 4.3 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Beckman Coulter | 4.18
Beckman AU | 65 | 10 | -0.28 | 3.9 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Beckman Coulter | 3.62
Beckman AU | 65 | 6.68 | -0.14 | 3.9 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Beckman Coulter | 3.76
Beckman AU | 40 | 8.29 | -0.34 | 3.6 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Beckman Coulter | 3.26
Beckman AU | 40 | 5.17 | -0.15 | 3.1 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Beckman Coulter | 2.95
Beckman AU | 42 | 5.81 | -0.03 | 4.5 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Beckman Coulter | 4.47
Beckman AU | 42 | 11.62 | -0.09 | 3.6 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Beckman Coulter | 3.51
Beckman AU | 42 | 9.42 | -0.11 | 3.5 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Beckman Coulter | 3.39
Beckman AU | 42 | 5.09 | 0.05 | 3.7 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Beckman Coulter | 3.75
Beckman AU | 42 | 7.32 | -0.06 | 3.6 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Beckman Coulter | 3.54
Beckman AU | 42-67 | 11.71 | 0.02 | 3.2 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Beckman Coulter | 3.22
Beckman AU | 42-67 | 5.98 | 0.05 | 2.9 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Beckman Coulter | 2.95
Beckman AU | 42-67 | 5.21 | 0.04 | 3.5 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Beckman Coulter | 3.54
Beckman AU | 42-67 | 8.06 | -0.08 | 2.5 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Beckman Coulter | 2.42
Beckman AU | 42-67 | 9.21 | -0.09 | 2.4 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Beckman Coulter | 2.31
Beckman AU | 50-83 | 5.29 | -0.03 | 3.4 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Beckman Coulter | 3.37
Beckman AU | 50-83 | 9.17 | 0.01 | 3.7 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Beckman Coulter | 3.71
Beckman AU | 50-83 | 5.29 | -0.02 | 3.2 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Beckman Coulter | 3.18
Beckman AU | 50-83 | 12.10 | 0.07 | 3.7 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Beckman Coulter | 3.77
Beckman AU | 50-83 | 5.90 | -0.04 | 3.3 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Beckman Coulter | 3.26
Beckman AU | 55 | 5.30 | 0.03 | 3.1 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Beckman Coulter | 3.13
Beckman AU | 55 | 10.59 | 0 | 4.1 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Beckman Coulter | 4.1
Beckman AU | 55 | 6.21 | 0.01 | 4.5 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Beckman Coulter | 4.51
Beckman AU | 55 | 12.32 | 0.09 | 4.1 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Beckman Coulter | 4.19
Beckman AU | 55 | 7.47 | -0.04 | 2.8 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Beckman Coulter | 2.76
Beckman AU | 57-89 | 8.98 | -0.13 | 4 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Beckman Coulter | 3.87
Beckman AU | 57-89 | 5.87 | -0.14 | 3.9 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Beckman Coulter | 3.76
Beckman AU | 57-89 | 11.60 | -0.11 | 4.6 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Beckman Coulter | 4.49
Beckman AU | 57-89 | 4.92 | -0.1 | 4.5 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Beckman Coulter | 4.4
Beckman AU | 57-89 | 7.42 | -0.16 | 4 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Beckman Coulter | 3.84
Beckman AU | 59-92 | 6.58 | 0.17 | 3.9 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Beckman Coulter | 4.07
Beckman AU | 59-92 | 9.6 | 0.07 | 3.4 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Beckman Coulter | 3.47
Beckman AU | 59-92 | 5.44 | 0.1 | 3.4 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Beckman Coulter | 3.5
Beckman AU | 59-92 | 8.65 | 0.14 | 4.1 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Beckman Coulter | 4.24
Beckman AU | 59-92 | 7.46 | 0.21 | 4.1 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Beckman Coulter | 4.31
Beckman AU | 87 | 7.16 | 0.01 | 3.9 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Beckman Coulter | 3.91
Beckman AU | 87 | 5.27 | 0.08 | 4.3 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Beckman Coulter | 4.38
Beckman AU | 87 | 8.32 | -0.1 | 4.7 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Beckman Coulter | 4.6
Beckman AU | 87 | 9.66 | -0.13 | 4 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Beckman Coulter | 3.87
Beckman AU | 87 | 6.16 | 0.04 | 3.4 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Beckman Coulter | 3.44
Beckman AU | 62 | 6.33 | 0.02 | 3.7 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Beckman Coulter | 3.72
Beckman AU | 62 | 9.04 | -0.07 | 4 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Beckman Coulter | 3.93
Beckman AU | 62 | 5.37 | 0.06 | 3.1 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Beckman Coulter | 3.16
Beckman AU | 62 | 8.34 | -0.01 | 4.4 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Beckman Coulter | 4.39
Beckman AU | 62 | 7.59 | -0.03 | 3.8 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Beckman Coulter | 3.77
Beckman AU | 63 | 7.51 | 0.09 | 3.2 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Beckman Coulter | 3.29
Beckman AU | 63 | 5.18 | -0.02 | 2.8 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Beckman Coulter | 2.78
Beckman AU | 63 | 8.30 | -0.01 | 2.9 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Beckman Coulter | 2.89
Beckman AU | 63 | 10.35 | 0.02 | 2.8 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Beckman Coulter | 2.82
Beckman AU | 63 | 5.90 | 0.03 | 3.3 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Beckman Coulter | 3.33
Beckman AU | 91 | 7.26 | 0.01 | 3 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Beckman Coulter | 0.01
Beckman AU | 91 | 6.83 | 0.02 | 3.6 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Beckman Coulter | 3.62
Beckman AU | 91 | 8.55 | -0.1 | 2.3 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Beckman Coulter | 2.2
Beckman AU | 91 | 9.43 | -0.07 | 2.3 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Beckman Coulter | 2.23
Beckman AU | 91 | 5.43 | -0.02 | 2.9 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Beckman Coulter | 2.88
Beckman AU HbA1c Advanced | 10 | 5.23 | -0.02 | 2.6 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Beckman Coulter | 2.58
Beckman AU HbA1c Advanced | 10 | 6.40 | -0.01 | 3.7 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Beckman Coulter | 3.69
Beckman AU HbA1c Advanced | 10 |  |  |  | GH5-13 | 8.21 | 2019 | CAP | Fresh | Beckman Coulter | 0
Beckman AU HbA1c Advanced | 10 |  |  |  | GH5-14 | 5.47 | 2019 | CAP | Fresh | Beckman Coulter | 0
Beckman AU HbA1c Advanced | 10 |  |  |  | GH5-15 | 9.48 | 2019 | CAP | Fresh | Beckman Coulter | 0
Beckman AU HbA1c Advanced | 15 | 7.84 | -0.13 | 1.8 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Beckman Coulter | 1.67
Beckman AU HbA1c Advanced | 15 | 5.78 | -0.11 | 2.2 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Beckman Coulter | 2.09
Beckman AU HbA1c Advanced | 15 | 5.03 | -0.1 | 2.4 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Beckman Coulter | 2.3
Beckman AU HbA1c Advanced | 15 |  |  |  | GH5-04 | 7.4 | 2020 | CAP | Fresh | Beckman Coulter | 0
Beckman AU HbA1c Advanced | 15 |  |  |  | GH5-05 | 9.17 | 2020 | CAP | Fresh | Beckman Coulter | 0
Beckman AU HbA1c Advanced | 13 | 7.67 | -0.14 | 3.7 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Beckman Coulter | 3.56
Beckman AU HbA1c Advanced | 13 | 5.87 | -0.04 | 3.2 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Beckman Coulter | 3.16
Beckman AU HbA1c Advanced | 13 | 8.48 | -0.21 | 5.1 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Beckman Coulter | 4.89
Beckman AU HbA1c Advanced | 13 | 5.82 | -0.04 | 4.2 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Beckman Coulter | 4.16
Beckman AU HbA1c Advanced | 13 | 9.33 | -0.23 | 4.6 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Beckman Coulter | 4.37
Beckman AU HbA1c Advanced | 24 | 5.4 | -0.08 | 2.8 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Beckman Coulter | 2.72
Beckman AU HbA1c Advanced | 24 | 8.15 | -0.11 | 2.9 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Beckman Coulter | 2.79
Beckman AU HbA1c Advanced | 24 | 5.1 | -0.05 | 2.9 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Beckman Coulter | 2.85
Beckman AU HbA1c Advanced | 24 | 9.88 | -0.17 | 3.5 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Beckman Coulter | 3.33
Beckman AU HbA1c Advanced | 24 | 5.92 | -0.07 | 4.1 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Beckman Coulter | 4.03
Beckman AU HbA1c Advanced | 28 | 8.84 | -0.02 | 3.6 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Beckman Coulter | 3.58
Beckman AU HbA1c Advanced | 28 | 6.8 | -0.04 | 2.8 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Beckman Coulter | 2.76
Beckman AU HbA1c Advanced | 28 | 9.17 | -0.22 | 2.5 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Beckman Coulter | 2.28
Beckman AU HbA1c Advanced | 28 | 6.07 | -0.06 | 3.5 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Beckman Coulter | 3.44
Beckman AU HbA1c Advanced | 28 | 7.45 | -0.07 | 2.9 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Beckman Coulter | 2.83
Beckman AU HbA1c Advanced | 19 | 5.71 | -0.09 | 2.3 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Beckman Coulter | 2.21
Beckman AU HbA1c Advanced | 19 | 8.87 | -0.25 | 2.3 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Beckman Coulter | 2.05
Beckman AU HbA1c Advanced | 19 | 5.42 | -0.05 | 2.4 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Beckman Coulter | 2.35
Beckman AU HbA1c Advanced | 19 | 7.19 | -0.06 | 2.4 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Beckman Coulter | 2.34
Beckman AU HbA1c Advanced | 19 | 9.03 | -0.23 | 2.4 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Beckman Coulter | 2.17
Beckman AU HbA1c Advanced | 27 | 5.11 | -0.1 | 2.9 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Beckman Coulter | 2.8
Beckman AU HbA1c Advanced | 27 | 6.61 | -0.07 | 2.4 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Beckman Coulter | 2.33
Beckman AU HbA1c Advanced | 27 | 8.06 |  | 3.8 | GH5-13 |  | 2021 | CAP | Fresh | Beckman Coulter | 3.8
Beckman AU HbA1c Advanced | 27 | 5.96 | -0.02 | 2.2 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Beckman Coulter | 2.18
Beckman AU HbA1c Advanced | 27 | 9.89 | -0.23 | 1.8 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Beckman Coulter | 1.57
Beckman AU systems | 17 | 5.26 | 0.06 | 5.3 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Beckman Coulter | 5.36
Beckman AU systems | 17 | 8.65 | -0.05 | 5 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Beckman Coulter | 4.95
Beckman AU systems | 17 | 6.21 | -0.09 | 4.8 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Beckman Coulter | 4.71
Beckman AU systems | 19 | 8.41 | -0.09 | 3.4 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Beckman Coulter | 3.31
Beckman AU systems | 19 | 5.43 | 0.03 | 3.2 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Beckman Coulter | 3.23
Beckman AU systems | 19 | 6.43 | 0.03 | 3.1 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Beckman Coulter | 3.13
Beckman AU systems | 37 | 5.34 | -0.06 | 3.3 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Beckman Coulter | 3.24
Beckman AU systems | 37 | 8.15 | -0.15 | 3.9 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Beckman Coulter | 3.75
Beckman AU systems | 37 | 6.66 | 1.01 | 3.6 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Beckman Coulter | 4.61
Beckman AU systems | 24 | 5.81 | -0.49 | 4.4 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Beckman Coulter | 3.91
Beckman AU systems | 25 | 6.30 | 0 | 5 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Beckman Coulter | 5
Beckman AU systems | 24 | 6.94 | -0.66 | 4.9 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Beckman Coulter | 4.24
Beckman AU systems | 25 | 7.64 | 0.04 | 4.5 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Beckman Coulter | 4.54
Beckman AU systems | 24 | 8.22 | -0.98 | 6.2 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Beckman Coulter | 5.22
Beckman AU systems | 25 | 9.22 | 0.02 | 5.5 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Beckman Coulter | 5.52
Beckman AU systems | 35 | 5.59 | -0.01 | 4.2 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Beckman Coulter | 4.19
Beckman AU systems | 35 | 9.31 | -0.09 | 3.6 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Beckman Coulter | 3.51
Beckman AU systems | 35 | 7.2 | 0 | 3.4 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Beckman Coulter | 3.4
Beckman AU systems | 37 | 6.92 | -0.19 | 5.5 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Beckman Coulter | 5.31
Beckman AU systems | 37 | 9.16 | -0.16 | 4.6 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Beckman Coulter | 4.44
Beckman AU systems | 37 | 5.89 | -0.18 | 5 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Beckman Coulter | 4.82
Beckman AU systems | 45 | 7.76 | -0.29 | 3.6 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Beckman Coulter | 3.31
Beckman AU systems | 45 | 5.15 | -0.15 | 3 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Beckman Coulter | 2.85
Beckman AU systems | 45 | 5.89 | -0.25 | 4.4 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Beckman Coulter | 4.15
Beckman AU Systems - Beckman reagent | 70 | 8.91 | 0.05 | 2.7 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Beckman Coulter | 2.75
Beckman AU Systems - Beckman reagent | 70 | 6.92 | 0.08 | 2.3 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Beckman Coulter | 2.38
Beckman AU Systems - Beckman reagent | 70 | 9.33 | -0.06 | 3.1 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Beckman Coulter | 3.04
Beckman AU Systems - Beckman reagent | 70 | 6.16 | 0.03 | 2.4 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Beckman Coulter | 2.43
Beckman AU Systems - Beckman reagent | 70 | 7.56 | 0.04 | 2.4 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Beckman Coulter | 2.44
Beckman AU Systems - Beckman reagent | 89 | 9.04 | -0.33 | 3.7 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Beckman Coulter | 3.37
Beckman AU Systems - Beckman reagent | 89 | 5.03 | 0.11 | 3.4 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Beckman Coulter | 3.51
Beckman AU Systems - Beckman reagent | 89 | 6.07 | 0.03 | 3.6 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Beckman Coulter | 3.63
Beckman AU Systems - Beckman reagent | 89 | 5.36 | 0.08 | 4 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Beckman Coulter | 4.08
Beckman AU Systems - Beckman reagent | 89 | 7.88 | -0.2 | 4.2 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Beckman Coulter | 4
Beckman AU Systems - Beckman reagent | 87 | 5.52 | 0.06 | 4.3 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Beckman Coulter | 4.36
Beckman AU Systems - Beckman reagent | 87 | 5.73 | 0.07 | 3.7 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Beckman Coulter | 3.77
Beckman AU Systems - Beckman reagent | 87 | 9.01 | -0.3 | 3.5 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Beckman Coulter | 3.2
Beckman AU Systems - Beckman reagent | 87 | 5.29 | 0.01 | 3.3 | GH5-04 | 5.28 | 2019 | CAP | Fresh | Beckman Coulter | 3.31
Beckman AU Systems - Beckman reagent | 87 | 7.34 | -0.07 | 3.8 | GH5-05 | 7.41 | 2019 | CAP | Fresh | Beckman Coulter | 3.73
Beckman AU Systems - Beckman reagent | 53 | 5.81 | 0.01 | 2.8 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Beckman Coulter | 2.81
Beckman AU Systems - Beckman reagent | 53 | 9.13 | 0.01 | 3.9 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Beckman Coulter | 3.91
Beckman AU Systems - Beckman reagent | 53 | 5.5 | 0.03 | 2.5 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Beckman Coulter | 2.53
Beckman AU Systems - Beckman reagent | 53 | 7.3 | 0.05 | 3.1 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Beckman Coulter | 3.15
Beckman AU Systems - Beckman reagent | 53 | 9.28 | 0.02 | 3.9 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Beckman Coulter | 3.92
Beckman AU Systems - Beckman reagent | 66 | 5.19 | -0.02 | 2.9 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Beckman Coulter | 2.88
Beckman AU Systems - Beckman reagent | 66 | 6.59 | -0.09 | 3.2 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Beckman Coulter | 3.11
Beckman AU Systems - Beckman reagent | 66 | 8.11 |  | 3.1 | GH5-13 |  | 2021 | CAP | Fresh | Beckman Coulter | 3.1
Beckman AU Systems - Beckman reagent | 66 | 5.93 | -0.05 | 2.4 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Beckman Coulter | 2.35
Beckman AU Systems - Beckman reagent | 66 | 9.78 | -0.34 | 3.4 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Beckman Coulter | 3.06
Beckman AU Systems - Beckman reagent | 64 | 6.33 | -0.08 | 3.9 | GH5-06 | 6.41 | 2019 | CAP | Fresh | Beckman Coulter | 3.82
Beckman AU Systems - Beckman reagent | 64 | 8.42 | -0.18 | 4 | GH5-07 | 8.6 | 2019 | CAP | Fresh | Beckman Coulter | 3.82
Beckman AU Systems - Beckman reagent | 64 | 5.32 | -0.1 | 3.8 | GH5-08 | 5.42 | 2019 | CAP | Fresh | Beckman Coulter | 3.7
Beckman AU Systems - Beckman reagent | 64 | 7.29 | -0.09 | 3.8 | GH5-09 | 7.38 | 2019 | CAP | Fresh | Beckman Coulter | 3.71
Beckman AU Systems - Beckman reagent | 64 | 9.42 | -0.33 | 5 | GH5-10 | 9.75 | 2019 | CAP | Fresh | Beckman Coulter | 4.67
Beckman AU Systems - Beckman reagent | 78 | 5.20 | -0.05 | 3 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Beckman Coulter | 2.95
Beckman AU Systems - Beckman reagent | 78 | 6.41 | 0 | 2.7 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Beckman Coulter | 2.7
Beckman AU Systems - Beckman reagent | 78 | 8.07 | -0.14 | 3.8 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Beckman Coulter | 3.66
Beckman AU Systems - Beckman reagent | 78 | 5.42 | -0.05 | 3 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Beckman Coulter | 2.95
Beckman AU Systems - Beckman reagent | 78 | 9.30 | -0.18 | 4 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Beckman Coulter | 3.82
Beckman AU Systems - Beckman reagent | 72 | 8.08 | 0.11 | 2.7 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Beckman Coulter | 2.81
Beckman AU Systems - Beckman reagent | 72 | 5.99 | 0.1 | 2.8 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Beckman Coulter | 2.9
Beckman AU Systems - Beckman reagent | 72 | 5.2 | 0.07 | 2.7 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Beckman Coulter | 2.77
Beckman AU Systems - Beckman reagent | 72 | 7.51 | 0.11 | 3 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Beckman Coulter | 0.11
Beckman AU Systems - Beckman reagent | 72 | 9.13 | -0.04 | 2.6 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Beckman Coulter | 2.56
Beckman AU Systems - Beckman reagent | 58 | 7.84 | 0.03 | 3.1 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Beckman Coulter | 3.13
Beckman AU Systems - Beckman reagent | 58 | 5.98 | 0.07 | 3 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Beckman Coulter | 0.07
Beckman AU Systems - Beckman reagent | 58 | 8.68 | -0.01 | 2.9 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Beckman Coulter | 2.89
Beckman AU Systems - Beckman reagent | 58 | 5.96 | 0.1 | 3.8 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Beckman Coulter | 3.9
Beckman AU Systems - Beckman reagent | 58 | 9.47 | -0.09 | 3 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Beckman Coulter | -0.09
Beckman AU Systems - Beckman reagent | 72 | 5.52 | 0.04 | 4.4 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Beckman Coulter | 4.44
Beckman AU Systems - Beckman reagent | 72 | 8.23 | -0.03 | 3.1 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Beckman Coulter | 3.07
Beckman AU Systems - Beckman reagent | 72 | 5.21 | 0.06 | 3.3 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Beckman Coulter | 3.36
Beckman AU Systems - Beckman reagent | 72 | 9.97 | -0.08 | 2.7 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Beckman Coulter | 2.62
Beckman AU Systems - Beckman reagent | 72 | 6.01 | 0.02 | 3.3 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Beckman Coulter | 3.32
Beckman Synchron CX System | 32 | 5 |  | 3.4 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Beckman Coulter | 3.4
Beckman Synchron CX System | 32 | 7.9 |  | 4.4 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Beckman Coulter | 4.4
Beckman Synchron CX System | 32 | 5.8 |  | 5.3 | GH2-03 | 6 | 2009 | CAP | Fresh | Beckman Coulter | 5.3
Beckman Synchron CX System | 27 | 6.45 | -0.15 | 5.1 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Beckman Coulter | 4.95
Beckman Synchron CX System | 27 | 9.21 | -0.29 | 5.6 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Beckman Coulter | 5.31
Beckman Synchron CX System | 27 | 7.18 | -0.22 | 5.5 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Beckman Coulter | 5.28
Beckman Synchron CX System | 24 | 5.79 | -0.11 | 5.2 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Beckman Coulter | 5.09
Beckman Synchron CX System | 24 | 9.51 | -0.29 | 4.3 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Beckman Coulter | 4.01
Beckman Synchron CX System | 24 | 7.11 | -0.29 | 6 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Beckman Coulter | 5.71
Beckman Synchron CX System | 16 | 5.25 | 0.05 | 4 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Beckman Coulter | 4.05
Beckman Synchron CX System | 16 | 8.15 | -0.55 | 4.1 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Beckman Coulter | 3.55
Beckman Synchron CX System | 16 | 6.12 | -0.18 | 3.6 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Beckman Coulter | 3.42
Beckman Synchron CX System | 12 | 8.43 | -0.07 | 3.9 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Beckman Coulter | 3.83
Beckman Synchron CX System | 12 | 5.42 | 0.02 | 7.2 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Beckman Coulter | 7.22
Beckman Synchron CX System | 12 | 6.48 | 0.08 | 7.2 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Beckman Coulter | 7.28
Beckman Synchron LX System | 85 | 5 |  | 4.9 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Beckman Coulter | 4.9
Beckman Synchron LX System | 85 | 8.2 |  | 3.7 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Beckman Coulter | 3.7
Beckman Synchron LX System | 85 | 5.8 |  | 3.8 | GH2-03 | 6 | 2009 | CAP | Fresh | Beckman Coulter | 3.8
Beckman Synchron LX System | 87 | 6.49 | -0.11 | 3.8 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Beckman Coulter | 3.69
Beckman Synchron LX System | 87 | 9.23 | -0.27 | 3.5 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Beckman Coulter | 3.23
Beckman Synchron LX System | 87 | 7.1 | -0.3 | 3.9 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Beckman Coulter | 3.6
Beckman Synchron LX System | 69 | 5.83 | -0.07 | 4.4 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Beckman Coulter | 4.33
Beckman Synchron LX System | 69 | 9.71 | -0.09 | 4.3 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Beckman Coulter | 4.21
Beckman Synchron LX System | 69 | 7.16 | -0.24 | 3.6 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Beckman Coulter | 3.36
Beckman Synchron LX System | 59 | 5.42 | 0.22 | 5.2 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Beckman Coulter | 5.42
Beckman Synchron LX System | 59 | 8.61 | -0.09 | 5.3 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Beckman Coulter | 5.21
Beckman Synchron LX System | 59 | 6.32 | 0.02 | 5 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Beckman Coulter | 5.02
Beckman Synchron LX System | 39 | 8.55 | 0.05 | 4.3 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Beckman Coulter | 4.35
Beckman Synchron LX System | 39 | 5.52 | 0.12 | 5.1 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Beckman Coulter | 5.22
Beckman Synchron LX System | 39 | 6.41 | 0.01 | 4.5 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Beckman Coulter | 4.51
Beckman Synchron LX System | 29 | 6.20 | -0.1 | 4.6 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Beckman Coulter | 4.5
Beckman Synchron LX System | 29 | 7.69 | 0.09 | 3.7 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Beckman Coulter | 3.79
Beckman Synchron LX System | 29 | 9.54 | 0.34 | 4.5 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Beckman Coulter | 4.84
Beckman Synchron LX System | 21 | 5.77 | 0.17 | 4.7 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Beckman Coulter | 4.87
Beckman Synchron LX System | 21 | 9.45 | 0.05 | 2.9 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Beckman Coulter | 2.95
Beckman Synchron LX System | 21 | 7.15 | -0.05 | 2.8 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Beckman Coulter | 2.75
Beckman Synchron LX System | 17 | 5.49 | 0.09 | 4 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Beckman Coulter | 4.09
Beckman Synchron LX System | 17 | 8.22 | -0.08 | 3.2 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Beckman Coulter | 3.12
Beckman Synchron LX System | 17 | 5.62 | -0.03 | 6.5 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Beckman Coulter | 6.47
Beckman Synchron LX System | 10 | 6.91 | -0.2 | 3.5 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Beckman Coulter | 3.3
Beckman Synchron LX System | 10 | 9.41 | 0.09 | 2.2 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Beckman Coulter | 2.29
Beckman Synchron LX System | 10 | 6.29 | 0.22 | 8.2 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Beckman Coulter | 8.42
Beckman Synchron System | 291 | 10.6 |  | 4.8 | GH2-01 | 10.7 | 2006 | CAP | Fresh | Beckman Coulter | 4.8
Beckman Synchron System | 291 | 5.2 |  | 5.4 | GH2-02 | 5.30 | 2006 | CAP | Fresh | Beckman Coulter | 5.4
Beckman Synchron System | 291 | 7.9 |  | 4.8 | GH2-03 | 8.40 | 2006 | CAP | Fresh | Beckman Coulter | 4.8
Beckman Synchron System | 310 | 6.6 |  | 4.4 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Beckman Coulter | 4.4
Beckman Synchron System | 310 | 6.6 |  | 4.3 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Beckman Coulter | 4.3
Beckman Synchron System | 310 | 11.4 |  | 4.7 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Beckman Coulter | 4.7
Beckman Synchron System | 310 | 6.6 |  | 4.4 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Beckman Coulter | 4.4
Beckman Synchron System | 310 | 6.6 |  | 4.3 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Beckman Coulter | 4.3
Beckman Synchron System | 310 | 11.4 |  | 4.7 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Beckman Coulter | 4.7
Beckman Synchron System | 385 | 5.3 |  | 4.6 | GH2-01 | 5.4 | 2007 | CAP | Fresh | Beckman Coulter | 4.6
Beckman Synchron System | 385 | 11.7 |  | 5.4 | GH2-02 | 11.6 | 2007 | CAP | Fresh | Beckman Coulter | 5.4
Beckman Synchron System | 385 | 7.3 |  | 4.6 | GH2-03 | 7.6 | 2007 | CAP | Fresh | Beckman Coulter | 4.6
Beckman Synchron System | 353 | 9 |  | 4.7 | GH2-04 | 9.2 | 2007 | CAP | Fresh | Beckman Coulter | 4.7
Beckman Synchron System | 353 | 11.2 |  | 4.4 | GH2-05 | 11.1 | 2007 | CAP | Fresh | Beckman Coulter | 4.4
Beckman Synchron System | 353 | 6 |  | 4.2 | GH2-06 | 6.2 | 2007 | CAP | Fresh | Beckman Coulter | 4.2
Beckman Synchron System | 369 | 13.2 |  | 4.6 | GH2-01 | 13 | 2008 | CAP | Fresh | Beckman Coulter | 4.6
Beckman Synchron System | 369 | 5.3 |  | 4.8 | GH2-02 | 5.3 | 2008 | CAP | Fresh | Beckman Coulter | 4.8
Beckman Synchron System | 369 | 8.7 |  | 4.5 | GH2-03 | 8.8 | 2008 | CAP | Fresh | Beckman Coulter | 4.5
Beckman Synchron System | 389 | 11.8 |  | 4.5 | GH2-04 | 11.9 | 2008 | CAP | Fresh | Beckman Coulter | 4.5
Beckman Synchron System | 389 | 6.5 |  | 4 | GH2-05 | 6.8 | 2008 | CAP | Fresh | Beckman Coulter | 4
Beckman Synchron System | 389 | 5 |  | 4.7 | GH2-06 | 5 | 2008 | CAP | Fresh | Beckman Coulter | 4.7
Beckman UniCel DxC | 172 | 6.57 | -0.01 | 2.7 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Beckman Coulter | 2.69
Beckman UniCel DxC | 172 | 8.49 | 0.1 | 2.8 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Beckman Coulter | 2.9
Beckman UniCel DxC | 172 | 5.61 | -0.04 | 2.6 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Beckman Coulter | 2.56
Beckman UniCel DxC | 191 | 6.5 | 0.01 | 3 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Beckman Coulter | 3.01
Beckman UniCel DxC | 191 | 6.95 | -0.02 | 2.6 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Beckman Coulter | 2.58
Beckman UniCel DxC | 191 | 9.66 | 0.01 | 2.8 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Beckman Coulter | 2.81
Beckman UniCel DxC | 149 | 6.78 | -0.01 | 2.6 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Beckman Coulter | 2.59
Beckman UniCel DxC | 149 | 10.37 | 0.09 | 3.1 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Beckman Coulter | 3.19
Beckman UniCel DxC | 149 | 6.77 | -0.05 | 2.5 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Beckman Coulter | 2.45
Beckman UniCel DxC | 114 | 8.7 | 0.07 | 2.6 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Beckman Coulter | 2.67
Beckman UniCel DxC | 114 | 5.31 | -0.01 | 2.5 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Beckman Coulter | 2.49
Beckman UniCel DxC | 117 | 5.78 | -0.06 | 2.4 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Beckman Coulter | 2.34
Beckman UniCel DxC | 117 | 11.87 | 0.16 | 3.3 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Beckman Coulter | 3.46
Beckman UniCel DxC | 117 | 9.62 | 0.09 | 2.8 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Beckman Coulter | 2.89
Beckman UniCel DxC | 117 | 5.05 | 0.01 | 2 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Beckman Coulter | 2.01
Beckman UniCel DxC | 117 | 7.32 | -0.06 | 2.5 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Beckman Coulter | 2.44
Beckman UniCel DxC | 119-157 | 11.87 | 0.18 | 3.5 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Beckman Coulter | 3.68
Beckman UniCel DxC | 119-157 | 5.89 | -0.04 | 2.4 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Beckman Coulter | 2.36
Beckman UniCel DxC | 119-157 | 5.19 | 0.02 | 2.3 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Beckman Coulter | 2.32
Beckman UniCel DxC | 119-157 | 8.17 | 0.03 | 2.2 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Beckman Coulter | 2.23
Beckman UniCel DxC | 119-157 | 9.41 | 0.11 | 3 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Beckman Coulter | 3.11
Beckman UniCel DxC | 111-142 | 5.31 | -0.01 | 2.6 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Beckman Coulter | 2.59
Beckman UniCel DxC | 111-142 | 9.35 | 0.19 | 2.4 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Beckman Coulter | 2.59
Beckman UniCel DxC | 111-142 | 5.28 | -0.03 | 2.4 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Beckman Coulter | 2.37
Beckman UniCel DxC | 111-142 | 12.27 | 0.24 | 3.1 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Beckman Coulter | 3.34
Beckman UniCel DxC | 111-142 | 5.86 | -0.08 | 2.8 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Beckman Coulter | 2.72
Beckman UniCel DxC | 108 | 5.23 | -0.04 | 1.5 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Beckman Coulter | 1.46
Beckman UniCel DxC | 108 | 10.68 | 0.09 | 2 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Beckman Coulter | 2.09
Beckman UniCel DxC | 108 | 6.07 | -0.13 | 1.6 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Beckman Coulter | 1.47
Beckman UniCel DxC | 108 | 12.40 | 0.17 | 2.2 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Beckman Coulter | 2.37
Beckman UniCel DxC | 108 | 7.50 | -0.01 | 1.8 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Beckman Coulter | 1.79
Beckman UniCel DxC | 101-133 | 9.21 | 0.1 | 2.4 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Beckman Coulter | 2.5
Beckman UniCel DxC | 101-133 | 5.93 | -0.08 | 2.2 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Beckman Coulter | 2.12
Beckman UniCel DxC | 101-133 | 11.83 | 0.12 | 2.6 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Beckman Coulter | 2.72
Beckman UniCel DxC | 101-133 | 5.02 | 0 | 2.4 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Beckman Coulter | 2.4
Beckman UniCel DxC | 101-133 | 7.53 | -0.05 | 1.9 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Beckman Coulter | 1.85
Beckman UniCel DxC | 108-136 | 6.44 | 0.03 | 3.3 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Beckman Coulter | 3.33
Beckman UniCel DxC | 108-136 | 9.68 | 0.15 | 2.8 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Beckman Coulter | 2.95
Beckman UniCel DxC | 108-136 | 5.33 | -0.01 | 3.1 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Beckman Coulter | 3.09
Beckman UniCel DxC | 108-136 | 8.59 | 0.08 | 2.7 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Beckman Coulter | 2.78
Beckman UniCel DxC | 108-136 | 7.27 | 0.02 | 2.3 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Beckman Coulter | 2.32
Beckman UniCel DxC | 111 | 7.24 | 0.09 | 3.3 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Beckman Coulter | 3.39
Beckman UniCel DxC | 111 | 5.27 | 0.08 | 3.7 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Beckman Coulter | 3.78
Beckman UniCel DxC | 111 | 8.50 | 0.08 | 2.9 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Beckman Coulter | 2.98
Beckman UniCel DxC | 111 | 9.95 | 0.16 | 2.3 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Beckman Coulter | 2.46
Beckman UniCel DxC | 111 | 6.12 | 0 | 2.5 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Beckman Coulter | 2.5
Beckman UniCel DxC | 86 | 6.34 | 0.03 | 2.5 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Beckman Coulter | 2.53
Beckman UniCel DxC | 86 | 9.23 | 0.12 | 2 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Beckman Coulter | 2.12
Beckman UniCel DxC | 86 | 5.27 | -0.04 | 2.3 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Beckman Coulter | 2.26
Beckman UniCel DxC | 86 | 8.44 | 0.09 | 2.1 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Beckman Coulter | 2.19
Beckman UniCel DxC | 86 | 7.67 | 0.05 | 2.6 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Beckman Coulter | 2.65
Beckman UniCel DxC | 103 | 7.51 | 0.09 | 2.9 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Beckman Coulter | 2.99
Beckman UniCel DxC | 103 | 5.21 | 0.01 | 3.2 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Beckman Coulter | 3.21
Beckman UniCel DxC | 103 | 8.45 | 0.14 | 3 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Beckman Coulter | 3.14
Beckman UniCel DxC | 103 | 10.63 | 0.3 | 3.2 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Beckman Coulter | 3.5
Beckman UniCel DxC | 103 | 5.88 | 0.01 | 2.4 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Beckman Coulter | 2.41
Beckman UniCel DxC | 118 | 7.33 | 0.08 | 3.1 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Beckman Coulter | 3.18
Beckman UniCel DxC | 118 | 6.84 | 0.03 | 3 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Beckman Coulter | 3.03
Beckman UniCel DxC | 118 | 8.77 | 0.12 | 2.6 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Beckman Coulter | 2.72
Beckman UniCel DxC | 118 | 9.78 | 0.28 | 2.7 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Beckman Coulter | 2.98
Beckman UniCel DxC | 118 | 5.45 | 0 | 2.3 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Beckman Coulter | 2.3
Beckman UniCel DxC Synchron | 233 | 7.01 | -0.1 | 3.5 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Beckman Coulter | 3.4
Beckman UniCel DxC Synchron | 233 | 9.45 | 0.13 | 3.6 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Beckman Coulter | 3.73
Beckman UniCel DxC Synchron | 233 | 6.06 | -0.01 | 4.2 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Beckman Coulter | 4.19
Beckman UniCel DxC Synchron | 187 | 7.97 | -0.08 | 3 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Beckman Coulter | 2.92
Beckman UniCel DxC Synchron | 187 | 5.3 | 0 | 3.1 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Beckman Coulter | 3.1
Beckman UniCel DxC Synchron | 187 | 6.09 | -0.05 | 3.1 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Beckman Coulter | 3.05
Beckman UniCel DxC Synchron | 143 | 7.01 | -0.1 | 3.6 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Beckman Coulter | 3.5
Beckman UniCel DxC Synchron (orig) | 143 | 9.46 | 0.14 | 3.8 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Beckman Coulter | 3.94
Beckman UniCel DxC Synchron (orig) | 143 | 6.05 | -0.02 | 4.2 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Beckman Coulter | 4.18
Beckman UniCel DxC Synchron Systems | 68 | 8.92 | 0.06 | 2.6 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Beckman Coulter | 2.66
Beckman UniCel DxC Synchron Systems | 68 | 6.83 | -0.01 | 2.4 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Beckman Coulter | 2.39
Beckman UniCel DxC Synchron Systems | 68 | 9.35 | -0.04 | 2.5 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Beckman Coulter | 2.46
Beckman UniCel DxC Synchron Systems | 68 | 6.1 | -0.03 | 1.8 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Beckman Coulter | 1.77
Beckman UniCel DxC Synchron Systems | 68 | 7.46 | -0.06 | 2.3 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Beckman Coulter | 2.24
Beckman UniCel DxC Synchron Systems | 93 | 9.35 | -0.02 | 2.1 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Beckman Coulter | 2.08
Beckman UniCel DxC Synchron Systems | 93 | 5.02 | 0.1 | 2.7 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Beckman Coulter | 2.8
Beckman UniCel DxC Synchron Systems | 93 | 6.00 | -0.04 | 2.3 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Beckman Coulter | 2.26
Beckman UniCel DxC Synchron Systems | 93 | 5.27 | -0.01 | 2.4 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Beckman Coulter | 2.39
Beckman UniCel DxC Synchron Systems | 93 | 8.05 | -0.03 | 2.3 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Beckman Coulter | 2.27
Beckman UniCel DxC Synchron Systems | 84 | 5.60 | 0.14 | 2.7 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Beckman Coulter | 2.84
Beckman UniCel DxC Synchron Systems | 84 | 5.70 | 0.04 | 2.5 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Beckman Coulter | 2.54
Beckman UniCel DxC Synchron Systems | 84 | 9.39 | 0.08 | 2.7 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Beckman Coulter | 2.78
Beckman UniCel DxC Synchron Systems | 84 | 5.34 | 0.06 | 2.9 | GH5-04 | 5.28 | 2019 | CAP | Fresh | Beckman Coulter | 2.96
Beckman UniCel DxC Synchron Systems | 84 | 7.32 | -0.09 | 2.8 | GH5-05 | 7.41 | 2019 | CAP | Fresh | Beckman Coulter | 2.71
Beckman UniCel DxC Synchron Systems | 81 | 5.24 | -0.01 | 3 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Beckman Coulter | 2.99
Beckman UniCel DxC Synchron Systems | 81 | 6.32 | -0.09 | 2.6 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Beckman Coulter | 2.51
Beckman UniCel DxC Synchron Systems | 81 | 8.08 | -0.13 | 2.2 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Beckman Coulter | 2.07
Beckman UniCel DxC Synchron Systems | 81 | 5.39 | -0.08 | 2.7 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Beckman Coulter | 2.62
Beckman UniCel DxC Synchron Systems | 81 | 9.39 | -0.09 | 3.1 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Beckman Coulter | 3.01
Beckman UniCel DxC Synchron Systems | 73 | 7.97 | 0 | 3 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Beckman Coulter | 0
Beckman UniCel DxC Synchron Systems | 73 | 5.81 | -0.08 | 2 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Beckman Coulter | -0.08
Beckman UniCel DxC Synchron Systems | 73 | 5.1 | -0.03 | 2.3 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Beckman Coulter | 2.27
Beckman UniCel DxC Synchron Systems | 73 | 7.29 | -0.11 | 2.5 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Beckman Coulter | 2.39
Beckman UniCel DxC Synchron Systems | 73 | 9.11 | -0.06 | 2.7 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Beckman Coulter | 2.64
Beckman UniCel DxC Synchron Systems | 60 | 7.78 | -0.03 | 2.4 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Beckman Coulter | 2.37
Beckman UniCel DxC Synchron Systems | 60 | 5.82 | -0.09 | 2.7 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Beckman Coulter | 2.61
Beckman UniCel DxC Synchron Systems | 60 | 8.7 | 0.01 | 3 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Beckman Coulter | 0.01
Beckman UniCel DxC Synchron Systems | 60 | 5.86 | 0 | 2.5 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Beckman Coulter | 2.5
Beckman UniCel DxC Synchron Systems | 60 | 9.56 | 0 | 2.8 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Beckman Coulter | 2.8
Beckman UniCel DxC Synchron Systems | 68 | 5.5 | 0.02 | 3.5 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Beckman Coulter | 3.52
Beckman UniCel DxC Synchron Systems | 68 | 8.24 | -0.02 | 3.1 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Beckman Coulter | 3.08
Beckman UniCel DxC Synchron Systems | 68 | 5.14 | -0.01 | 2.9 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Beckman Coulter | 2.89
Beckman UniCel DxC Synchron Systems | 68 | 10.12 | 0.07 | 2.7 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Beckman Coulter | 2.77
Beckman UniCel DxC Synchron Systems | 68 | 5.93 | -0.06 | 3.1 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Beckman Coulter | 3.04
Beckman UniCel DxC Synchron Systems | 59 | 5.88 | 0.08 | 3 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Beckman Coulter | 0.08
Beckman UniCel DxC Synchron Systems | 59 | 9.08 | -0.04 | 2.4 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Beckman Coulter | 2.36
Beckman UniCel DxC Synchron Systems | 59 | 5.47 | 0 | 2.3 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Beckman Coulter | 2.3
Beckman UniCel DxC Synchron Systems | 59 | 7.24 | -0.01 | 2.2 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Beckman Coulter | 2.19
Beckman UniCel DxC Synchron Systems | 59 | 9.25 | -0.01 | 2.5 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Beckman Coulter | 2.49
Beckman UniCel DxC Synchron Systems | 65 | 5.28 | 0.07 | 3.1 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Beckman Coulter | 3.17
Beckman UniCel DxC Synchron Systems | 65 | 6.66 | -0.02 | 2.4 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Beckman Coulter | 2.38
Beckman UniCel DxC Synchron Systems | 65 | 8.32 |  | 2.7 | GH5-13 |  | 2021 | CAP | Fresh | Beckman Coulter | 2.7
Beckman UniCel DxC Synchron Systems | 65 | 5.98 | 0 | 2.1 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Beckman Coulter | 2.1
Beckman UniCel DxC Synchron Systems | 65 | 10.3 | 0.18 | 2.6 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Beckman Coulter | 2.78
Beckman UniCel DxC Synchron Systems | 69 | 6.32 | -0.09 | 3 | GH5-06 | 6.41 | 2019 | CAP | Fresh | Beckman Coulter | 2.91
Beckman UniCel DxC Synchron Systems | 69 | 8.58 | -0.02 | 2.5 | GH5-07 | 8.6 | 2019 | CAP | Fresh | Beckman Coulter | 2.48
Beckman UniCel DxC Synchron Systems | 69 | 5.39 | -0.03 | 3.7 | GH5-08 | 5.42 | 2019 | CAP | Fresh | Beckman Coulter | 3.67
Beckman UniCel DxC Synchron Systems | 69 | 7.29 | -0.09 | 3.1 | GH5-09 | 7.38 | 2019 | CAP | Fresh | Beckman Coulter | 3.01
Beckman UniCel DxC Synchron Systems | 69 | 9.81 | 0.06 | 3 | GH5-10 | 9.75 | 2019 | CAP | Fresh | Beckman Coulter | 3.06
Bio-Rad D-10 | 106 | 10.9 |  | 2.1 | GH2-01 | 10.7 | 2006 | CAP | Fresh | Bio-Rad | 2.1
Bio-Rad D-10 | 106 | 5.3 |  | 2.9 | GH2-02 | 5.30 | 2006 | CAP | Fresh | Bio-Rad | 2.9
Bio-Rad D-10 | 106 | 8.5 |  | 2 | GH2-03 | 8.40 | 2006 | CAP | Fresh | Bio-Rad | 2
Bio-Rad D-10 | 130 | 7.2 |  | 2.6 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Bio-Rad | 2.6
Bio-Rad D-10 | 130 | 7.2 |  | 2.6 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Bio-Rad | 2.6
Bio-Rad D-10 | 130 | 12.3 |  | 2.4 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Bio-Rad | 2.4
Bio-Rad D-10 | 130 | 7.2 |  | 2.6 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Bio-Rad | 2.6
Bio-Rad D-10 | 130 | 7.2 |  | 2.6 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Bio-Rad | 2.6
Bio-Rad D-10 | 130 | 12.3 |  | 2.4 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Bio-Rad | 2.4
Bio-Rad D-10 | 169 | 5.5 |  | 2.7 | GH2-01 | 5.4 | 2007 | CAP | Fresh | Bio-Rad | 2.7
Bio-Rad D-10 | 169 | 12.6 |  | 2.3 | GH2-02 | 11.6 | 2007 | CAP | Fresh | Bio-Rad | 2.3
Bio-Rad D-10 | 169 | 8.1 |  | 2.5 | GH2-03 | 7.6 | 2007 | CAP | Fresh | Bio-Rad | 2.5
Bio-Rad D-10 | 167 | 9.6 |  | 2.7 | GH2-04 | 9.2 | 2007 | CAP | Fresh | Bio-Rad | 2.7
Bio-Rad D-10 | 167 | 11.5 |  | 2.6 | GH2-05 | 11.1 | 2007 | CAP | Fresh | Bio-Rad | 2.6
Bio-Rad D-10 | 167 | 6.3 |  | 2.5 | GH2-06 | 6.2 | 2007 | CAP | Fresh | Bio-Rad | 2.5
Bio-Rad D-10 | 177 | 13.2 |  | 2.4 | GH2-01 | 13 | 2008 | CAP | Fresh | Bio-Rad | 2.4
Bio-Rad D-10 | 177 | 5.3 |  | 3.1 | GH2-02 | 5.3 | 2008 | CAP | Fresh | Bio-Rad | 3.1
Bio-Rad D-10 | 177 | 9 |  | 2.3 | GH2-03 | 8.8 | 2008 | CAP | Fresh | Bio-Rad | 2.3
Bio-Rad D-10 | 200 | 12.1 |  | 2.8 | GH2-04 | 11.9 | 2008 | CAP | Fresh | Bio-Rad | 2.8
Bio-Rad D-10 | 200 | 6.9 |  | 2.9 | GH2-05 | 6.8 | 2008 | CAP | Fresh | Bio-Rad | 2.9
Bio-Rad D-10 | 200 | 5 |  | 3.9 | GH2-06 | 5 | 2008 | CAP | Fresh | Bio-Rad | 3.9
Bio-Rad D-10 | 191 | 5.1 |  | 3.2 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Bio-Rad | 3.2
Bio-Rad D-10 | 191 | 8.5 |  | 2.3 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Bio-Rad | 2.3
Bio-Rad D-10 | 191 | 6.1 |  | 3 | GH2-03 | 6 | 2009 | CAP | Fresh | Bio-Rad | 3
Bio-Rad D-10 | 209 | 6.74 | 0.14 | 2.9 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Bio-Rad | 3.04
Bio-Rad D-10 | 209 | 9.54 | 0.04 | 2.8 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Bio-Rad | 2.84
Bio-Rad D-10 | 209 | 7.6 | 0.2 | 3.2 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Bio-Rad | 3.4
Bio-Rad D-10 | 198 | 6.09 | 0.19 | 2.6 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Bio-Rad | 2.79
Bio-Rad D-10 | 198 | 10 | 0.2 | 2.5 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Bio-Rad | 2.7
Bio-Rad D-10 | 198 | 7.56 | 0.16 | 2.4 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Bio-Rad | 2.56
Bio-Rad D-10 | 207 | 5.29 | 0.09 | 2.8 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Bio-Rad | 2.89
Bio-Rad D-10 | 207 | 8.94 | 0.24 | 2.5 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Bio-Rad | 2.74
Bio-Rad D-10 | 207 | 6.45 | 0.15 | 2.7 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Bio-Rad | 2.85
Bio-Rad D-10 | 199 | 8.72 | 0.22 | 2.4 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Bio-Rad | 2.62
Bio-Rad D-10 | 199 | 5.41 | 0.01 | 2.9 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Bio-Rad | 2.91
Bio-Rad D-10 | 199 | 6.44 | 0.04 | 2.8 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Bio-Rad | 2.84
Bio-Rad D-10 | 215 | 6.55 | 0.25 | 2.8 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Bio-Rad | 3.05
Bio-Rad D-10 | 215 | 7.92 | 0.32 | 2.3 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Bio-Rad | 2.62
Bio-Rad D-10 | 215 | 9.46 | 0.26 | 2.7 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Bio-Rad | 2.96
Bio-Rad D-10 | 244 | 5.68 | 0.08 | 2.7 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Bio-Rad | 2.78
Bio-Rad D-10 | 244 | 9.64 | 0.24 | 2.6 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Bio-Rad | 2.84
Bio-Rad D-10 | 244 | 7.55 | 0.35 | 2.7 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Bio-Rad | 3.05
Bio-Rad D-10 | 230 | 5.55 | 0.15 | 2.9 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Bio-Rad | 3.05
Bio-Rad D-10 | 230 | 8.5 | 0.2 | 2.5 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Bio-Rad | 2.7
Bio-Rad D-10 | 230 | 5.8 | 0.15 | 3 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Bio-Rad | 3.15
Bio-Rad D-10 | 210 | 7.16 | 0.05 | 2.7 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Bio-Rad | 2.75
Bio-Rad D-10 | 210 | 9.41 | 0.09 | 2.6 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Bio-Rad | 2.69
Bio-Rad D-10 | 210 | 6.14 | 0.07 | 2.6 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Bio-Rad | 2.67
Bio-Rad D-10 | 205 | 8.18 | 0.13 | 2.7 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Bio-Rad | 2.83
Bio-Rad D-10 | 205 | 5.32 | 0.02 | 2.8 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Bio-Rad | 2.82
Bio-Rad D-10 | 205 | 6.2 | 0.06 | 2.8 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Bio-Rad | 2.86
Bio-Rad D-10 | 217 | 6.6 | 0.02 | 2.8 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Bio-Rad | 2.82
Bio-Rad D-10 | 217 | 8.45 | 0.06 | 2.2 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Bio-Rad | 2.26
Bio-Rad D-10 | 217 | 5.69 | 0.04 | 2.7 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Bio-Rad | 2.74
Bio-Rad D-10 | 209 | 6.58 | 0.09 | 2.3 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Bio-Rad | 2.39
Bio-Rad D-10 | 209 | 7.13 | 0.16 | 2.1 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Bio-Rad | 2.26
Bio-Rad D-10 | 209 | 9.77 | 0.12 | 2.4 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Bio-Rad | 2.52
Bio-Rad D-10 | 198 | 6.92 | 0.13 | 2.7 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Bio-Rad | 2.83
Bio-Rad D-10 | 198 | 10.36 | 0.08 | 2.3 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Bio-Rad | 2.38
Bio-Rad D-10 | 198 | 6.92 | 0.1 | 2.4 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Bio-Rad | 2.5
Bio-Rad D-10 | 136 | 8.72 | 0.09 | 2.3 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Bio-Rad | 2.39
Bio-Rad D-10 | 136 | 5.36 | 0.04 | 2.6 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Bio-Rad | 2.64
Bio-Rad D-10 | 133 | 6 | 0.16 | 2.7 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Bio-Rad | 2.86
Bio-Rad D-10 | 133 | 11.83 | 0.12 | 2.6 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Bio-Rad | 2.72
Bio-Rad D-10 | 133 | 9.67 | 0.14 | 2.7 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Bio-Rad | 2.84
Bio-Rad D-10 | 133 | 5.09 | 0.05 | 3.3 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Bio-Rad | 3.35
Bio-Rad D-10 | 133 | 7.43 | 0.05 | 2.6 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Bio-Rad | 2.65
Bio-Rad D-10 | 137-198 | 11.94 | 0.25 | 2.7 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Bio-Rad | 2.95
Bio-Rad D-10 | 137-198 | 5.99 | 0.06 | 2.5 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Bio-Rad | 2.56
Bio-Rad D-10 | 137-198 | 5.14 | -0.03 | 2.9 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Bio-Rad | 2.87
Bio-Rad D-10 | 137-198 | 8.06 | -0.08 | 2.6 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Bio-Rad | 2.52
Bio-Rad D-10 | 137-198 | 9.4 | 0.1 | 2.7 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Bio-Rad | 2.8
Bio-Rad D-10 | 127-176 | 5.42 | 0.1 | 2.8 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Bio-Rad | 2.9
Bio-Rad D-10 | 127-176 | 9.42 | 0.26 | 2.5 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Bio-Rad | 2.76
Bio-Rad D-10 | 127-176 | 5.41 | 0.1 | 2.9 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Bio-Rad | 3
Bio-Rad D-10 | 127-176 | 12.46 | 0.43 | 2.4 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Bio-Rad | 2.83
Bio-Rad D-10 | 127-176 | 6.11 | 0.17 | 2.4 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Bio-Rad | 2.57
Bio-Rad D-10 | 128 | 5.32 | 0.05 | 2.9 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Bio-Rad | 2.95
Bio-Rad D-10 | 128 | 10.83 | 0.24 | 2.4 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Bio-Rad | 2.64
Bio-Rad D-10 | 128 | 6.33 | 0.13 | 2.5 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Bio-Rad | 2.63
Bio-Rad D-10 | 128 | 12.39 | 0.16 | 2.4 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Bio-Rad | 2.56
Bio-Rad D-10 | 128 | 7.68 | 0.17 | 2.4 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Bio-Rad | 2.57
Bio-Rad D-10 | 125-172 | 9.16 | 0.05 | 2.5 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Bio-Rad | 2.55
Bio-Rad D-10 | 125-172 | 6.05 | 0.04 | 2.4 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Bio-Rad | 2.44
Bio-Rad D-10 | 125-172 | 11.85 | 0.14 | 2.6 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Bio-Rad | 2.74
Bio-Rad D-10 | 125-172 | 5.11 | 0.09 | 2.3 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Bio-Rad | 2.39
Bio-Rad D-10 | 125-172 | 7.65 | 0.07 | 2.2 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Bio-Rad | 2.27
Bio-Rad D-10 | 120-164 | 6.53 | 0.12 | 2.4 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Bio-Rad | 2.52
Bio-Rad D-10 | 120-164 | 9.81 | 0.28 | 2.7 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Bio-Rad | 2.98
Bio-Rad D-10 | 120-164 | 5.36 | 0.02 | 2.6 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Bio-Rad | 2.62
Bio-Rad D-10 | 120-164 | 8.75 | 0.24 | 2.2 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Bio-Rad | 2.44
Bio-Rad D-10 | 120-164 | 7.36 | 0.11 | 2.4 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Bio-Rad | 2.51
Bio-Rad D-10 | 121 | 7.58 | 0.16 | 2.5 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Bio-Rad | 2.66
Bio-Rad D-10 | 121 | 5.45 | 0.25 | 2.3 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Bio-Rad | 2.55
Bio-Rad D-10 | 121 | 8.38 | 0.07 | 1.9 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Bio-Rad | 1.97
Bio-Rad D-10 | 121 | 10.39 | 0.06 | 1.9 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Bio-Rad | 1.96
Bio-Rad D-10 | 121 | 6.01 | 0.14 | 2.3 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Bio-Rad | 2.44
Bio-Rad D-10 | 156 | 7.36 | 0.11 | 2.3 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Bio-Rad | 2.41
Bio-Rad D-10 | 156 | 6.90 | 0.09 | 2.4 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Bio-Rad | 2.49
Bio-Rad D-10 | 156 | 8.71 | 0.06 | 2.1 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Bio-Rad | 2.16
Bio-Rad D-10 | 156 | 9.65 | 0.15 | 2.4 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Bio-Rad | 2.55
Bio-Rad D-10 | 156 | 5.48 | 0.03 | 2.7 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Bio-Rad | 2.73
Bio-Rad D-10 | 156 | 7.25 | 0.1 | 2.4 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Bio-Rad | 2.5
Bio-Rad D-10 | 156 | 5.20 | 0.01 | 2.8 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Bio-Rad | 2.81
Bio-Rad D-10 | 156 | 8.60 | 0.18 | 2.4 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Bio-Rad | 2.58
Bio-Rad D-10 | 156 | 9.94 | 0.15 | 2.4 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Bio-Rad | 2.55
Bio-Rad D-10 | 156 | 6.25 | 0.13 | 2.6 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Bio-Rad | 2.73
Bio-Rad D-10 | 125 | 6.45 | 0.14 | 2.7 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Bio-Rad | 2.84
Bio-Rad D-10 | 125 | 9.15 | 0.04 | 2.6 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Bio-Rad | 2.64
Bio-Rad D-10 | 125 | 5.29 | -0.02 | 2.8 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Bio-Rad | 2.78
Bio-Rad D-10 | 125 | 8.50 | 0.15 | 2.4 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Bio-Rad | 2.55
Bio-Rad D-10 | 125 | 7.73 | 0.11 | 2.5 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Bio-Rad | 2.61
Bio-Rad D-10 | 153 | 9.45 | 0.08 | 3 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Bio-Rad | 3.08
Bio-Rad D-10 | 153 | 4.85 | -0.07 | 3.3 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Bio-Rad | 3.23
Bio-Rad D-10 | 153 | 6.13 | 0.09 | 3 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Bio-Rad | 3.09
Bio-Rad D-10 | 153 | 5.39 | 0.11 | 2.9 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Bio-Rad | 3.01
Bio-Rad D-10 | 153 | 8.28 | 0.2 | 2.6 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Bio-Rad | 2.8
Bio-Rad D-10 | 155 | 5.41 | -0.05 | 3.5 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Bio-Rad | 3.45
Bio-Rad D-10 | 155 | 5.51 | -0.15 | 4.4 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Bio-Rad | 4.25
Bio-Rad D-10 | 155 | 9.49 | 0.18 | 2.3 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Bio-Rad | 2.48
Bio-Rad D-10 | 155 | 5.23 | -0.05 | 3 | GH5-04 | 5.28 | 2019 | CAP | Fresh | Bio-Rad | 2.95
Bio-Rad D-10 | 155 | 7.54 | 0.13 | 2.2 | GH5-05 | 7.41 | 2019 | CAP | Fresh | Bio-Rad | 2.33
Bio-Rad D-10 | 123 | 6.55 | 0.14 | 2.4 | GH5-06 | 6.41 | 2019 | CAP | Fresh | Bio-Rad | 2.54
Bio-Rad D-10 | 123 | 8.80 | 0.2 | 2.3 | GH5-07 | 8.6 | 2019 | CAP | Fresh | Bio-Rad | 2.5
Bio-Rad D-10 | 123 | 5.45 | 0.03 | 2.7 | GH5-08 | 5.42 | 2019 | CAP | Fresh | Bio-Rad | 2.73
Bio-Rad D-10 | 123 | 7.47 | 0.09 | 2.4 | GH5-09 | 7.38 | 2019 | CAP | Fresh | Bio-Rad | 2.49
Bio-Rad D-10 | 123 | 9.89 | 0.14 | 2.4 | GH5-10 | 9.75 | 2019 | CAP | Fresh | Bio-Rad | 2.54
Bio-Rad D-10 | 135 | 5.17 | -0.08 | 2.4 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Bio-Rad | 2.32
Bio-Rad D-10 | 135 | 6.51 | 0.1 | 2.1 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Bio-Rad | 2.2
Bio-Rad D-10 | 135 | 8.16 | -0.05 | 2.1 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Bio-Rad | 2.05
Bio-Rad D-10 | 135 | 5.44 | -0.03 | 2.2 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Bio-Rad | 2.17
Bio-Rad D-10 | 135 | 9.58 | 0.1 | 1.9 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Bio-Rad | 2
Bio-Rad D-10 | 117 | 7.89 | -0.08 | 2.8 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Bio-Rad | 2.72
Bio-Rad D-10 | 117 | 5.9 | 0.01 | 2.3 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Bio-Rad | 2.31
Bio-Rad D-10 | 117 | 5.07 | -0.06 | 2.3 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Bio-Rad | 2.24
Bio-Rad D-10 | 117 | 7.33 | -0.07 | 2.8 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Bio-Rad | 2.73
Bio-Rad D-10 | 117 | 9.23 | 0.06 | 2.3 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Bio-Rad | 2.36
Bio-Rad D-10 | 103 | 7.8 | -0.01 | 1.8 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Bio-Rad | 1.79
Bio-Rad D-10 | 103 | 5.92 | 0.01 | 1.8 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Bio-Rad | 1.81
Bio-Rad D-10 | 103 | 8.78 | 0.09 | 1.7 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Bio-Rad | 1.79
Bio-Rad D-10 | 103 | 6.03 | 0.17 | 1.9 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Bio-Rad | 2.07
Bio-Rad D-10 | 103 | 9.59 | 0.03 | 1.9 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Bio-Rad | 1.93
Bio-Rad D-10 | 124 | 5.58 | 0.1 | 2.6 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Bio-Rad | 2.7
Bio-Rad D-10 | 124 | 8.29 | 0.03 | 2.5 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Bio-Rad | 2.53
Bio-Rad D-10 | 124 | 5.09 | -0.06 | 2.6 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Bio-Rad | 2.54
Bio-Rad D-10 | 124 | 10.14 | 0.09 | 2.5 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Bio-Rad | 2.59
Bio-Rad D-10 | 124 | 6.01 | 0.02 | 2.2 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Bio-Rad | 2.22
Bio-Rad D-10 | 116 | 9.12 | 0.26 | 2.3 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Bio-Rad | 2.56
Bio-Rad D-10 | 116 | 6.98 | 0.14 | 2.5 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Bio-Rad | 2.64
Bio-Rad D-10 | 116 | 9.53 | 0.14 | 1.9 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Bio-Rad | 2.04
Bio-Rad D-10 | 116 | 6.19 | 0.06 | 2.2 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Bio-Rad | 2.26
Bio-Rad D-10 | 116 | 7.55 | 0.03 | 1.8 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Bio-Rad | 1.83
Bio-Rad D-10 | 100 | 5.99 | 0.19 | 2.2 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Bio-Rad | 2.39
Bio-Rad D-10 | 100 | 9.25 | 0.13 | 2.1 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Bio-Rad | 2.23
Bio-Rad D-10 | 100 | 5.58 | 0.11 | 2.1 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Bio-Rad | 2.21
Bio-Rad D-10 | 100 | 7.24 | -0.01 | 2 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Bio-Rad | -0.01
Bio-Rad D-10 | 100 | 9.49 | 0.23 | 2 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Bio-Rad | 0.23
Bio-Rad D-10 | 112 | 5.3 | 0.09 | 2.2 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Bio-Rad | 2.29
Bio-Rad D-10 | 112 | 6.82 | 0.14 | 1.9 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Bio-Rad | 2.04
Bio-Rad D-10 | 112 | 8.49 |  | 2.4 | GH5-13 |  | 2021 | CAP | Fresh | Bio-Rad | 2.4
Bio-Rad D-10 | 112 | 6.13 | 0.15 | 1.9 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Bio-Rad | 2.05
Bio-Rad D-10 | 112 | 10.25 | 0.13 | 1.9 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Bio-Rad | 2.03
Bio-Rad D-100 | 17-24 | 6.38 | -0.03 | 1.6 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Bio-Rad | 1.57
Bio-Rad D-100 | 17-24 | 9.41 | -0.12 | 2.1 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Bio-Rad | 1.98
Bio-Rad D-100 | 17-24 | 5.25 | -0.09 | 2.1 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Bio-Rad | 2.01
Bio-Rad D-100 | 17-24 | 8.47 | -0.04 | 2 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Bio-Rad | 1.96
Bio-Rad D-100 | 17-24 | 7.18 | -0.07 | 1.7 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Bio-Rad | 1.63
Bio-Rad D-100 | 23 | 7.33 | -0.09 | 2.4 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Bio-Rad | 2.31
Bio-Rad D-100 | 23 | 5.20 | 0 | 2 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Bio-Rad | 2
Bio-Rad D-100 | 23 | 8.13 | -0.18 | 1.7 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Bio-Rad | 1.52
Bio-Rad D-100 | 23 | 10.06 | -0.27 | 1.9 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Bio-Rad | 1.63
Bio-Rad D-100 | 23 | 5.83 | -0.04 | 1.7 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Bio-Rad | 1.66
Bio-Rad D-100 | 41 | 7.19 | -0.06 | 1.7 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Bio-Rad | 1.64
Bio-Rad D-100 | 41 | 6.76 | -0.05 | 1.7 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Bio-Rad | 1.65
Bio-Rad D-100 | 41 | 8.49 | -0.16 | 1.3 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Bio-Rad | 1.14
Bio-Rad D-100 | 41 | 9.32 | -0.18 | 2 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Bio-Rad | 1.82
Bio-Rad D-100 | 41 | 5.42 | -0.03 | 1.7 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Bio-Rad | 1.67
Bio-Rad D-100 | 60 | 7.02 | -0.13 | 1.4 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Bio-Rad | 1.27
Bio-Rad D-100 | 60 | 5.15 | -0.04 | 2 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Bio-Rad | 1.96
Bio-Rad D-100 | 60 | 8.18 | -0.24 | 1.8 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Bio-Rad | 1.56
Bio-Rad D-100 | 60 | 9.52 | -0.27 | 1.6 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Bio-Rad | 1.33
Bio-Rad D-100 | 60 | 6.10 | -0.02 | 1.5 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Bio-Rad | 1.48
Bio-Rad D-100 | 59 | 6.23 | -0.08 | 1.9 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Bio-Rad | 1.82
Bio-Rad D-100 | 59 | 8.85 | -0.26 | 1.5 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Bio-Rad | 1.24
Bio-Rad D-100 | 59 | 5.18 | -0.13 | 2.1 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Bio-Rad | 1.97
Bio-Rad D-100 | 59 | 8.13 | -0.22 | 1.6 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Bio-Rad | 1.38
Bio-Rad D-100 | 59 | 7.48 | -0.14 | 1.7 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Bio-Rad | 1.56
Bio-Rad D-100 | 85 | 9.14 | -0.23 | 1.8 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Bio-Rad | 1.57
Bio-Rad D-100 | 85 | 4.86 | -0.06 | 2.6 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Bio-Rad | 2.54
Bio-Rad D-100 | 85 | 6.00 | -0.04 | 2.2 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Bio-Rad | 2.16
Bio-Rad D-100 | 85 | 5.28 | 0 | 2.4 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Bio-Rad | 2.4
Bio-Rad D-100 | 85 | 8.00 | -0.08 | 2 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Bio-Rad | 1.92
Bio-Rad D-100 | 102 | 5.33 | -0.13 | 2.2 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Bio-Rad | 2.07
Bio-Rad D-100 | 102 | 5.58 | -0.08 | 2.4 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Bio-Rad | 2.32
Bio-Rad D-100 | 102 | 9.11 | -0.2 | 1.7 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Bio-Rad | 1.5
Bio-Rad D-100 | 102 | 5.11 | -0.17 | 2.4 | GH5-04 | 5.28 | 2019 | CAP | Fresh | Bio-Rad | 2.23
Bio-Rad D-100 | 102 | 7.23 | -0.18 | 2 | GH5-05 | 7.41 | 2019 | CAP | Fresh | Bio-Rad | -0.18
Bio-Rad D-100 | 94 | 6.30 | -0.11 | 1.9 | GH5-06 | 6.41 | 2019 | CAP | Fresh | Bio-Rad | 1.79
Bio-Rad D-100 | 94 | 8.45 | -0.15 | 1.5 | GH5-07 | 8.6 | 2019 | CAP | Fresh | Bio-Rad | 1.35
Bio-Rad D-100 | 94 | 5.25 | -0.17 | 2.2 | GH5-08 | 5.42 | 2019 | CAP | Fresh | Bio-Rad | 2.03
Bio-Rad D-100 | 94 | 7.21 | -0.17 | 1.6 | GH5-09 | 7.38 | 2019 | CAP | Fresh | Bio-Rad | 1.43
Bio-Rad D-100 | 94 | 9.55 | -0.2 | 1.2 | GH5-10 | 9.75 | 2019 | CAP | Fresh | Bio-Rad | 1
Bio-Rad D-100 | 119 | 4.99 | -0.26 | 2.3 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Bio-Rad | 2.04
Bio-Rad D-100 | 119 | 6.27 | -0.14 | 2.1 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Bio-Rad | 1.96
Bio-Rad D-100 | 119 | 7.95 | -0.26 | 2 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Bio-Rad | 1.74
Bio-Rad D-100 | 119 | 5.26 | -0.21 | 2.2 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Bio-Rad | 1.99
Bio-Rad D-100 | 119 | 9.27 | -0.21 | 2.1 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Bio-Rad | 1.89
Bio-Rad D-100 | 116 | 7.86 | -0.11 | 2.2 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Bio-Rad | 2.09
Bio-Rad D-100 | 116 | 5.75 | -0.14 | 2.2 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Bio-Rad | 2.06
Bio-Rad D-100 | 116 | 4.92 | -0.21 | 2.5 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Bio-Rad | 2.29
Bio-Rad D-100 | 116 | 7.22 | -0.18 | 2.2 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Bio-Rad | 2.02
Bio-Rad D-100 | 116 | 9.02 | -0.15 | 2.1 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Bio-Rad | 1.95
Bio-Rad D-100 | 109 | 7.72 | -0.09 | 1.8 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Bio-Rad | 1.71
Bio-Rad D-100 | 109 | 5.83 | -0.08 | 1.9 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Bio-Rad | 1.82
Bio-Rad D-100 | 109 | 8.64 | -0.05 | 1.9 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Bio-Rad | 1.85
Bio-Rad D-100 | 109 | 5.91 | 0.05 | 2.1 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Bio-Rad | 2.15
Bio-Rad D-100 | 109 | 9.45 | -0.11 | 1.7 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Bio-Rad | 1.59
Bio-Rad D-100 | 131 | 5.46 | -0.02 | 1.8 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Bio-Rad | 1.78
Bio-Rad D-100 | 131 | 8.22 | -0.04 | 1.7 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Bio-Rad | 1.66
Bio-Rad D-100 | 131 | 5.03 | -0.12 | 2.1 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Bio-Rad | 1.98
Bio-Rad D-100 | 131 | 10.02 | -0.03 | 1.7 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Bio-Rad | 1.67
Bio-Rad D-100 | 131 | 5.92 | -0.07 | 1.9 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Bio-Rad | 1.83
Bio-Rad D-100 | 131 | 8.93 | 0.07 | 1.7 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Bio-Rad | 1.77
Bio-Rad D-100 | 131 | 6.86 | 0.02 | 1.6 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Bio-Rad | 1.62
Bio-Rad D-100 | 131 | 9.35 | -0.04 | 1.4 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Bio-Rad | 1.36
Bio-Rad D-100 | 131 | 6.02 | -0.11 | 1.6 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Bio-Rad | 1.49
Bio-Rad D-100 | 131 | 7.42 | -0.1 | 1.7 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Bio-Rad | 1.6
Bio-Rad D-100 | 112 | 5.82 | 0.02 | 1.7 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Bio-Rad | 1.72
Bio-Rad D-100 | 112 | 9.06 | -0.06 | 1.5 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Bio-Rad | 1.44
Bio-Rad D-100 | 112 | 5.42 | -0.05 | 2.1 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Bio-Rad | 2.05
Bio-Rad D-100 | 112 | 7.13 | -0.12 | 1.5 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Bio-Rad | 1.38
Bio-Rad D-100 | 112 | 9.27 | 0.01 | 1.4 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Bio-Rad | 1.41
Bio-Rad D-100 | 133 | 5.12 | -0.09 | 2.1 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Bio-Rad | 2.01
Bio-Rad D-100 | 133 | 6.66 | -0.02 | 1.7 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Bio-Rad | 1.68
Bio-Rad D-100 | 133 | 8.33 |  | 2.1 | GH5-13 |  | 2021 | CAP | Fresh | Bio-Rad | 2.1
Bio-Rad D-100 | 133 | 5.93 | -0.05 | 1.5 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Bio-Rad | 1.45
Bio-Rad D-100 | 133 | 9.99 | -0.13 | 1.7 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Bio-Rad | 1.57
Bio-Rad Diastat | 15 | 10 |  | 6.1 | GH2-01 | 10.7 | 2006 | CAP | Fresh | Bio-Rad | 6.1
Bio-Rad Diastat | 15 | 5 |  | 5.9 | GH2-02 | 5.30 | 2006 | CAP | Fresh | Bio-Rad | 5.9
Bio-Rad Diastat | 15 | 8 |  | 5.5 | GH2-03 | 8.40 | 2006 | CAP | Fresh | Bio-Rad | 5.5
Bio-Rad Diastat | 13 | 6.9 |  | 5.1 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Bio-Rad | 5.1
Bio-Rad Diastat | 13 | 6.8 |  | 5.5 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Bio-Rad | 5.5
Bio-Rad Diastat | 13 | 11.8 |  | 3.6 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Bio-Rad | 3.6
Bio-Rad Diastat | 13 | 6.9 |  | 5.1 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Bio-Rad | 5.1
Bio-Rad Diastat | 13 | 6.8 |  | 5.5 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Bio-Rad | 5.5
Bio-Rad Diastat | 13 | 11.8 |  | 3.6 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Bio-Rad | 3.6
Bio-Rad Diastat | 11 | 5.2 |  | 3.5 | GH2-01 | 5.4 | 2007 | CAP | Fresh | Bio-Rad | 3.5
Bio-Rad Diastat | 11 | 11.8 |  | 4.4 | GH2-02 | 11.6 | 2007 | CAP | Fresh | Bio-Rad | 4.4
Bio-Rad Diastat | 11 | 7.3 |  | 3.7 | GH2-03 | 7.6 | 2007 | CAP | Fresh | Bio-Rad | 3.7
Bio-Rad Diastat | 7 | 8.9 |  |  | GH2-04 | 9.2 | 2007 | CAP | Fresh | Bio-Rad
Bio-Rad Diastat | 7 | 10.9 |  |  | GH2-05 | 11.1 | 2007 | CAP | Fresh | Bio-Rad
Bio-Rad Diastat | 7 | 5.7 |  |  | GH2-06 | 6.2 | 2007 | CAP | Fresh | Bio-Rad
Bio-Rad in2it | 10 | 6.1 | 0.2 | 3.6 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Bio-Rad | 3.8
Bio-Rad in2it | 10 | 10.28 | 0.48 | 5.6 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Bio-Rad | 6.08
Bio-Rad in2it | 10 | 7.41 | 0.01 | 5.7 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Bio-Rad | 5.71
Bio-Rad in2it | 10 | 5.41 | -0.19 | 7 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Bio-Rad | 6.81
Bio-Rad in2it | 10 | 9.01 | -0.39 | 3 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Bio-Rad | 2.61
Bio-Rad in2it | 10 | 7.24 | 0.04 | 5.4 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Bio-Rad | 5.44
Bio-Rad in2it | 11 | 5.3 | -0.1 | 5.2 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Bio-Rad | 5.1
Bio-Rad in2it | 11 | 8.13 | -0.17 | 3.9 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Bio-Rad | 3.73
Bio-Rad in2it | 11 | 5.55 | -0.1 | 7.1 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Bio-Rad | 7.0
Bio-Rad Variant | 10 | 13.6 |  | 2.7 | GH2-04 | 10.7 | 2006 | CAP | Fresh | Bio-Rad | 2.7
Bio-Rad Variant | 10 | 5.6 |  | 8 | GH2-05 | 5.30 | 2006 | CAP | Fresh | Bio-Rad | 8.0
Bio-Rad Variant | 10 | 10.2 |  | 5.9 | GH2-06 | 8.40 | 2006 | CAP | Fresh | Bio-Rad | 5.9
Bio-Rad Variant | 9 | 7.6 |  |  | GH2-04 | 10.7 | 2006 | CAP | Fresh | Bio-Rad
Bio-Rad Variant | 9 | 7.6 |  |  | GH2-05 | 5.30 | 2006 | CAP | Fresh | Bio-Rad
Bio-Rad Variant | 9 | 14.8 |  |  | GH2-06 | 8.40 | 2006 | CAP | Fresh | Bio-Rad
Bio-Rad Variant | 10 | 5.4 |  | 6.5 | GH2-01 | 5.4 | 2007 | CAP | Fresh | Bio-Rad | 6.5
Bio-Rad Variant | 10 | 14.1 |  | 4.2 | GH2-02 | 11.6 | 2007 | CAP | Fresh | Bio-Rad | 4.2
Bio-Rad Variant | 10 | 8.8 |  | 3.9 | GH2-03 | 7.6 | 2007 | CAP | Fresh | Bio-Rad | 3.9
Bio-Rad Variant A1c | 20 | 10.5 |  | 1.7 | GH2-01 | 10.7 | 2006 | CAP | Fresh | Bio-Rad | 1.7
Bio-Rad Variant A1c | 20 | 5.2 |  | 1.3 | GH2-02 | 5.30 | 2006 | CAP | Fresh | Bio-Rad | 1.3
Bio-Rad Variant A1c | 20 | 8.4 |  | 1.9 | GH2-03 | 8.40 | 2006 | CAP | Fresh | Bio-Rad | 1.9
Bio-Rad Variant A1c | 23 | 6.9 |  | 2.9 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Bio-Rad | 2.9
Bio-Rad Variant A1c | 23 | 6.9 |  | 3.1 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Bio-Rad | 3.1
Bio-Rad Variant A1c | 23 | 11.6 |  | 3.6 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Bio-Rad | 3.6
Bio-Rad Variant A1c | 23 | 6.9 |  | 2.9 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Bio-Rad | 2.9
Bio-Rad Variant A1c | 23 | 6.9 |  | 3.1 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Bio-Rad | 3.1
Bio-Rad Variant A1c | 23 | 11.6 |  | 3.6 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Bio-Rad | 3.6
Bio-Rad Variant A1c | 17 | 5.2 |  | 2.6 | GH2-01 | 5.4 | 2007 | CAP | Fresh | Bio-Rad | 2.6
Bio-Rad Variant A1c | 17 | 11.5 |  | 2.2 | GH2-02 | 11.6 | 2007 | CAP | Fresh | Bio-Rad | 2.2
Bio-Rad Variant A1c | 17 | 7.5 |  | 3.7 | GH2-03 | 7.6 | 2007 | CAP | Fresh | Bio-Rad | 3.7
Bio-Rad Variant A1c | 15 | 9.1 |  | 3.2 | GH2-04 | 9.2 | 2007 | CAP | Fresh | Bio-Rad | 3.2
Bio-Rad Variant A1c | 15 | 11 |  | 3 | GH2-05 | 11.1 | 2007 | CAP | Fresh | Bio-Rad | 3
Bio-Rad Variant A1c | 15 | 6.2 |  | 2.3 | GH2-06 | 6.2 | 2007 | CAP | Fresh | Bio-Rad | 2.3
Bio-Rad Variant A1c | 10 | 12.7 |  | 2.5 | GH2-01 | 13 | 2008 | CAP | Fresh | Bio-Rad | 2.5
Bio-Rad Variant A1c | 10 | 5.1 |  | 4.1 | GH2-02 | 5.3 | 2008 | CAP | Fresh | Bio-Rad | 4.1
Bio-Rad Variant A1c | 10 | 8.7 |  | 1.3 | GH2-03 | 8.8 | 2008 | CAP | Fresh | Bio-Rad | 1.3
Bio-Rad Variant A1c | 13 | 12 |  | 3 | GH2-04 | 11.9 | 2008 | CAP | Fresh | Bio-Rad | 3
Bio-Rad Variant A1c | 13 | 6.8 |  | 1.6 | GH2-05 | 6.8 | 2008 | CAP | Fresh | Bio-Rad | 1.6
Bio-Rad Variant A1c | 13 | 4.9 |  | 2.9 | GH2-06 | 5 | 2008 | CAP | Fresh | Bio-Rad | 2.9
Bio-Rad Variant II | 139 | 6.81 | 0.21 | 2.8 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Bio-Rad | 3.01
Bio-Rad Variant II | 139 | 9.75 | 0.25 | 2.7 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Bio-Rad | 2.95
Bio-Rad Variant II | 139 | 7.67 | 0.27 | 3.4 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Bio-Rad | 3.67
Bio-Rad Variant II | 118 | 6.14 | 0.24 | 3.1 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Bio-Rad | 3.34
Bio-Rad Variant II | 118 | 10.07 | 0.27 | 2.3 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Bio-Rad | 2.57
Bio-Rad Variant II | 118 | 7.64 | 0.24 | 2.7 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Bio-Rad | 2.94
Bio-Rad Variant II | 110 | 5.32 | 0.12 | 2.5 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Bio-Rad | 2.62
Bio-Rad Variant II | 110 | 8.87 | 0.17 | 2.5 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Bio-Rad | 2.67
Bio-Rad Variant II | 110 | 6.4 | 0.1 | 2.2 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Bio-Rad | 2.3
Bio-Rad Variant II | 107 | 8.65 | 0.15 | 2.2 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Bio-Rad | 2.35
Bio-Rad Variant II | 107 | 5.47 | 0.07 | 2.8 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Bio-Rad | 2.87
Bio-Rad Variant II | 107 | 6.48 | 0.08 | 2.2 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Bio-Rad | 2.28
Bio-Rad Variant II | 11 | 6.18 | -0.12 | 3.8 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Bio-Rad | 3.68
Bio-Rad Variant II | 11 | 7.55 | -0.05 | 2.6 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Bio-Rad | 2.55
Bio-Rad Variant II | 11 | 9.11 | -0.09 | 4 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Bio-Rad | 3.91
Bio-Rad Variant II | 111 | 5.6 | 0 | 2.8 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Bio-Rad | 2.8
Bio-Rad Variant II | 111 | 9.61 | 0.21 | 2.5 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Bio-Rad | 2.71
Bio-Rad Variant II | 111 | 7.46 | 0.26 | 2.5 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Bio-Rad | 2.76
Bio-Rad Variant II | 103 | 5.32 | -0.08 | 2.7 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Bio-Rad | 2.62
Bio-Rad Variant II | 103 | 8.36 | 0.06 | 2.3 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Bio-Rad | 2.36
Bio-Rad Variant II | 103 | 5.68 | 0.03 | 3.7 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Bio-Rad | 3.73
Bio-Rad Variant II | 97 | 7.04 | -0.07 | 2.1 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Bio-Rad | 2.03
Bio-Rad Variant II | 97 | 9.37 | 0.05 | 2.2 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Bio-Rad | 2.25
Bio-Rad Variant II | 97 | 5.97 | -0.1 | 2.2 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Bio-Rad | 2.1
Bio-Rad Variant II | 88 | 8.06 | 0.01 | 2.4 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Bio-Rad | 2.41
Bio-Rad Variant II | 88 | 5.24 | -0.06 | 2.4 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Bio-Rad | 2.34
Bio-Rad Variant II | 88 | 6.06 | -0.08 | 2.4 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Bio-Rad | 2.32
Bio-Rad Variant II | 27 | 9.41 | 0.04 | 2.3 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Bio-Rad | 2.34
Bio-Rad Variant II | 27 | 4.83 | -0.09 | 3.2 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Bio-Rad | 3.11
Bio-Rad Variant II | 27 | 6.10 | 0.06 | 2.5 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Bio-Rad | 2.56
Bio-Rad Variant II | 27 | 5.33 | 0.05 | 2.5 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Bio-Rad | 2.55
Bio-Rad Variant II | 27 | 8.18 | 0.1 | 2.1 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Bio-Rad | 2.2
Bio-Rad Variant II | 22 | 5.35 | -0.11 | 2.6 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Bio-Rad | 2.49
Bio-Rad Variant II | 22 | 5.69 | 0.03 | 3.4 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Bio-Rad | 3.43
Bio-Rad Variant II | 22 | 9.47 | 0.16 | 2 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Bio-Rad | 2.16
Bio-Rad Variant II | 22 | 5.15 | -0.13 | 2 | GH5-04 | 5.28 | 2019 | CAP | Fresh | Bio-Rad | 1.87
Bio-Rad Variant II | 22 | 7.45 | 0.04 | 1.9 | GH5-05 | 7.41 | 2019 | CAP | Fresh | Bio-Rad | 1.94
Bio-Rad Variant II | 18 | 6.47 | 0.06 | 1.2 | GH5-06 | 6.41 | 2019 | CAP | Fresh | Bio-Rad | 1.26
Bio-Rad Variant II | 18 | 8.79 | 0.19 | 1 | GH5-07 | 8.6 | 2019 | CAP | Fresh | Bio-Rad | 1.19
Bio-Rad Variant II | 18 | 5.31 | -0.11 | 2 | GH5-08 | 5.42 | 2019 | CAP | Fresh | Bio-Rad | 1.89
Bio-Rad Variant II | 18 | 7.44 | 0.06 | 1.2 | GH5-09 | 7.38 | 2019 | CAP | Fresh | Bio-Rad | 1.26
Bio-Rad Variant II | 18 | 9.91 | 0.16 | 1.2 | GH5-10 | 9.75 | 2019 | CAP | Fresh | Bio-Rad | 1.36
Bio-Rad Variant II | 21 | 5.16 | -0.09 | 3.4 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Bio-Rad | 3.31
Bio-Rad Variant II | 21 | 6.49 | 0.08 | 2.7 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Bio-Rad | 2.78
Bio-Rad Variant II | 21 | 8.34 | 0.13 | 1.9 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Bio-Rad | 2.03
Bio-Rad Variant II | 21 | 5.40 | -0.07 | 2.9 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Bio-Rad | 2.83
Bio-Rad Variant II | 21 | 9.64 | 0.16 | 2.7 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Bio-Rad | 2.86
Bio-Rad Variant II | 16 | 7.93 | -0.04 | 2.7 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Bio-Rad | 2.66
Bio-Rad Variant II | 16 | 5.84 | -0.05 | 2.9 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Bio-Rad | 2.85
Bio-Rad Variant II | 16 | 4.98 | -0.15 | 2.6 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Bio-Rad | 2.45
Bio-Rad Variant II | 16 | 7.35 | -0.05 | 4.2 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Bio-Rad | 4.15
Bio-Rad Variant II | 16 | 9.29 | 0.12 | 2.4 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Bio-Rad | 2.52
Bio-Rad Variant II | 17 | 7.91 | 0.1 | 2.9 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Bio-Rad | 3
Bio-Rad Variant II | 17 | 5.93 | 0.02 | 2.6 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Bio-Rad | 2.62
Bio-Rad Variant II | 17 | 8.89 | 0.2 | 2.8 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Bio-Rad | 3
Bio-Rad Variant II | 17 | 6.1 | 0.24 | 3.4 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Bio-Rad | 3.64
Bio-Rad Variant II | 17 | 9.67 | 0.11 | 2.6 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Bio-Rad | 2.71
Bio-Rad Variant II | 18 | 5.57 | 0.09 | 1.9 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Bio-Rad | 1.99
Bio-Rad Variant II | 18 | 8.37 | 0.11 | 2.3 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Bio-Rad | 2.41
Bio-Rad Variant II | 18 | 5.15 | 0 | 3.3 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Bio-Rad | 3.3
Bio-Rad Variant II | 18 | 10.31 | 0.26 | 2.2 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Bio-Rad | 2.46
Bio-Rad Variant II | 18 | 6.13 | 0.14 | 2.1 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Bio-Rad | 2.24
Bio-Rad Variant II | 17 | 9.11 | 0.25 | 1.6 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Bio-Rad | 1.85
Bio-Rad Variant II | 17 | 6.96 | 0.12 | 2.3 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Bio-Rad | 2.42
Bio-Rad Variant II | 17 | 9.55 | 0.16 | 2.1 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Bio-Rad | 2.26
Bio-Rad Variant II | 17 | 6.18 | 0.05 | 1.3 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Bio-Rad | 1.35
Bio-Rad Variant II | 17 | 7.59 | 0.07 | 2.1 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Bio-Rad | 2.17
Bio-Rad Variant II | 18 | 5.91 | 0.11 | 2.6 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Bio-Rad | 2.71
Bio-Rad Variant II | 18 | 9.28 | 0.16 | 1.5 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Bio-Rad | 1.66
Bio-Rad Variant II | 18 | 5.48 | 0.01 | 3.1 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Bio-Rad | 3.11
Bio-Rad Variant II | 18 | 7.23 | -0.02 | 2.6 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Bio-Rad | 2.58
Bio-Rad Variant II | 18 | 9.46 | 0.2 | 1.9 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Bio-Rad | 2.1
Bio-Rad Variant II | 15 | 5.31 | 0.1 | 5 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Bio-Rad | 0.1
Bio-Rad Variant II | 15 | 6.81 | 0.13 | 2.7 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Bio-Rad | 2.83
Bio-Rad Variant II | 15 | 8.51 |  | 2.2 | GH5-13 |  | 2021 | CAP | Fresh | Bio-Rad | 2.2
Bio-Rad Variant II | 15 | 6.11 | 0.13 | 3.5 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Bio-Rad | 3.63
Bio-Rad Variant II | 15 | 10.24 | 0.12 | 2.2 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Bio-Rad | 2.32
Bio-Rad Variant II A1c | 254 | 11.2 |  | 2.8 | GH2-01 | 10.7 | 2006 | CAP | Fresh | Bio-Rad | 2.8
Bio-Rad Variant II A1c | 254 | 5.3 |  | 2.6 | GH2-02 | 5.30 | 2006 | CAP | Fresh | Bio-Rad | 2.6
Bio-Rad Variant II A1c | 254 | 8.7 |  | 2.9 | GH2-03 | 8.40 | 2006 | CAP | Fresh | Bio-Rad | 2.9
Bio-Rad Variant II A1c | 239 | 7.1 |  | 3.2 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Bio-Rad | 3.2
Bio-Rad Variant II A1c | 239 | 7.1 |  | 3 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Bio-Rad | 3
Bio-Rad Variant II A1c | 239 | 12.2 |  | 3 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Bio-Rad | 3
Bio-Rad Variant II A1c | 239 | 7.1 |  | 3.2 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Bio-Rad | 3.2
Bio-Rad Variant II A1c | 239 | 7.1 |  | 3 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Bio-Rad | 3
Bio-Rad Variant II A1c | 239 | 12.2 |  | 3 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Bio-Rad | 3
Bio-Rad Variant II A1c | 253 | 5.4 |  | 3.4 | GH2-01 | 5.4 | 2007 | CAP | Fresh | Bio-Rad | 3.4
Bio-Rad Variant II A1c | 253 | 12.3 |  | 2.5 | GH2-02 | 11.6 | 2007 | CAP | Fresh | Bio-Rad | 2.5
Bio-Rad Variant II A1c | 253 | 7.9 |  | 2.6 | GH2-03 | 7.6 | 2007 | CAP | Fresh | Bio-Rad | 2.6
Bio-Rad Variant II A1c | 200 | 9.6 |  | 2.3 | GH2-04 | 9.2 | 2007 | CAP | Fresh | Bio-Rad | 2.3
Bio-Rad Variant II A1c | 200 | 11.5 |  | 2.1 | GH2-05 | 11.1 | 2007 | CAP | Fresh | Bio-Rad | 2.1
Bio-Rad Variant II A1c | 200 | 6.3 |  | 2.8 | GH2-06 | 6.2 | 2007 | CAP | Fresh | Bio-Rad | 2.8
Bio-Rad Variant II A1c | 175 | 13.4 |  | 2.5 | GH2-01 | 13 | 2008 | CAP | Fresh | Bio-Rad | 2.5
Bio-Rad Variant II A1c | 175 | 5.3 |  | 3.5 | GH2-02 | 5.3 | 2008 | CAP | Fresh | Bio-Rad | 3.5
Bio-Rad Variant II A1c | 175 | 9.2 |  | 2.4 | GH2-03 | 8.8 | 2008 | CAP | Fresh | Bio-Rad | 2.4
Bio-Rad Variant II A1c | 167 | 12.4 |  | 2.6 | GH2-04 | 11.9 | 2008 | CAP | Fresh | Bio-Rad | 2.6
Bio-Rad Variant II A1c | 167 | 7 |  | 3 | GH2-05 | 6.8 | 2008 | CAP | Fresh | Bio-Rad | 3
Bio-Rad Variant II A1c | 167 | 5 |  | 4 | GH2-06 | 5 | 2008 | CAP | Fresh | Bio-Rad | 4
Bio-Rad Variant II A1c | 146 | 5.1 |  | 3.6 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Bio-Rad | 3.6
Bio-Rad Variant II A1c | 146 | 8.7 |  | 2.8 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Bio-Rad | 2.8
Bio-Rad Variant II A1c | 146 | 6.2 |  | 2.8 | GH2-03 | 6 | 2009 | CAP | Fresh | Bio-Rad | 2.8
Bio-Rad Variant II Turbo | 147 | 6.76 | 0.16 | 2.9 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Bio-Rad | 3.06
Bio-Rad Variant II Turbo | 147 | 9.45 | -0.05 | 2.4 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Bio-Rad | 2.35
Bio-Rad Variant II Turbo | 147 | 7.58 | 0.18 | 2.6 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Bio-Rad | 2.78
Bio-Rad Variant II Turbo | 143 | 6.04 | 0.14 | 3 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Bio-Rad | 3.14
Bio-Rad Variant II Turbo | 143 | 9.66 | -0.14 | 2.3 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Bio-Rad | 2.16
Bio-Rad Variant II Turbo | 143 | 7.5 | 0.1 | 2.3 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Bio-Rad | 2.4
Bio-Rad Variant II Turbo | 146 | 5.28 | 0.08 | 2.7 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Bio-Rad | 2.78
Bio-Rad Variant II Turbo | 146 | 8.74 | 0.04 | 2.2 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Bio-Rad | 2.24
Bio-Rad Variant II Turbo | 146 | 6.38 | 0.08 | 2.2 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Bio-Rad | 2.28
Bio-Rad Variant II Turbo | 148 | 8.59 | 0.09 | 2.4 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Bio-Rad | 2.49
Bio-Rad Variant II Turbo | 148 | 5.42 | 0.02 | 2.7 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Bio-Rad | 2.72
Bio-Rad Variant II Turbo | 148 | 6.45 | 0.05 | 2.4 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Bio-Rad | 2.45
Bio-Rad Variant II Turbo | 96 | 6.45 | 0.15 | 3 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Bio-Rad | 3.15
Bio-Rad Variant II Turbo | 96 | 7.82 | 0.22 | 2.7 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Bio-Rad | 2.92
Bio-Rad Variant II Turbo | 96 | 9.45 | 0.25 | 2.3 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Bio-Rad | 2.55
Bio-Rad Variant II Turbo | 183 | 5.53 | -0.07 | 2.6 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Bio-Rad | 2.53
Bio-Rad Variant II Turbo | 183 | 9.52 | 0.12 | 2.2 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Bio-Rad | 2.32
Bio-Rad Variant II Turbo | 183 | 7.35 | 0.15 | 2.1 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Bio-Rad | 2.25
Bio-Rad Variant II Turbo | 174 | 5.4 | 0 | 2.9 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Bio-Rad | 2.9
Bio-Rad Variant II Turbo | 174 | 8.42 | 0.12 | 2.5 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Bio-Rad | 2.62
Bio-Rad Variant II Turbo | 174 | 5.66 | 0.01 | 4.2 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Bio-Rad | 4.21
Bio-Rad Variant II Turbo | 152 | 7.16 | 0.05 | 2.6 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Bio-Rad | 2.65
Bio-Rad Variant II Turbo | 152 | 9.43 | 0.11 | 2 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Bio-Rad | 2.11
Bio-Rad Variant II Turbo | 152 | 6.05 | -0.02 | 2.6 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Bio-Rad | 2.58
Bio-Rad Variant II Turbo | 133 | 8.15 | 0.1 | 2.3 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Bio-Rad | 2.4
Bio-Rad Variant II Turbo | 133 | 5.31 | 0.01 | 2.8 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Bio-Rad | 2.81
Bio-Rad Variant II Turbo | 133 | 6.2 | 0.06 | 2.3 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Bio-Rad | 2.36
Bio-Rad Variant II Turbo | 55 | 9.44 | 0.07 | 2.1 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Bio-Rad | 2.17
Bio-Rad Variant II Turbo | 55 | 4.89 | -0.03 | 2.3 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Bio-Rad | 2.27
Bio-Rad Variant II Turbo | 55 | 6.09 | 0.05 | 1.8 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Bio-Rad | 1.85
Bio-Rad Variant II Turbo | 55 | 5.32 | 0.04 | 2.1 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Bio-Rad | 2.14
Bio-Rad Variant II Turbo | 55 | 8.24 | 0.16 | 2 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Bio-Rad | 2.16
Bio-Rad Variant II Turbo | 42 | 5.40 | -0.06 | 2.5 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Bio-Rad | 2.44
Bio-Rad Variant II Turbo | 42 | 5.84 | 0.18 | 2.7 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Bio-Rad | 2.88
Bio-Rad Variant II Turbo | 42 | 9.45 | 0.14 | 1.7 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Bio-Rad | 1.84
Bio-Rad Variant II Turbo | 42 | 5.20 | -0.08 | 2 | GH5-04 | 5.28 | 2019 | CAP | Fresh | Bio-Rad | 1.92
Bio-Rad Variant II Turbo | 42 | 7.50 | 0.09 | 1.8 | GH5-05 | 7.41 | 2019 | CAP | Fresh | Bio-Rad | 1.89
Bio-Rad Variant II Turbo | 39 | 6.46 | 0.05 | 2.9 | GH5-06 | 6.41 | 2019 | CAP | Fresh | Bio-Rad | 2.95
Bio-Rad Variant II Turbo | 39 | 8.85 | 0.25 | 2.4 | GH5-07 | 8.6 | 2019 | CAP | Fresh | Bio-Rad | 2.65
Bio-Rad Variant II Turbo | 39 | 5.37 | -0.05 | 3.3 | GH5-08 | 5.42 | 2019 | CAP | Fresh | Bio-Rad | 3.25
Bio-Rad Variant II Turbo | 39 | 7.44 | 0.06 | 2.4 | GH5-09 | 7.38 | 2019 | CAP | Fresh | Bio-Rad | 2.46
Bio-Rad Variant II Turbo | 39 | 9.92 | 0.17 | 2.5 | GH5-10 | 9.75 | 2019 | CAP | Fresh | Bio-Rad | 2.67
Bio-Rad Variant II Turbo | 29 | 5.06 | -0.19 | 2.7 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Bio-Rad | 2.51
Bio-Rad Variant II Turbo | 29 | 6.42 | 0.01 | 1.9 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Bio-Rad | 1.91
Bio-Rad Variant II Turbo | 29 | 8.19 | -0.02 | 1.7 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Bio-Rad | 1.68
Bio-Rad Variant II Turbo | 29 | 5.31 | -0.16 | 2 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Bio-Rad | 1.84
Bio-Rad Variant II Turbo | 29 | 9.53 | 0.05 | 1.7 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Bio-Rad | 1.75
Bio-Rad Variant II Turbo 2.0 | 19 | 6.27 | 0.37 | 2 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Bio-Rad | 2.37
Bio-Rad Variant II Turbo 2.0 | 19 | 9.93 | 0.13 | 2 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Bio-Rad | 2.13
Bio-Rad Variant II Turbo 2.0 | 19 | 7.67 | 0.27 | 1.7 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Bio-Rad | 1.97
Bio-Rad Variant II Turbo 2.0 | 22 | 5.39 | 0.19 | 1.5 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Bio-Rad | 1.69
Bio-Rad Variant II Turbo 2.0 | 22 | 8.91 | 0.21 | 1.4 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Bio-Rad | 1.61
Bio-Rad Variant II Turbo 2.0 | 22 | 6.51 | 0.21 | 1.5 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Bio-Rad | 1.71
Bio-Rad Variant II Turbo 2.0 | 30 | 8.70 | 0.2 | 2.1 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Bio-Rad | 2.3
Bio-Rad Variant II Turbo 2.0 | 30 | 5.57 | 0.17 | 2.4 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Bio-Rad | 2.57
Bio-Rad Variant II Turbo 2.0 | 30 | 6.59 | 0.19 | 2.2 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Bio-Rad | 2.39
Bio-Rad Variant II Turbo 2.0 | 166 | 6.43 | 0.13 | 2.4 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Bio-Rad | 2.53
Bio-Rad Variant II Turbo 2.0 | 166 | 7.82 | 0.22 | 2.5 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Bio-Rad | 2.72
Bio-Rad Variant II Turbo 2.0 | 166 | 9.36 | 0.16 | 2.8 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Bio-Rad | 2.96
Bio-Rad Variant II Turbo 2.0 | 47 | 5.69 | 0.09 | 3 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Bio-Rad | 3.09
Bio-Rad Variant II Turbo 2.0 | 47 | 9.64 | 0.24 | 1.7 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Bio-Rad | 1.94
Bio-Rad Variant II Turbo 2.0 | 47 | 7.54 | 0.34 | 2.5 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Bio-Rad | 2.84
Bio-Rad Variant II Turbo 2.0 | 48 | 5.5 | 0.1 | 2.1 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Bio-Rad | 2.2
Bio-Rad Variant II Turbo 2.0 | 48 | 8.52 | 0.22 | 1.9 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Bio-Rad | 2.12
Bio-Rad Variant II Turbo 2.0 | 48 | 5.96 | 0.31 | 3.8 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Bio-Rad | 4.11
Bio-Rad Variant II Turbo 2.0 | 51 | 7.12 | 0.01 | 2.2 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Bio-Rad | 2.21
Bio-Rad Variant II Turbo 2.0 | 51 | 9.35 | 0.03 | 2.3 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Bio-Rad | 2.33
Bio-Rad Variant II Turbo 2.0 | 51 | 6.11 | 0.04 | 2.6 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Bio-Rad | 2.64
Bio-Rad Variant II Turbo 2.0 | 60 | 8.13 | 0.08 | 1.9 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Bio-Rad | 1.98
Bio-Rad Variant II Turbo 2.0 | 60 | 5.27 | -0.03 | 2.5 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Bio-Rad | 2.47
Bio-Rad Variant II Turbo 2.0 | 60 | 6.15 | 0.01 | 1.9 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Bio-Rad | 1.91
Bio-Rad Variant II Turbo 2.0 | 125 | 9.50 | 0.13 | 2.1 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Bio-Rad | 2.23
Bio-Rad Variant II Turbo 2.0 | 125 | 4.89 | -0.03 | 2.7 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Bio-Rad | 2.67
Bio-Rad Variant II Turbo 2.0 | 125 | 6.12 | 0.08 | 1.9 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Bio-Rad | 1.98
Bio-Rad Variant II Turbo 2.0 | 125 | 5.38 | 0.1 | 2.3 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Bio-Rad | 2.4
Bio-Rad Variant II Turbo 2.0 | 125 | 8.24 | 0.16 | 2 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Bio-Rad | 2.16
Bio-Rad Variant II Turbo 2.0 | 120 | 5.47 | 0.01 | 2.5 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Bio-Rad | 2.51
Bio-Rad Variant II Turbo 2.0 | 120 | 5.93 | 0.27 | 2.3 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Bio-Rad | 2.57
Bio-Rad Variant II Turbo 2.0 | 120 | 9.48 | 0.17 | 1.6 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Bio-Rad | 1.77
Bio-Rad Variant II Turbo 2.0 | 120 | 5.23 | -0.05 | 2.6 | GH5-04 | 5.28 | 2019 | CAP | Fresh | Bio-Rad | 2.55
Bio-Rad Variant II Turbo 2.0 | 120 | 7.51 | 0.1 | 1.9 | GH5-05 | 7.41 | 2019 | CAP | Fresh | Bio-Rad | 2
Bio-Rad Variant II Turbo 2.0 | 106 | 6.48 | 0.07 | 2.1 | GH5-06 | 6.41 | 2019 | CAP | Fresh | Bio-Rad | 2.17
Bio-Rad Variant II Turbo 2.0 | 106 | 8.85 | 0.25 | 2 | GH5-07 | 8.6 | 2019 | CAP | Fresh | Bio-Rad | 2.25
Bio-Rad Variant II Turbo 2.0 | 106 | 5.39 | -0.03 | 2.7 | GH5-08 | 5.42 | 2019 | CAP | Fresh | Bio-Rad | 2.67
Bio-Rad Variant II Turbo 2.0 | 106 | 7.41 | 0.03 | 1.9 | GH5-09 | 7.38 | 2019 | CAP | Fresh | Bio-Rad | 1.93
Bio-Rad Variant II Turbo 2.0 | 106 | 9.88 | 0.13 | 1.9 | GH5-10 | 9.75 | 2019 | CAP | Fresh | Bio-Rad | 2.03
Bio-Rad Variant II Turbo 2.0 | 131 | 5.05 | -0.2 | 2.7 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Bio-Rad | 2.5
Bio-Rad Variant II Turbo 2.0 | 131 | 6.41 | 0 | 1.9 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Bio-Rad | 1.9
Bio-Rad Variant II Turbo 2.0 | 131 | 8.19 | -0.02 | 1.8 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Bio-Rad | 1.78
Bio-Rad Variant II Turbo 2.0 | 131 | 5.33 | -0.14 | 2.4 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Bio-Rad | 2.26
Bio-Rad Variant II Turbo 2.0 | 131 | 9.52 | 0.04 | 1.7 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Bio-Rad | 1.74
Bio-Rad Variant II Turbo 2.0 | 138 | 7.96 | -0.01 | 2.6 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Bio-Rad | 2.59
Bio-Rad Variant II Turbo 2.0 | 138 | 5.79 | -0.1 | 2 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Bio-Rad | -0.1
Bio-Rad Variant II Turbo 2.0 | 138 | 5 | -0.13 | 2.3 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Bio-Rad | 2.17
Bio-Rad Variant II Turbo 2.0 | 138 | 7.28 | -0.12 | 2.4 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Bio-Rad | 2.28
Bio-Rad Variant II Turbo 2.0 | 138 | 9.18 | 0.01 | 2 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Bio-Rad | 0.01
Bio-Rad Variant II Turbo 2.0 | 131 | 7.8 | -0.01 | 2.2 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Bio-Rad | 2.19
Bio-Rad Variant II Turbo 2.0 | 131 | 5.85 | -0.06 | 2.2 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Bio-Rad | 2.14
Bio-Rad Variant II Turbo 2.0 | 131 | 8.84 | 0.15 | 2.1 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Bio-Rad | 2.25
Bio-Rad Variant II Turbo 2.0 | 131 | 5.97 | 0.11 | 2.3 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Bio-Rad | 2.41
Bio-Rad Variant II Turbo 2.0 | 131 | 9.59 | 0.03 | 2.4 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Bio-Rad | 2.43
Bio-Rad Variant II Turbo 2.0 | 139 | 5.5 | 0.02 | 2.3 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Bio-Rad | 2.32
Bio-Rad Variant II Turbo 2.0 | 139 | 8.32 | 0.06 | 1.9 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Bio-Rad | 1.96
Bio-Rad Variant II Turbo 2.0 | 139 | 5.04 | -0.11 | 2.6 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Bio-Rad | 2.49
Bio-Rad Variant II Turbo 2.0 | 139 | 10.22 | 0.17 | 1.9 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Bio-Rad | 2.07
Bio-Rad Variant II Turbo 2.0 | 139 | 5.94 | -0.05 | 2.3 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Bio-Rad | 2.25
Bio-Rad Variant II Turbo 2.0 | 134 | 9.03 | 0.17 | 2.2 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Bio-Rad | 2.37
Bio-Rad Variant II Turbo 2.0 | 134 | 6.9 | 0.06 | 2 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Bio-Rad | 0.06
Bio-Rad Variant II Turbo 2.0 | 134 | 9.49 | 0.1 | 2 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Bio-Rad | 0.1
Bio-Rad Variant II Turbo 2.0 | 134 | 6.05 | -0.08 | 2.2 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Bio-Rad | 2.12
Bio-Rad Variant II Turbo 2.0 | 134 | 7.52 | 0 | 1.8 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Bio-Rad | 1.8
Bio-Rad Variant II Turbo 2.0 | 113 | 5.85 | 0.05 | 2.2 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Bio-Rad | 2.25
Bio-Rad Variant II Turbo 2.0 | 113 | 9.21 | 0.09 | 1.7 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Bio-Rad | 1.79
Bio-Rad Variant II Turbo 2.0 | 113 | 5.43 | -0.04 | 2.1 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Bio-Rad | 2.06
Bio-Rad Variant II Turbo 2.0 | 113 | 7.16 | -0.09 | 1.8 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Bio-Rad | 1.71
Bio-Rad Variant II Turbo 2.0 | 113 | 9.37 | 0.11 | 1.7 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Bio-Rad | 1.81
Bio-Rad Variant II Turbo 2.0 | 121 | 5.16 | -0.05 | 2.4 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Bio-Rad | 2.35
Bio-Rad Variant II Turbo 2.0 | 121 | 6.74 | 0.06 | 1.8 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Bio-Rad | 1.86
Bio-Rad Variant II Turbo 2.0 | 121 | 8.43 |  | 2.2 | GH5-13 |  | 2021 | CAP | Fresh | Bio-Rad | 2.2
Bio-Rad Variant II Turbo 2.0 | 121 | 5.97 | -0.01 | 2 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Bio-Rad | -0.01
Bio-Rad Variant II Turbo 2.0 | 121 | 10.09 | -0.03 | 1.5 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Bio-Rad | 1.47
Bio-Rad Variant II Turbo A1c | 40 | 10.9 |  | 2.4 | GH2-01 | 10.7 | 2006 | CAP | Fresh | Bio-Rad | 2.4
Bio-Rad Variant II Turbo A1c | 40 | 5.3 |  | 2.3 | GH2-02 | 5.30 | 2006 | CAP | Fresh | Bio-Rad | 2.3
Bio-Rad Variant II Turbo A1c | 40 | 8.4 |  | 2.5 | GH2-03 | 8.40 | 2006 | CAP | Fresh | Bio-Rad | 2.5
Bio-Rad Variant II Turbo A1c | 54 | 6.9 |  | 2.6 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Bio-Rad | 2.6
Bio-Rad Variant II Turbo A1c | 54 | 6.9 |  | 2.2 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Bio-Rad | 2.2
Bio-Rad Variant II Turbo A1c | 54 | 12 |  | 2.6 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Bio-Rad | 2.6
Bio-Rad Variant II Turbo A1c | 54 | 6.9 |  | 2.6 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Bio-Rad | 2.6
Bio-Rad Variant II Turbo A1c | 54 | 6.9 |  | 2.2 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Bio-Rad | 2.2
Bio-Rad Variant II Turbo A1c | 54 | 12 |  | 2.6 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Bio-Rad | 2.6
Bio-Rad Variant II Turbo A1c | 95 | 5.4 |  | 2.7 | GH2-01 | 5.4 | 2007 | CAP | Fresh | Bio-Rad | 2.7
Bio-Rad Variant II Turbo A1c | 95 | 12 |  | 2.6 | GH2-02 | 11.6 | 2007 | CAP | Fresh | Bio-Rad | 2.6
Bio-Rad Variant II Turbo A1c | 95 | 7.8 |  | 2.3 | GH2-03 | 7.6 | 2007 | CAP | Fresh | Bio-Rad | 2.3
Bio-Rad Variant II Turbo A1c | 103 | 9.4 |  | 2.7 | GH2-04 | 9.2 | 2007 | CAP | Fresh | Bio-Rad | 2.7
Bio-Rad Variant II Turbo A1c | 103 | 11.4 |  | 2.6 | GH2-05 | 11.1 | 2007 | CAP | Fresh | Bio-Rad | 2.6
Bio-Rad Variant II Turbo A1c | 103 | 6.3 |  | 2.4 | GH2-06 | 6.2 | 2007 | CAP | Fresh | Bio-Rad | 2.4
Bio-Rad Variant II Turbo A1c | 128 | 13.5 |  | 3.3 | GH2-01 | 13 | 2008 | CAP | Fresh | Bio-Rad | 3.3
Bio-Rad Variant II Turbo A1c | 128 | 5.5 |  | 2.1 | GH2-02 | 5.3 | 2008 | CAP | Fresh | Bio-Rad | 2.1
Bio-Rad Variant II Turbo A1c | 128 | 9.3 |  | 2.7 | GH2-03 | 8.8 | 2008 | CAP | Fresh | Bio-Rad | 2.7
Bio-Rad Variant II Turbo A1c | 132 | 12 |  | 2.5 | GH2-04 | 11.9 | 2008 | CAP | Fresh | Bio-Rad | 2.5
Bio-Rad Variant II Turbo A1c | 132 | 7 |  | 2.7 | GH2-05 | 6.8 | 2008 | CAP | Fresh | Bio-Rad | 2.7
Bio-Rad Variant II Turbo A1c | 132 | 5.2 |  | 3.2 | GH2-06 | 5 | 2008 | CAP | Fresh | Bio-Rad | 3.2
Bio-Rad Variant II Turbo A1c | 130 | 5.1 |  | 2.8 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Bio-Rad | 2.8
Bio-Rad Variant II Turbo A1c | 130 | 8.5 |  | 2.5 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Bio-Rad | 2.5
Bio-Rad Variant II Turbo A1c | 130 | 6 |  | 2.6 | GH2-03 | 6 | 2009 | CAP | Fresh | Bio-Rad | 2.6
Bio-Rad VII | 83 | 6.58 | 0 | 2.6 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Bio-Rad | 2.6
Bio-Rad VII | 83 | 8.51 | 0.12 | 2.8 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Bio-Rad | 2.92
Bio-Rad VII | 83 | 5.57 | -0.08 | 2.7 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Bio-Rad | 2.62
Bio-Rad VII | 84 | 6.44 | -0.05 | 2.3 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Bio-Rad | 2.25
Bio-Rad VII | 84 | 6.98 | 0.01 | 2.1 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Bio-Rad | 2.11
Bio-Rad VII | 84 | 9.8 | 0.15 | 2.6 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Bio-Rad | 2.75
Bio-Rad VII | 70 | 6.9 | 0.11 | 2.6 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Bio-Rad | 2.71
Bio-Rad VII | 70 | 10.43 | 0.15 | 2.6 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Bio-Rad | 2.75
Bio-Rad VII | 70 | 6.92 | 0.1 | 2.9 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Bio-Rad | 3
Bio-Rad VII | 60 | 8.77 | 0.14 | 2.7 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Bio-Rad | 2.84
Bio-Rad VII | 60 | 5.29 | -0.03 | 2.9 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Bio-Rad | 2.87
Bio-Rad VII | 60 | 5.94 | 0.1 | 3.2 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Bio-Rad | 3.3
Bio-Rad VII | 60 | 12.14 | 0.43 | 3.1 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Bio-Rad | 3.53
Bio-Rad VII | 60 | 9.92 | 0.39 | 3.1 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Bio-Rad | 3.49
Bio-Rad VII | 60 | 5.04 | 0 | 3.2 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Bio-Rad | 3.2
Bio-Rad VII | 60 | 7.51 | 0.13 | 3.1 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Bio-Rad | 3.23
Bio-Rad VII | 52-65 | 12.12 | 0.43 | 3.1 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Bio-Rad | 3.53
Bio-Rad VII | 52-65 | 6.08 | 0.15 | 2.8 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Bio-Rad | 2.95
Bio-Rad VII | 52-65 | 5.21 | 0.04 | 3.3 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Bio-Rad | 3.34
Bio-Rad VII | 52-65 | 8.3 | 0.16 | 3.4 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Bio-Rad | 3.56
Bio-Rad VII | 52-65 | 9.53 | 0.23 | 2.8 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Bio-Rad | 3.03
Bio-Rad VII | 50-59 | 5.45 | 0.13 | 3 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Bio-Rad | 3.13
Bio-Rad VII | 50-59 | 9.53 | 0.37 | 2.4 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Bio-Rad | 2.77
Bio-Rad VII | 50-59 | 5.45 | 0.14 | 2.5 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Bio-Rad | 2.64
Bio-Rad VII | 50-59 | 12.55 | 0.52 | 2.4 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Bio-Rad | 2.92
Bio-Rad VII | 50-59 | 6.10 | 0.16 | 2.3 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Bio-Rad | 2.46
Bio-Rad VII | 44 | 5.39 | 0.12 | 3.1 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Bio-Rad | 3.22
Bio-Rad VII | 44 | 10.85 | 0.26 | 2.4 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Bio-Rad | 2.66
Bio-Rad VII | 44 | 6.37 | 0.17 | 2.2 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Bio-Rad | 2.37
Bio-Rad VII | 44 | 12.54 | 0.31 | 3 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Bio-Rad | 3.31
Bio-Rad VII | 44 | 7.74 | 0.23 | 2.2 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Bio-Rad | 2.43
Bio-Rad VII | 44-53 | 9.23 | 0.12 | 2.2 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Bio-Rad | 2.32
Bio-Rad VII | 44-53 | 6.12 | 0.11 | 2.7 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Bio-Rad | 2.81
Bio-Rad VII | 44-53 | 11.91 | 0.2 | 2.3 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Bio-Rad | 2.5
Bio-Rad VII | 44-53 | 5.11 | 0.09 | 2.8 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Bio-Rad | 2.89
Bio-Rad VII | 44-53 | 7.67 | 0.09 | 1.7 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Bio-Rad | 1.79
Bio-Rad VII | 43-50 | 6.51 | 0.1 | 2.6 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Bio-Rad | 2.7
Bio-Rad VII | 43-50 | 9.75 | 0.22 | 2 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Bio-Rad | 0.22
Bio-Rad VII | 43-50 | 5.34 | 0 | 2.2 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Bio-Rad | 2.2
Bio-Rad VII | 43-50 | 8.73 | 0.22 | 1.9 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Bio-Rad | 2.12
Bio-Rad VII | 43-50 | 7.29 | 0.04 | 2.1 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Bio-Rad | 2.14
Bio-Rad VII | 34 | 7.18 | 0.03 | 1.9 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Bio-Rad | 1.93
Bio-Rad VII | 34 | 5.19 | 0 | 1.8 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Bio-Rad | 1.8
Bio-Rad VII | 34 | 8.47 | 0.05 | 1.7 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Bio-Rad | 1.75
Bio-Rad VII | 34 | 9.81 | 0.02 | 1.9 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Bio-Rad | 1.92
Bio-Rad VII | 34 | 6.23 | 0.11 | 2.4 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Bio-Rad | 2.51
Bio-Rad VII | 26 | 6.39 | 0.08 | 1.7 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Bio-Rad | 1.78
Bio-Rad VII | 26 | 9.17 | 0.06 | 1.5 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Bio-Rad | 1.56
Bio-Rad VII | 26 | 5.31 | 0 | 3 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Bio-Rad | 3
Bio-Rad VII | 26 | 8.40 | 0.05 | 1.6 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Bio-Rad | 1.65
Bio-Rad VII | 26 | 7.70 | 0.08 | 2.1 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Bio-Rad | 2.18
Bio-Rad VII | 38 | 7.59 | 0.17 | 2.2 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Bio-Rad | 2.37
Bio-Rad VII | 38 | 5.33 | 0.13 | 1.5 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Bio-Rad | 1.63
Bio-Rad VII | 38 | 8.42 | 0.11 | 2.5 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Bio-Rad | 2.61
Bio-Rad VII | 38 | 10.56 | 0.23 | 2.7 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Bio-Rad | 2.93
Bio-Rad VII | 38 | 6.01 | 0.14 | 1.9 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Bio-Rad | 2.04
Bio-Rad VII | 39 | 7.32 | 0.07 | 2.4 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Bio-Rad | 2.47
Bio-Rad VII | 39 | 6.92 | 0.11 | 2.5 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Bio-Rad | 2.61
Bio-Rad VII | 39 | 8.66 | 0.01 | 2.4 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Bio-Rad | 2.41
Bio-Rad VII | 39 | 9.56 | 0.06 | 2.6 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Bio-Rad | 2.66
Bio-Rad VII | 39 | 5.47 | 0.02 | 2.7 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Bio-Rad | 2.72
Bio-Rad VII Turbo | 127 | 6.7 | 0.12 | 2.2 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Bio-Rad | 2.32
Bio-Rad VII Turbo | 127 | 8.56 | 0.17 | 2.1 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Bio-Rad | 2.27
Bio-Rad VII Turbo | 127 | 5.7 | 0.05 | 2.6 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Bio-Rad | 2.65
Bio-Rad VII Turbo | 151 | 6.57 | 0.08 | 2.5 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Bio-Rad | 2.58
Bio-Rad VII Turbo | 151 | 7.13 | 0.16 | 2.2 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Bio-Rad | 2.36
Bio-Rad VII Turbo | 151 | 9.84 | 0.19 | 2.2 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Bio-Rad | 2.39
Bio-Rad VII Turbo | 104 | 6.79 | 0 | 3.2 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Bio-Rad | 3.2
Bio-Rad VII Turbo | 104 | 10.33 | 0.05 | 2.5 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Bio-Rad | 2.55
Bio-Rad VII Turbo | 104 | 6.82 | 0 | 3 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Bio-Rad | 3
Bio-Rad VII Turbo | 85 | 8.69 | 0.06 | 2.5 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Bio-Rad | 2.56
Bio-Rad VII Turbo | 85 | 5.3 | -0.02 | 2.9 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Bio-Rad | 2.88
Bio-Rad VII Turbo | 82 | 5.8 | -0.04 | 1.7 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Bio-Rad | 1.66
Bio-Rad VII Turbo | 82 | 11.73 | 0.02 | 2.5 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Bio-Rad | 2.52
Bio-Rad VII Turbo | 82 | 9.6 | 0.07 | 2.4 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Bio-Rad | 2.47
Bio-Rad VII Turbo | 82 | 5.01 | -0.03 | 2.3 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Bio-Rad | 2.27
Bio-Rad VII Turbo | 82 | 7.38 | 0 | 2.1 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Bio-Rad | 2.1
Bio-Rad VII Turbo | 83-99 | 11.73 | 0.04 | 2.4 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Bio-Rad | 2.44
Bio-Rad VII Turbo | 83-99 | 5.92 | -0.01 | 1.9 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Bio-Rad | 1.89
Bio-Rad VII Turbo | 83-99 | 5.15 | -0.02 | 2 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Bio-Rad | 1.98
Bio-Rad VII Turbo | 83-99 | 8.15 | 0.01 | 2.1 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Bio-Rad | 2.11
Bio-Rad VII Turbo | 83-99 | 9.32 | 0.02 | 2.3 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Bio-Rad | 2.32
Bio-Rad VII Turbo | 82-89 | 5.34 | 0.02 | 2.5 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Bio-Rad | 2.52
Bio-Rad VII Turbo | 82-89 | 9.34 | 0.18 | 1.9 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Bio-Rad | 2.08
Bio-Rad VII Turbo | 82-89 | 5.33 | 0.02 | 2.1 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Bio-Rad | 2.12
Bio-Rad VII Turbo | 82-89 | 12.30 | 0.27 | 2.2 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Bio-Rad | 2.47
Bio-Rad VII Turbo | 82-89 | 5.98 | 0.04 | 1.6 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Bio-Rad | 1.64
Bio-Rad VII Turbo | 79 | 5.27 | 0 | 2.3 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Bio-Rad | 2.3
Bio-Rad VII Turbo | 79 | 10.73 | 0.14 | 2.4 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Bio-Rad | 2.54
Bio-Rad VII Turbo | 79 | 6.24 | 0.04 | 2.6 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Bio-Rad | 2.64
Bio-Rad VII Turbo | 79 | 12.36 | 0.13 | 2.8 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Bio-Rad | 2.93
Bio-Rad VII Turbo | 79 | 7.66 | 0.15 | 2.4 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Bio-Rad | 2.55
Bio-Rad VII Turbo | 77-84 | 9.16 | 0.05 | 2.4 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Bio-Rad | 2.45
Bio-Rad VII Turbo | 77-84 | 5.99 | -0.02 | 2.1 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Bio-Rad | 2.08
Bio-Rad VII Turbo | 77-84 | 11.73 | 0.02 | 2.8 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Bio-Rad | 2.82
Bio-Rad VII Turbo | 77-84 | 5.00 | -0.02 | 2.3 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Bio-Rad | 2.28
Bio-Rad VII Turbo | 77-84 | 7.65 | 0.07 | 2 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Bio-Rad | 2.07
Bio-Rad VII Turbo | 73-78 | 6.48 | 0.07 | 2.6 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Bio-Rad | 2.67
Bio-Rad VII Turbo | 73-78 | 9.84 | 0.31 | 2.3 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Bio-Rad | 2.61
Bio-Rad VII Turbo | 73-78 | 5.36 | 0.02 | 2.8 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Bio-Rad | 2.82
Bio-Rad VII Turbo | 73-78 | 8.81 | 0.3 | 2.6 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Bio-Rad | 2.9
Bio-Rad VII Turbo | 73-78 | 7.40 | 0.15 | 2.5 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Bio-Rad | 2.65
Bio-Rad VII Turbo | 62 | 7.12 | -0.03 | 2.2 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Bio-Rad | 2.17
Bio-Rad VII Turbo | 62 | 5.13 | -0.06 | 2.5 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Bio-Rad | 2.44
Bio-Rad VII Turbo | 62 | 8.48 | 0.06 | 2.4 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Bio-Rad | 2.46
Bio-Rad VII Turbo | 62 | 9.83 | 0.04 | 2.7 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Bio-Rad | 2.74
Bio-Rad VII Turbo | 62 | 6.13 | 0.01 | 2.9 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Bio-Rad | 2.91
Bio-Rad VII Turbo | 56 | 6.29 | -0.02 | 3.3 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Bio-Rad | 3.28
Bio-Rad VII Turbo | 56 | 9.15 | 0.04 | 2.6 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Bio-Rad | 2.64
Bio-Rad VII Turbo | 56 | 5.21 | -0.1 | 3.2 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Bio-Rad | 3.1
Bio-Rad VII Turbo | 56 | 8.41 | 0.06 | 2.4 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Bio-Rad | 2.46
Bio-Rad VII Turbo | 56 | 7.68 | 0.06 | 2.4 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Bio-Rad | 2.46
Bio-Rad VII Turbo | 62 | 7.51 | 0.09 | 1.7 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Bio-Rad | 1.79
Bio-Rad VII Turbo | 62 | 5.26 | 0.06 | 2.1 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Bio-Rad | 2.16
Bio-Rad VII Turbo | 62 | 8.42 | 0.11 | 2 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Bio-Rad | 2.11
Bio-Rad VII Turbo | 62 | 10.50 | 0.17 | 2.2 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Bio-Rad | 2.37
Bio-Rad VII Turbo | 62 | 5.90 | 0.03 | 1.9 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Bio-Rad | 1.93
Bio-Rad VII Turbo | 67 | 7.31 | 0.06 | 2.5 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Bio-Rad | 2.56
Bio-Rad VII Turbo | 67 | 6.85 | 0.04 | 2.3 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Bio-Rad | 2.34
Bio-Rad VII Turbo | 67 | 8.72 | 0.07 | 2.8 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Bio-Rad | 2.87
Bio-Rad VII Turbo | 67 | 9.66 | 0.16 | 2.9 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Bio-Rad | 3.06
Bio-Rad VII Turbo | 67 | 5.41 | -0.04 | 2.5 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Bio-Rad | 2.46
Bio-Rad VII Turbo 2.0 | 124 | 6.64 | 0.06 | 2.6 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Bio-Rad | 2.66
Bio-Rad VII Turbo 2.0 | 124 | 8.54 | 0.15 | 2.3 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Bio-Rad | 2.45
Bio-Rad VII Turbo 2.0 | 124 | 5.65 | 0 | 3.1 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Bio-Rad | 3.1
Bio-Rad VII Turbo 2.0 | 103 | 6.45 | -0.04 | 2 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Bio-Rad | 1.96
Bio-Rad VII Turbo 2.0 | 103 | 7.03 | 0.06 | 2 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Bio-Rad | 2.06
Bio-Rad VII Turbo 2.0 | 103 | 9.71 | 0.06 | 1.7 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Bio-Rad | 1.76
Bio-Rad VII Turbo 2.0 | 144 | 6.89 | 0.1 | 2.3 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Bio-Rad | 2.4
Bio-Rad VII Turbo 2.0 | 144 | 10.26 | -0.02 | 2.5 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Bio-Rad | 2.48
Bio-Rad VII Turbo 2.0 | 144 | 6.88 | 0.06 | 2.6 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Bio-Rad | 2.66
Bio-Rad VII Turbo 2.0 | 121 | 8.67 | 0.04 | 2.7 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Bio-Rad | 2.74
Bio-Rad VII Turbo 2.0 | 121 | 5.25 | -0.07 | 2.7 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Bio-Rad | 2.63
Bio-Rad VII Turbo 2.0 | 132 | 5.84 | 0 | 2.3 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Bio-Rad | 2.3
Bio-Rad VII Turbo 2.0 | 132 | 11.57 | -0.14 | 2.3 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Bio-Rad | 2.16
Bio-Rad VII Turbo 2.0 | 132 | 9.55 | 0.02 | 2.5 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Bio-Rad | 2.52
Bio-Rad VII Turbo 2.0 | 132 | 5.01 | -0.03 | 2.7 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Bio-Rad | 2.67
Bio-Rad VII Turbo 2.0 | 132 | 7.35 | -0.03 | 2.3 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Bio-Rad | 2.27
Bio-Rad VII Turbo 2.0 | 141-171 | 11.59 | -0.1 | 2.4 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Bio-Rad | 2.3
Bio-Rad VII Turbo 2.0 | 141-171 | 5.92 | -0.01 | 2 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Bio-Rad | 1.99
Bio-Rad VII Turbo 2.0 | 141-171 | 5.14 | -0.03 | 2.2 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Bio-Rad | 2.17
Bio-Rad VII Turbo 2.0 | 141-171 | 8.09 | -0.05 | 1.6 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Bio-Rad | 1.55
Bio-Rad VII Turbo 2.0 | 141-171 | 9.18 | -0.12 | 1.6 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Bio-Rad | 1.48
Bio-Rad VII Turbo 2.0 | 146-173 | 5.37 | 0.05 | 2.1 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Bio-Rad | 2.15
Bio-Rad VII Turbo 2.0 | 146-173 | 9.24 | 0.08 | 1.7 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Bio-Rad | 1.78
Bio-Rad VII Turbo 2.0 | 146-173 | 5.35 | 0.04 | 2.3 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Bio-Rad | 2.34
Bio-Rad VII Turbo 2.0 | 146-173 | 12.00 | -0.03 | 1.7 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Bio-Rad | 1.67
Bio-Rad VII Turbo 2.0 | 146-173 | 5.99 | 0.05 | 1.9 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Bio-Rad | 1.95
Bio-Rad VII Turbo 2.0 | 145 | 5.23 | -0.04 | 2.6 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Bio-Rad | 2.56
Bio-Rad VII Turbo 2.0 | 145 | 10.62 | 0.03 | 1.9 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Bio-Rad | 1.93
Bio-Rad VII Turbo 2.0 | 145 | 6.23 | 0.03 | 2.2 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Bio-Rad | 2.23
Bio-Rad VII Turbo 2.0 | 145 | 12.09 | -0.14 | 2.2 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Bio-Rad | 2.06
Bio-Rad VII Turbo 2.0 | 145 | 7.53 | 0.02 | 2 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Bio-Rad | 2.02
Bio-Rad VII Turbo 2.0 | 147-173 | 9.13 | 0.02 | 2 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Bio-Rad | 2.02
Bio-Rad VII Turbo 2.0 | 147-173 | 5.99 | -0.02 | 1.8 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Bio-Rad | 1.78
Bio-Rad VII Turbo 2.0 | 147-173 | 11.49 | -0.22 | 2 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Bio-Rad | 1.78
Bio-Rad VII Turbo 2.0 | 147-173 | 4.98 | -0.04 | 2 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Bio-Rad | 1.96
Bio-Rad VII Turbo 2.0 | 147-173 | 7.54 | -0.04 | 1.7 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Bio-Rad | 1.66
Bio-Rad VII Turbo 2.0 | 138-157 | 6.41 | 0 | 1.7 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Bio-Rad | 1.7
Bio-Rad VII Turbo 2.0 | 138-157 | 9.63 | 0.1 | 1.7 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Bio-Rad | 1.8
Bio-Rad VII Turbo 2.0 | 138-157 | 5.31 | -0.03 | 2.1 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Bio-Rad | 2.07
Bio-Rad VII Turbo 2.0 | 138-157 | 8.57 | 0.06 | 1.6 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Bio-Rad | 1.66
Bio-Rad VII Turbo 2.0 | 138-157 | 7.24 | -0.01 | 1.6 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Bio-Rad | 1.59
Bio-Rad VII Turbo 2.0 | 142 | 7.50 | 0.08 | 2.5 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Bio-Rad | 2.58
Bio-Rad VII Turbo 2.0 | 142 | 5.30 | 0.1 | 2.6 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Bio-Rad | 2.7
Bio-Rad VII Turbo 2.0 | 142 | 8.31 | 0 | 2.1 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Bio-Rad | 2.1
Bio-Rad VII Turbo 2.0 | 142 | 10.40 | 0.07 | 2.1 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Bio-Rad | 2.17
Bio-Rad VII Turbo 2.0 | 142 | 5.92 | 0.05 | 2.2 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Bio-Rad | 2.25
Bio-Rad VII Turbo 2.0 | 152 | 7.29 | 0.04 | 2 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Bio-Rad | 0.04
Bio-Rad VII Turbo 2.0 | 152 | 6.86 | 0.05 | 2 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Bio-Rad | 2.05
Bio-Rad VII Turbo 2.0 | 152 | 8.58 | -0.07 | 1.8 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Bio-Rad | 1.73
Bio-Rad VII Turbo 2.0 | 152 | 9.54 | 0.04 | 1.9 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Bio-Rad | 1.94
Bio-Rad VII Turbo 2.0 | 152 | 5.49 | 0.04 | 2.2 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Bio-Rad | 2.24
Bio-Rad VII Turbo 2.0 | 142 | 7.16 | 0.01 | 1.8 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Bio-Rad | 1.81
Bio-Rad VII Turbo 2.0 | 142 | 5.11 | -0.08 | 2.7 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Bio-Rad | 2.62
Bio-Rad VII Turbo 2.0 | 142 | 8.46 | 0.04 | 2.2 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Bio-Rad | 2.24
Bio-Rad VII Turbo 2.0 | 142 | 9.75 | -0.04 | 2 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Bio-Rad | 1.96
Bio-Rad VII Turbo 2.0 | 142 | 6.13 | 0.01 | 1.8 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Bio-Rad | 1.81
Bio-Rad VII Turbo 2.0 | 115 | 6.31 | 0 | 2.3 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Bio-Rad | 2.3
Bio-Rad VII Turbo 2.0 | 115 | 9.11 | 0 | 2 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Bio-Rad | 2
Bio-Rad VII Turbo 2.0 | 115 | 5.20 | -0.11 | 2.9 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Bio-Rad | 2.79
Bio-Rad VII Turbo 2.0 | 115 | 8.36 | 0.01 | 2.1 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Bio-Rad | 2.11
Bio-Rad VII Turbo 2.0 | 115 | 7.58 | -0.04 | 1.9 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Bio-Rad | 1.86
Dade Behring Dimension | 490 | 10.5 |  | 3.1 | GH2-01 | 10.7 | 2006 | CAP | Fresh | Dade Behring | 3.1
Dade Behring Dimension | 490 | 5.6 |  | 3.6 | GH2-02 | 5.30 | 2006 | CAP | Fresh | Dade Behring | 3.6
Dade Behring Dimension | 490 | 8.1 |  | 3.2 | GH2-03 | 8.40 | 2006 | CAP | Fresh | Dade Behring | 3.2
Dade Behring Dimension | 494 | 7.1 |  | 3.7 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Dade Behring | 3.7
Dade Behring Dimension | 494 | 7.1 |  | 3.6 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Dade Behring | 3.6
Dade Behring Dimension | 494 | 11.5 |  | 3.7 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Dade Behring | 3.7
Dade Behring Dimension | 494 | 7.1 |  | 3.7 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Dade Behring | 3.7
Dade Behring Dimension | 494 | 7.1 |  | 3.6 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Dade Behring | 3.6
Dade Behring Dimension | 494 | 11.5 |  | 3.7 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Dade Behring | 3.7
Dade Behring Dimension | 593 | 5.8 |  | 3.1 | GH2-01 | 5.4 | 2007 | CAP | Fresh | Dade Behring | 3.1
Dade Behring Dimension | 593 | 11.3 |  | 3.5 | GH2-02 | 11.6 | 2007 | CAP | Fresh | Dade Behring | 3.5
Dade Behring Dimension | 593 | 7.6 |  | 2.8 | GH2-03 | 7.6 | 2007 | CAP | Fresh | Dade Behring | 2.8
Dade Behring Dimension | 527 | 8.9 |  | 3.1 | GH2-04 | 9.2 | 2007 | CAP | Fresh | Dade Behring | 3.1
Dade Behring Dimension | 527 | 10.9 |  | 3.1 | GH2-05 | 11.1 | 2007 | CAP | Fresh | Dade Behring | 3.1
Dade Behring Dimension | 527 | 6.3 |  | 3.3 | GH2-06 | 6.2 | 2007 | CAP | Fresh | Dade Behring | 3.3
Dade Behring Dimension | 574 | 12.9 |  | 4.4 | GH2-01 | 13 | 2008 | CAP | Fresh | Dade Behring | 4.4
Dade Behring Dimension | 574 | 5.6 |  | 3.1 | GH2-02 | 5.3 | 2008 | CAP | Fresh | Dade Behring | 3.1
Dade Behring Dimension | 574 | 8.6 |  | 3 | GH2-03 | 8.8 | 2008 | CAP | Fresh | Dade Behring | 3
Dade Behring Dimension | 578 | 12.3 |  | 4.6 | GH2-04 | 11.9 | 2008 | CAP | Fresh | Dade Behring | 4.6
Dade Behring Dimension | 578 | 6.9 |  | 3.7 | GH2-05 | 6.8 | 2008 | CAP | Fresh | Dade Behring | 3.7
Dade Behring Dimension | 578 | 5.4 |  | 3.5 | GH2-06 | 5 | 2008 | CAP | Fresh | Dade Behring | 3.5
Dade Behring Dimension Vista | 13 | 11.4 |  | 2.4 | GH2-04 | 11.9 | 2008 | CAP | Fresh | Dade Behring | 2.4
Dade Behring Dimension Vista | 13 | 7.2 |  | 4.6 | GH2-05 | 6.8 | 2008 | CAP | Fresh | Dade Behring | 4.6
Dade Behring Dimension Vista | 13 | 5 |  | 4.7 | GH2-06 | 5 | 2008 | CAP | Fresh | Dade Behring | 4.7
JEOL Biomajesty JCA-BM series | 8-10 | 11.48 | -0.21 | 2.9 | GH5-11 | 11.69 | 2015 | CAP | Fresh | JEOL | 2.69
JEOL Biomajesty JCA-BM series | 8-10 |  |  |  | GH5-12 | 5.93 | 2015 | CAP | Fresh | JEOL | 0
JEOL Biomajesty JCA-BM series | 8-10 |  |  |  | GH5-13 | 5.17 | 2015 | CAP | Fresh | JEOL | 0
JEOL Biomajesty JCA-BM series | 8-10 |  |  |  | GH5-14 | 8.14 | 2015 | CAP | Fresh | JEOL | 0
JEOL Biomajesty JCA-BM series | 8-10 |  |  |  | GH5-15 | 9.3 | 2015 | CAP | Fresh | JEOL | 0
JEOL Biomajesty JCA-BM series | 8-10 | 5.28 | -0.04 | 2.8 | GH5-01 | 5.32 | 2016 | CAP | Fresh | JEOL | 2.76
JEOL Biomajesty JCA-BM series | 8-10 | 9.10 | -0.06 | 1.7 | GH5-02 | 9.16 | 2016 | CAP | Fresh | JEOL | 1.64
JEOL Biomajesty JCA-BM series | 8-10 | 5.30 | -0.01 | 2.8 | GH5-03 | 5.31 | 2016 | CAP | Fresh | JEOL | 2.79
JEOL Biomajesty JCA-BM series | 8-10 |  |  |  | GH5-04 | 12.03 | 2016 | CAP | Fresh | JEOL | 0
JEOL Biomajesty JCA-BM series | 8-10 |  |  |  | GH5-05 | 5.94 | 2016 | CAP | Fresh | JEOL | 0
JEOL Biomajesty JCA-BM series | 7-10 | 8.97 | -0.14 | 1.8 | GH5-11 | 9.11 | 2016 | CAP | Fresh | JEOL | 1.66
JEOL Biomajesty JCA-BM series | 7-10 |  |  |  | GH5-12 | 6.01 | 2016 | CAP | Fresh | JEOL | 0
JEOL Biomajesty JCA-BM series | 7-10 | 12.02 | 0.31 | 1.7 | GH5-13 | 11.71 | 2016 | CAP | Fresh | JEOL | 2.01
JEOL Biomajesty JCA-BM series | 7-10 |  |  |  | GH5-14 | 5.02 | 2016 | CAP | Fresh | JEOL | 0
JEOL Biomajesty JCA-BM series | 7-10 |  |  |  | GH5-15 | 7.58 | 2016 | CAP | Fresh | JEOL | 0
Metrika A1cNOW | 15 | 10.4 |  | 8.1 | GH2-01 | 10.7 | 2006 | CAP | Fresh | Metrika | 8.1
Metrika A1cNOW | 15 | 5.3 |  | 6.3 | GH2-02 | 5.30 | 2006 | CAP | Fresh | Metrika | 6.3
Metrika A1cNOW | 15 | 8.2 |  | 7.6 | GH2-03 | 8.40 | 2006 | CAP | Fresh | Metrika | 7.6
Metrika A1cNOW | 18 | 6.7 |  | 7 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Metrika | 7
Metrika A1cNOW | 18 | 6.7 |  | 7.7 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Metrika | 7.7
Metrika A1cNOW | 18 | 11.4 |  | 7.5 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Metrika | 7.5
Metrika A1cNOW | 18 | 6.7 |  | 7 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Metrika | 7
Metrika A1cNOW | 18 | 6.7 |  | 7.7 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Metrika | 7.7
Metrika A1cNOW | 18 | 11.4 |  | 7.5 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Metrika | 7.5
Metrika A1cNOW | 23 | 5.3 |  | 6.1 | GH2-01 | 5.4 | 2007 | CAP | Fresh | Metrika | 6.1
Metrika A1cNOW | 23 | 11.7 |  | 6.3 | GH2-02 | 11.6 | 2007 | CAP | Fresh | Metrika | 6.3
Metrika A1cNOW | 23 | 7.4 |  | 7.7 | GH2-03 | 7.6 | 2007 | CAP | Fresh | Metrika | 7.7
Metrika A1cNOW | 19 | 8.2 |  | 7.4 | GH2-04 | 9.2 | 2007 | CAP | Fresh | Metrika | 7.4
Metrika A1cNOW | 19 | 10.2 |  | 6.5 | GH2-05 | 11.1 | 2007 | CAP | Fresh | Metrika | 6.5
Metrika A1cNOW | 19 | 5.7 |  | 5.7 | GH2-06 | 6.2 | 2007 | CAP | Fresh | Metrika | 5.7
Metrika A1cNOW | 19 | 11.9 |  | 5.6 | GH2-01 | 13 | 2008 | CAP | Fresh | Metrika | 5.6
Metrika A1cNOW | 19 | 4.9 |  | 3.7 | GH2-02 | 5.3 | 2008 | CAP | Fresh | Metrika | 3.7
Metrika A1cNOW | 19 | 8.2 |  | 6.6 | GH2-03 | 8.8 | 2008 | CAP | Fresh | Metrika | 6.6
Metrika A1cNOW | 20 | 10.4 |  | 6.9 | GH2-04 | 11.9 | 2008 | CAP | Fresh | Metrika | 6.9
Metrika A1cNOW | 20 | 6.1 |  | 7.7 | GH2-05 | 6.8 | 2008 | CAP | Fresh | Metrika | 7.7
Metrika A1cNOW | 20 | 4.8 |  | 7 | GH2-06 | 5 | 2008 | CAP | Fresh | Metrika | 7
Metrika A1cNOW | 17 | 4.6 |  | 5 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Metrika | 5
Metrika A1cNOW | 17 | 7.2 |  | 6 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Metrika | 6
Metrika A1cNOW | 17 | 5.3 |  | 5.2 | GH2-03 | 6 | 2009 | CAP | Fresh | Metrika | 5.2
Metrika A1cNOW | 19 | 5.98 | -0.62 | 7.9 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Metrika | 7.28
Metrika A1cNOW | 19 | 8.51 | -0.99 | 6.5 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Metrika | 5.51
Metrika A1cNOW | 19 | 6.72 | -0.68 | 6 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Metrika | 5.32
Metrika A1cNOW | 21 | 5.63 | -0.27 | 5.6 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Metrika | 5.33
Metrika A1cNOW | 21 | 9.09 | -0.71 | 4.6 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Metrika | 3.89
Metrika A1cNOW | 21 | 6.84 | -0.56 | 3.5 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Metrika | 2.94
Olympus AU 400/400e | 15 | 4.8 |  | 5.3 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Olympus | 5.3
Olympus AU 400/400e | 15 | 8.1 |  | 7.3 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Olympus | 7.3
Olympus AU 400/400e | 15 | 5.6 |  | 5.7 | GH2-03 | 6 | 2009 | CAP | Fresh | Olympus | 5.7
Olympus AU system | 22 | 11.2 |  | 4 | GH2-01 | 10.7 | 2006 | CAP | Fresh | Olympus | 4
Olympus AU system | 22 | 5.5 |  | 3.9 | GH2-02 | 5.30 | 2006 | CAP | Fresh | Olympus | 3.9
Olympus AU system | 22 | 8.9 |  | 5.5 | GH2-03 | 8.40 | 2006 | CAP | Fresh | Olympus | 5.5
Olympus AU system | 26 | 7.3 |  | 4.3 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Olympus | 4.3
Olympus AU system | 26 | 7.2 |  | 4.3 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Olympus | 4.3
Olympus AU system | 26 | 12.4 |  | 4.5 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Olympus | 4.5
Olympus AU system | 26 | 7.3 |  | 4.3 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Olympus | 4.3
Olympus AU system | 26 | 7.2 |  | 4.3 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Olympus | 4.3
Olympus AU system | 26 | 12.4 |  | 4.5 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Olympus | 4.5
Olympus AU system | 30 | 5.6 |  | 6.2 | GH2-01 | 5.4 | 2007 | CAP | Fresh | Olympus | 6.2
Olympus AU system | 30 | 12.2 |  | 6.1 | GH2-02 | 11.6 | 2007 | CAP | Fresh | Olympus | 6.1
Olympus AU system | 30 | 8.2 |  | 5.5 | GH2-03 | 7.6 | 2007 | CAP | Fresh | Olympus | 5.5
Olympus AU system | 25 | 9.4 |  | 3.6 | GH2-04 | 9.2 | 2007 | CAP | Fresh | Olympus | 3.6
Olympus AU system | 25 | 11 |  | 4.1 | GH2-05 | 11.1 | 2007 | CAP | Fresh | Olympus | 4.1
Olympus AU system | 25 | 6.1 |  | 3.6 | GH2-06 | 6.2 | 2007 | CAP | Fresh | Olympus | 3.6
Olympus AU system | 30 | 12.8 |  | 6 | GH2-01 | 13 | 2008 | CAP | Fresh | Olympus | 6
Olympus AU system | 30 | 5.2 |  | 3.5 | GH2-02 | 5.3 | 2008 | CAP | Fresh | Olympus | 3.5
Olympus AU system | 30 | 8.9 |  | 5.4 | GH2-03 | 8.8 | 2008 | CAP | Fresh | Olympus | 5.4
Olympus AU system | 25 | 11.9 |  | 6.3 | GH2-04 | 11.9 | 2008 | CAP | Fresh | Olympus | 6.3
Olympus AU system | 25 | 6.7 |  | 4.9 | GH2-05 | 6.8 | 2008 | CAP | Fresh | Olympus | 4.9
Olympus AU system | 25 | 4.9 |  | 4 | GH2-06 | 5 | 2008 | CAP | Fresh | Olympus | 4
Olympus AU system | 25 | 6.46 | -0.14 | 4.4 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Olympus | 4.26
Olympus AU system | 25 | 9.24 | -0.26 | 4.8 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Olympus | 4.54
Olympus AU system | 25 | 7.32 | -0.08 | 4.8 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Olympus | 4.72
Olympus AU system | 19 | 5.74 | -0.16 | 3.9 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Olympus | 3.74
Olympus AU system | 19 | 9.67 | -0.13 | 5.1 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Olympus | 4.97
Olympus AU system | 19 | 7.39 | -0.01 | 4.5 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Olympus | 4.49
Pointe Scientific | 5 | 5 |  |  | GH2-01 | 5.4 | 2007 | CAP | Fresh | Pointe Scientific | 0
Pointe Scientific | 5 | 11 |  |  | GH2-02 | 11.6 | 2007 | CAP | Fresh | Pointe Scientific | 0
Pointe Scientific | 5 | 7.5 |  |  | GH2-03 | 7.6 | 2007 | CAP | Fresh | Pointe Scientific | 0
Primus | 9 | 15.2 |  |  | GH2-04 | 10.7 | 2006 | CAP | Fresh | Primus | 0
Primus | 9 | 6 |  |  | GH2-05 | 5.30 | 2006 | CAP | Fresh | Primus | 0
Primus | 9 | 11.2 |  |  | GH2-06 | 8.40 | 2006 | CAP | Fresh | Primus | 0
Primus | 6 | 8.3 |  |  | GH2-04 | 7.1 | 2006 | CAP | Fresh | Primus | 0
Primus | 6 | 8.3 |  |  | GH2-05 | 7.0 | 2006 | CAP | Fresh | Primus | 0
Primus | 6 | 16.9 |  |  | GH2-06 | 11.7 | 2006 | CAP | Fresh | Primus | 0
Primus | 15 | 6.2 |  | 4.3 | GH2-01 | 5.4 | 2007 | CAP | Fresh | Primus | 4.3
Primus | 15 | 17.1 |  | 2.3 | GH2-02 | 11.6 | 2007 | CAP | Fresh | Primus | 2.3
Primus | 15 | 10 |  | 2.9 | GH2-03 | 7.6 | 2007 | CAP | Fresh | Primus | 2.9
Primus | 16 | 12.1 |  | 5.2 | GH2-04 | 9.2 | 2007 | CAP | Fresh | Primus | 5.2
Primus | 16 | 15.4 |  | 5.1 | GH2-05 | 11.1 | 2007 | CAP | Fresh | Primus | 5.1
Primus | 16 | 7.3 |  | 5.1 | GH2-06 | 6.2 | 2007 | CAP | Fresh | Primus | 5.1
Primus | 16 | 19.2 |  | 5.1 | GH2-01 | 13 | 2008 | CAP | Fresh | Primus | 5.1
Primus | 16 | 6.1 |  | 5.7 | GH2-02 | 5.3 | 2008 | CAP | Fresh | Primus | 5.7
Primus | 16 | 11.8 |  | 6 | GH2-03 | 8.8 | 2008 | CAP | Fresh | Primus | 6
Primus | 20 | 17.5 |  | 2.1 | GH2-04 | 11.9 | 2008 | CAP | Fresh | Primus | 2.1
Primus | 20 | 8.7 |  | 1.7 | GH2-05 | 6.8 | 2008 | CAP | Fresh | Primus | 1.7
Primus | 20 | 5.7 |  | 4.2 | GH2-06 | 5 | 2008 | CAP | Fresh | Primus | 4.2
Primus HPLC (affinity) | 25 | 10.9 |  | 3.2 | GH2-01 | 10.7 | 2006 | CAP | Fresh | Primus | 3.2
Primus HPLC (affinity) | 25 | 5.1 |  | 3.5 | GH2-02 | 5.30 | 2006 | CAP | Fresh | Primus | 3.5
Primus HPLC (affinity) | 25 | 8.4 |  | 3.3 | GH2-03 | 8.40 | 2006 | CAP | Fresh | Primus | 3.3
Primus HPLC (affinity) | 27 | 6.5 |  | 4 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Primus | 4
Primus HPLC (affinity) | 27 | 6.5 |  | 3.9 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Primus | 3.9
Primus HPLC (affinity) | 27 | 11.5 |  | 3.2 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Primus | 3.2
Primus HPLC (affinity) | 27 | 6.5 |  | 4 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Primus | 4
Primus HPLC (affinity) | 27 | 6.5 |  | 3.9 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Primus | 3.9
Primus HPLC (affinity) | 27 | 11.5 |  | 3.2 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Primus | 3.2
Primus HPLC (affinity) | 36 | 5.3 |  | 2.7 | GH2-01 | 5.4 | 2007 | CAP | Fresh | Primus | 2.7
Primus HPLC (affinity) | 36 | 11.7 |  | 2.5 | GH2-02 | 11.6 | 2007 | CAP | Fresh | Primus | 2.5
Primus HPLC (affinity) | 36 | 7.6 |  | 2.3 | GH2-03 | 7.6 | 2007 | CAP | Fresh | Primus | 2.3
Primus HPLC (affinity) | 25 | 9 |  | 4.8 | GH2-04 | 9.2 | 2007 | CAP | Fresh | Primus | 4.8
Primus HPLC (affinity) | 25 | 11 |  | 4.7 | GH2-05 | 11.1 | 2007 | CAP | Fresh | Primus | 4.7
Primus HPLC (affinity) | 25 | 6 |  | 4.6 | GH2-06 | 6.2 | 2007 | CAP | Fresh | Primus | 4.6
Primus HPLC (affinity) | 24 | 12.8 |  | 3.2 | GH2-01 | 13 | 2008 | CAP | Fresh | Primus | 3.2
Primus HPLC (affinity) | 24 | 5.3 |  | 2.8 | GH2-02 | 5.3 | 2008 | CAP | Fresh | Primus | 2.8
Primus HPLC (affinity) | 24 | 8.6 |  | 3 | GH2-03 | 8.8 | 2008 | CAP | Fresh | Primus | 3
Primus HPLC (affinity) | 32 | 12 |  | 2.2 | GH2-04 | 11.9 | 2008 | CAP | Fresh | Primus | 2.2
Primus HPLC (affinity) | 32 | 6.8 |  | 3.2 | GH2-05 | 6.8 | 2008 | CAP | Fresh | Primus | 3.2
Primus HPLC (affinity) | 32 | 5.1 |  | 2.7 | GH2-06 | 5 | 2008 | CAP | Fresh | Primus | 2.7
Primus HPLC (affinity) | 12 | 5.1 |  | 2.8 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Primus | 2.8
Primus HPLC (affinity) | 12 | 8.3 |  | 2.8 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Primus | 2.8
Primus HPLC (affinity) | 12 | 5.9 |  | 2.3 | GH2-03 | 6 | 2009 | CAP | Fresh | Primus | 2.3
Roche Cobas | 33 | 6.58 | 0.28 | 2.9 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Roche | 3.18
Roche Cobas | 33 | 7.95 | 0.35 | 2.9 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Roche | 3.25
Roche Cobas | 33 | 9.41 | 0.21 | 2.4 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Roche | 2.61
Roche Cobas | 12 | 5.55 | -0.05 | 5 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Roche | 4.95
Roche Cobas | 12 | 9.53 | 0.13 | 5.7 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Roche | 5.83
Roche Cobas | 12 | 7.23 | 0.03 | 4.1 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Roche | 4.13
Roche Cobas | 15 | 5.5 | 0.1 | 5 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Roche | 5.1
Roche Cobas | 15 | 8.33 | 0.03 | 5.8 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Roche | 5.83
Roche Cobas | 15 | 5.7 | 0.05 | 12.2 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Roche | 12.25
Roche Cobas | 15 | 7.09 | -0.02 | 3.3 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Roche | 3.28
Roche Cobas | 15 | 9.44 | 0.12 | 2.9 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Roche | 3.02
Roche Cobas | 15 | 6.03 | -0.04 | 4 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Roche | 3.96
Roche Cobas | 23 | 8.07 | 0.02 | 3 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Roche | 3.02
Roche Cobas | 23 | 5.38 | 0.08 | 4.5 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Roche | 4.58
Roche Cobas | 23 | 6.21 | 0.07 | 4.1 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Roche | 4.17
Roche Cobas | 25 | 6.68 | 0.1 | 2.6 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Roche | 2.7
Roche Cobas | 25 | 8.66 | 0.27 | 2.4 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Roche | 2.67
Roche Cobas | 25 | 5.7 | 0.05 | 4.6 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Roche | 4.65
Roche Cobas | 29 | 6.6 | 0.11 | 3.1 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Roche | 3.21
Roche Cobas | 29 | 7.11 | 0.14 | 2.4 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Roche | 2.54
Roche Cobas | 29 | 9.94 | 0.29 | 2.8 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Roche | 3.09
Roche Cobas | 30 | 6.8 | 0.01 | 2.7 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Roche | 2.71
Roche Cobas | 30 | 10.39 | 0.11 | 2.6 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Roche | 2.71
Roche Cobas | 30 | 6.78 | -0.04 | 2.7 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Roche | 2.66
Roche Cobas | 19 | 8.73 | 0.1 | 2.8 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Roche | 2.9
Roche Cobas | 19 | 5.35 | 0.03 | 4.1 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Roche | 4.13
Roche Cobas | 18 | 5.83 | -0.01 | 3.8 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Roche | 3.79
Roche Cobas | 18 | 11.82 | 0.11 | 3.1 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Roche | 3.21
Roche Cobas | 18 | 9.72 | 0.19 | 2.7 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Roche | 2.89
Roche Cobas | 18 | 5.05 | 0.01 | 3.8 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Roche | 3.81
Roche Cobas | 18 | 7.38 | 0 | 3.4 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Roche | 3.4
Roche Cobas | 18-26 | 11.62 | -0.07 | 3.6 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Roche | 3.53
Roche Cobas | 18-26 | 5.81 | -0.12 | 4.8 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Roche | 4.68
Roche Cobas | 18-26 | 5.07 | -0.1 | 5.8 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Roche | 5.7
Roche Cobas | 18-26 | 8.19 | 0.05 | 2.5 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Roche | 2.55
Roche Cobas | 18-26 | 9.38 | 0.08 | 3.3 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Roche | 3.38
Roche Cobas | 16-25 | 5.15 | -0.17 | 4.6 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Roche | 4.43
Roche Cobas | 16-25 | 9.28 | 0.12 | 3.6 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Roche | 3.72
Roche Cobas | 16-25 | 5.17 | -0.14 | 4.2 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Roche | 4.06
Roche Cobas | 16-25 | 12.11 | 0.08 | 2.6 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Roche | 2.68
Roche Cobas | 16-25 | 5.85 | -0.09 | 2.7 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Roche | 2.61
Roche Cobas | 18 | 5.06 | -0.21 | 3.5 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Roche | 3.29
Roche Cobas | 18 | 10.42 | -0.17 | 3 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Roche | 2.83
Roche Cobas | 18 | 5.94 | -0.26 | 2.5 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Roche | 2.24
Roche Cobas | 18 | 11.95 | -0.28 | 3.1 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Roche | 2.82
Roche Cobas | 18 | 7.38 | -0.13 | 2.7 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Roche | 2.57
Roche Cobas | 16-28 | 9.15 | 0.04 | 2.3 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Roche | 2.34
Roche Cobas | 16-28 | 5.83 | -0.18 | 2.8 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Roche | 2.62
Roche Cobas | 16-28 | 11.52 | -0.19 | 2 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Roche | 1.81
Roche Cobas | 16-28 | 4.75 | -0.27 | 3.1 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Roche | 2.83
Roche Cobas | 16-28 | 7.40 | -0.18 | 1.8 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Roche | 1.62
Roche Cobas | 20-28 | 6.16 | -0.25 | 3.8 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Roche | 3.55
Roche Cobas | 20-28 | 9.44 | -0.09 | 3.3 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Roche | 3.21
Roche Cobas | 20-28 | 5.07 | -0.27 | 5.2 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Roche | 4.93
Roche Cobas | 20-28 | 8.39 | -0.12 | 2.7 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Roche | 2.58
Roche Cobas | 20-28 | 7.15 | -0.1 | 2.5 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Roche | 2.4
Roche Cobas | 20 | 7.21 | -0.21 | 2.2 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Roche | 1.99
Roche Cobas | 20 | 4.87 | -0.33 | 3.4 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Roche | 3.07
Roche Cobas | 20 | 8.19 | -0.12 | 3.3 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Roche | 3.18
Roche Cobas | 20 | 10.32 | -0.01 | 3 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Roche | 2.99
Roche Cobas | 20 | 5.56 | -0.31 | 3.3 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Roche | 2.99
Roche Cobas | 26 | 7.10 | -0.15 | 1.8 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Roche | 1.65
Roche Cobas | 26 | 6.56 | -0.25 | 2.2 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Roche | 1.95
Roche Cobas | 26 | 8.64 | -0.01 | 1.6 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Roche | 1.59
Roche Cobas | 26 | 9.55 | 0.05 | 1.4 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Roche | 1.45
Roche Cobas | 26 | 5.16 | -0.29 | 2.6 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Roche | 2.31
Roche Cobas | 21 | 6.93 | -0.22 | 2.9 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Roche | 2.68
Roche Cobas | 21 | 4.89 | -0.3 | 4 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Roche | 3.7
Roche Cobas | 21 | 8.35 | -0.07 | 2.1 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Roche | 2.03
Roche Cobas | 21 | 9.88 | 0.09 | 1.8 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Roche | 1.89
Roche Cobas | 21 | 5.88 | -0.24 | 3.1 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Roche | 2.86
Roche Cobas | 12 | 6.24 | -0.07 | 1.4 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Roche | 1.33
Roche Cobas | 12 | 9.26 | 0.15 | 1.5 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Roche | 1.65
Roche Cobas | 12 | 5.18 | -0.13 | 2.5 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Roche | 2.37
Roche Cobas | 12 | 8.33 | -0.02 | 1.4 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Roche | 1.38
Roche Cobas | 12 | 7.64 | 0.02 | 1 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Roche | 1.02
Roche Cobas | 18 | 9.36 | -0.01 | 1.7 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Roche | 1.69
Roche Cobas | 18 | 4.77 | -0.15 | 3.1 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Roche | 2.95
Roche Cobas | 18 | 5.91 | -0.13 | 2.3 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Roche | 2.17
Roche Cobas | 18 | 5.18 | -0.1 | 1.1 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Roche | 1
Roche Cobas | 18 | 7.96 | -0.12 | 1.4 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Roche | 1.28
Roche Cobas | 19 | 5.18 | -0.28 | 2.7 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Roche | 2.42
Roche Cobas | 19 | 5.43 | -0.23 | 1.9 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Roche | 1.67
Roche Cobas | 19 | 9.33 | 0.02 | 1.2 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Roche | 1.22
Roche Cobas | 19 | 4.99 | -0.29 | 3.3 | GH5-04 | 5.28 | 2019 | CAP | Fresh | Roche | 3.01
Roche Cobas | 19 | 7.31 | -0.1 | 1.1 | GH5-05 | 7.41 | 2019 | CAP | Fresh | Roche | 1
Roche Cobas | 14 | 6.36 | -0.05 | 1.8 | GH5-06 | 6.41 | 2019 | CAP | Fresh | Roche | 1.75
Roche Cobas | 14 | 8.71 | 0.11 | 1.8 | GH5-07 | 8.6 | 2019 | CAP | Fresh | Roche | 1.91
Roche Cobas | 14 | 5.29 | -0.13 | 2.3 | GH5-08 | 5.42 | 2019 | CAP | Fresh | Roche | 2.17
Roche Cobas | 14 | 7.39 | 0.01 | 2.1 | GH5-09 | 7.38 | 2019 | CAP | Fresh | Roche | 2.11
Roche Cobas | 14 | 9.96 | 0.21 | 2.6 | GH5-10 | 9.75 | 2019 | CAP | Fresh | Roche | 2.81
Roche Cobas | 21 | 5.14 | -0.11 | 2.6 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Roche | 2.49
Roche Cobas | 21 | 6.35 | -0.06 | 2.2 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Roche | 2.14
Roche Cobas | 21 | 8.31 | 0.1 | 2.3 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Roche | 2.4
Roche Cobas | 21 | 5.34 | -0.13 | 2.7 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Roche | 2.57
Roche Cobas | 21 | 9.75 | 0.27 | 1.4 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Roche | 1.67
Roche Cobas | 18 | 8.14 | 0.17 | 2.1 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Roche | 2.27
Roche Cobas | 18 | 5.85 | -0.04 | 2.4 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Roche | 2.36
Roche Cobas | 18 | 5.03 | -0.1 | 3.3 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Roche | 3.2
Roche Cobas | 18 | 7.53 | 0.13 | 2.9 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Roche | 3.03
Roche Cobas | 18 | 9.31 | 0.14 | 1.6 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Roche | 1.74
Roche Cobas | 17 | 7.9 | 0.09 | 2.1 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Roche | 2.19
Roche Cobas | 17 | 5.85 | -0.06 | 2.5 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Roche | 2.44
Roche Cobas | 17 | 8.82 | 0.13 | 2.2 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Roche | 2.33
Roche Cobas | 17 | 5.76 | -0.1 | 3.3 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Roche | 3.2
Roche Cobas | 17 | 9.72 | 0.16 | 1.9 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Roche | 2.06
Roche Cobas | 21 | 5.44 | -0.04 | 2.8 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Roche | 2.76
Roche Cobas | 21 | 8.3 | 0.04 | 2.2 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Roche | 2.24
Roche Cobas | 21 | 5.06 | -0.09 | 3.4 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Roche | 3.31
Roche Cobas | 21 | 10.12 | 0.07 | 2.4 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Roche | 2.47
Roche Cobas | 21 | 5.92 | -0.07 | 2.3 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Roche | 2.23
Roche Cobas | 23 | 8.9 | 0.04 | 2.7 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Roche | 2.74
Roche Cobas | 23 | 6.75 | -0.09 | 2.5 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Roche | 2.41
Roche Cobas | 23 | 9.43 | 0.04 | 3.3 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Roche | 3.34
Roche Cobas | 23 | 6.06 | -0.07 | 2.5 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Roche | 2.43
Roche Cobas | 23 | 7.48 | -0.04 | 2.8 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Roche | 2.76
Roche Cobas | 12 | 5.63 | -0.17 | 1.9 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Roche | 1.73
Roche Cobas | 12 | 9.24 | 0.12 | 1.9 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Roche | 2.02
Roche Cobas | 12 | 5.32 | -0.15 | 2.7 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Roche | 2.55
Roche Cobas | 12 | 7.28 | 0.03 | 1.6 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Roche | 1.63
Roche Cobas | 12 | 9.35 | 0.09 | 1.4 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Roche | 1.49
Roche Cobas | 21 | 4.95 | -0.26 | 3.4 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Roche | 3.14
Roche Cobas | 21 | 6.53 | -0.15 | 2.6 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Roche | 2.45
Roche Cobas | 21 | 8.25 |  | 2.6 | GH5-13 |  | 2021 | CAP | Fresh | Roche | 2.6
Roche Cobas | 21 | 5.83 | -0.15 | 2.7 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Roche | 2.55
Roche Cobas | 21 | 10.21 | 0.09 | 2.2 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Roche | 2.29
Roche Cobas c500 series | 237 | 6.95 | -0.16 | 2.4 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Roche | 2.24
Roche Cobas c500 series | 237 | 9.14 | -0.18 | 2.9 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Roche | 2.72
Roche Cobas c500 series | 237 | 6.06 | -0.01 | 3 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Roche | 2.99
Roche Cobas c500 series | 274 | 7.82 | -0.23 | 2.8 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Roche | 2.57
Roche Cobas c500 series | 274 | 5.36 | 0.06 | 3.1 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Roche | 3.16
Roche Cobas c500 series | 274 | 6.12 | -0.02 | 3.1 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Roche | 3.08
Roche Cobas c500 series | 350 | 6.63 | 0.05 | 2.8 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Roche | 2.85
Roche Cobas c500 series | 350 | 8.39 | 0 | 2.8 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Roche | 2.8
Roche Cobas c500 series | 350 | 5.73 | 0.08 | 2.9 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Roche | 2.98
Roche Cobas c500 series | 316 | 6.52 | 0.03 | 2.5 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Roche | 2.53
Roche Cobas c500 series | 316 | 6.95 | -0.02 | 2.6 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Roche | 2.58
Roche Cobas c500 series | 316 | 9.51 | -0.14 | 2.7 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Roche | 2.56
Roche Cobas c500 series | 358 | 6.77 | -0.02 | 2.7 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Roche | 2.68
Roche Cobas c500 series | 358 | 10.01 | -0.27 | 2.8 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Roche | 2.53
Roche Cobas c500 series | 358 | 6.76 | -0.06 | 2.5 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Roche | 2.44
Roche Cobas c500 series | 293 | 8.51 | -0.12 | 2.9 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Roche | 2.78
Roche Cobas c500 series | 293 | 5.39 | 0.07 | 2.8 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Roche | 2.87
Roche Cobas c500 series | 294 | 5.84 | 0 | 2.7 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Roche | 2.7
Roche Cobas c500 series | 294 | 11.22 | -0.49 | 3.2 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Roche | 2.71
Roche Cobas c500 series | 294 | 9.32 | -0.21 | 3.1 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Roche | 2.89
Roche Cobas c500 series | 294 | 5.11 | 0.07 | 3.5 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Roche | 3.57
Roche Cobas c500 series | 294 | 7.27 | -0.11 | 2.4 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Roche | 2.29
Roche Cobas c500 series | 307-385 | 11.37 | -0.32 | 2.9 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Roche | 2.58
Roche Cobas c500 series | 307-385 | 5.89 | -0.04 | 2.7 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Roche | 2.66
Roche Cobas c500 series | 307-385 | 5.19 | 0.02 | 3.5 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Roche | 3.52
Roche Cobas c500 series | 307-385 | 8.04 | -0.1 | 2.5 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Roche | 2.4
Roche Cobas c500 series | 307-385 | 9.17 | -0.13 | 2.9 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Roche | 2.77
Roche Cobas c500 series | 319-381 | 5.23 | -0.09 | 3.2 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Roche | 3.11
Roche Cobas c500 series | 319-381 | 9.11 | -0.05 | 2.8 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Roche | 2.75
Roche Cobas c500 series | 319-381 | 5.23 | -0.08 | 3.3 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Roche | 3.22
Roche Cobas c500 series | 319-381 | 11.82 | -0.21 | 3 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Roche | 2.79
Roche Cobas c500 series | 319-381 | 5.90 | -0.04 | 2.3 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Roche | 2.26
Roche Cobas c500 series | 321 | 5.19 | -0.08 | 2.8 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Roche | 2.72
Roche Cobas c500 series | 321 | 10.32 | -0.27 | 2.7 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Roche | 2.43
Roche Cobas c500 series | 321 | 6.07 | -0.13 | 2.2 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Roche | 2.07
Roche Cobas c500 series | 321 | 11.86 | -0.37 | 2.7 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Roche | 2.33
Roche Cobas c500 series | 321 | 7.49 | -0.02 | 2.3 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Roche | 2.28
Roche Cobas c500 series | 331-399 | 9.03 | -0.08 | 2.2 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Roche | 2.12
Roche Cobas c500 series | 331-399 | 5.90 | -0.11 | 2.3 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Roche | 2.19
Roche Cobas c500 series | 331-399 | 11.43 | -0.28 | 2.6 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Roche | 2.32
Roche Cobas c500 series | 331-399 | 4.87 | -0.15 | 2.9 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Roche | 2.75
Roche Cobas c500 series | 331-399 | 7.52 | -0.06 | 2.2 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Roche | 2.14
Roche Cobas c500 series | 340-396 | 6.28 | -0.13 | 2.6 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Roche | 2.47
Roche Cobas c500 series | 340-396 | 9.45 | -0.08 | 2.5 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Roche | 2.42
Roche Cobas c500 series | 340-396 | 5.21 | -0.13 | 2.9 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Roche | 2.77
Roche Cobas c500 series | 340-396 | 8.40 | -0.11 | 2.4 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Roche | 2.29
Roche Cobas c500 series | 340-396 | 7.20 | -0.05 | 2.3 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Roche | 2.25
Roche Cobas c500 series | 346 | 7.31 | -0.11 | 2.5 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Roche | 2.39
Roche Cobas c500 series | 346 | 5.00 | -0.2 | 2.9 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Roche | 2.7
Roche Cobas c500 series | 346 | 8.23 | -0.08 | 2.6 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Roche | 2.52
Roche Cobas c500 series | 346 | 10.28 | -0.05 | 2.7 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Roche | 2.65
Roche Cobas c500 series | 346 | 5.70 | -0.17 | 2.9 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Roche | 2.73
Roche Cobas c500 series | 406 | 7.16 | -0.09 | 2.1 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Roche | 2.01
Roche Cobas c500 series | 406 | 6.67 | -0.14 | 2.2 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Roche | 2.06
Roche Cobas c500 series | 406 | 8.58 | -0.07 | 2.2 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Roche | 2.13
Roche Cobas c500 series | 406 | 9.48 | -0.02 | 2.3 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Roche | 2.28
Roche Cobas c500 series | 406 | 5.31 | -0.14 | 2.2 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Roche | 2.06
Roche Cobas c500 series | 427 | 7.05 | -0.1 | 2.5 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Roche | 2.4
Roche Cobas c500 series | 427 | 5.06 | -0.13 | 3.3 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Roche | 3.17
Roche Cobas c500 series | 427 | 8.36 | -0.06 | 2.4 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Roche | 2.34
Roche Cobas c500 series | 427 | 9.77 | -0.02 | 2.6 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Roche | 2.58
Roche Cobas c500 series | 427 | 6.03 | -0.09 | 2.4 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Roche | 2.31
Roche Cobas c500 series | 367 | 6.24 | -0.07 | 2.4 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Roche | 2.33
Roche Cobas c500 series | 367 | 9.13 | 0.02 | 2.6 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Roche | 2.62
Roche Cobas c500 series | 367 | 5.25 | -0.06 | 2.3 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Roche | 2.24
Roche Cobas c500 series | 367 | 8.34 | -0.01 | 2.5 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Roche | 2.49
Roche Cobas c500 series | 367 | 7.63 | 0.01 | 2.2 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Roche | 2.21
Roche Cobas c500 series | 413 | 9.28 | -0.09 | 2.4 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Roche | 2.31
Roche Cobas c500 series | 413 | 4.86 | -0.06 | 2.7 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Roche | 2.64
Roche Cobas c500 series | 413 | 6.00 | -0.04 | 2.3 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Roche | 2.26
Roche Cobas c500 series | 413 | 5.23 | -0.05 | 2.6 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Roche | 2.55
Roche Cobas c500 series | 413 | 8.02 | -0.06 | 2.5 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Roche | 2.44
Roche Cobas c500 series | 416 | 5.26 | -0.2 | 3.2 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Roche | 3
Roche Cobas c500 series | 416 | 5.52 | -0.14 | 2.8 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Roche | 2.66
Roche Cobas c500 series | 416 | 9.20 | -0.11 | 2.4 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Roche | 2.29
Roche Cobas c500 series | 416 | 5.08 | -0.2 | 2.8 | GH5-04 | 5.28 | 2019 | CAP | Fresh | Roche | 2.6
Roche Cobas c500 series | 416 | 7.34 | -0.07 | 2.2 | GH5-05 | 7.41 | 2019 | CAP | Fresh | Roche | 2.13
Roche Cobas c500 series | 364 | 6.38 | -0.03 | 2.3 | GH5-06 | 6.41 | 2019 | CAP | Fresh | Roche | 2.27
Roche Cobas c500 series | 364 | 8.63 | 0.03 | 2.1 | GH5-07 | 8.6 | 2019 | CAP | Fresh | Roche | 2.13
Roche Cobas c500 series | 364 | 5.36 | -0.06 | 2.4 | GH5-08 | 5.42 | 2019 | CAP | Fresh | Roche | 2.34
Roche Cobas c500 series | 364 | 7.41 | 0.03 | 2.4 | GH5-09 | 7.38 | 2019 | CAP | Fresh | Roche | 2.43
Roche Cobas c500 series | 364 | 9.78 | 0.03 | 2.3 | GH5-10 | 9.75 | 2019 | CAP | Fresh | Roche | 2.33
Roche Cobas c500 series | 406 | 5.18 | -0.07 | 2.3 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Roche | 2.23
Roche Cobas c500 series | 406 | 6.42 | 0.01 | 2.4 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Roche | 2.41
Roche Cobas c500 series | 406 | 8.32 | 0.11 | 2.3 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Roche | 2.41
Roche Cobas c500 series | 406 | 5.43 | -0.04 | 2.2 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Roche | 2.16
Roche Cobas c500 series | 406 | 9.57 | 0.09 | 2.3 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Roche | 2.39
Roche Cobas c500 series | 357 | 8.11 | 0.14 | 2 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Roche | 0.14
Roche Cobas c500 series | 357 | 5.91 | 0.02 | 2.3 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Roche | 2.32
Roche Cobas c500 series | 357 | 5.06 | -0.07 | 2.5 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Roche | 2.43
Roche Cobas c500 series | 357 | 7.49 | 0.09 | 2.2 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Roche | 2.29
Roche Cobas c500 series | 357 | 9.22 | 0.05 | 2.5 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Roche | 2.55
Roche Cobas c500 series | 347 | 7.85 | 0.04 | 2 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Roche | 0.04
Roche Cobas c500 series | 347 | 5.9 | -0.01 | 2.2 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Roche | 2.19
Roche Cobas c500 series | 347 | 8.74 | 0.05 | 2.3 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Roche | 2.35
Roche Cobas c500 series | 347 | 5.85 | -0.01 | 2.5 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Roche | 2.49
Roche Cobas c500 series | 347 | 9.58 | 0.02 | 2.5 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Roche | 2.52
Roche Cobas c500 series | 400 | 5.47 | -0.01 | 2.4 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Roche | 2.39
Roche Cobas c500 series | 400 | 8.34 | 0.08 | 2.3 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Roche | 2.38
Roche Cobas c500 series | 400 | 5.13 | -0.02 | 2.7 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Roche | 2.68
Roche Cobas c500 series | 400 | 10.09 | 0.04 | 2.3 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Roche | 2.34
Roche Cobas c500 series | 400 | 5.99 | 0 | 2.4 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Roche | 2.4
Roche Cobas c500 series | 394 | 8.89 | 0.03 | 2.5 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Roche | 2.53
Roche Cobas c500 series | 394 | 6.78 | -0.06 | 2.3 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Roche | 2.24
Roche Cobas c500 series | 394 | 9.41 | 0.02 | 2.6 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Roche | 2.62
Roche Cobas c500 series | 394 | 6.09 | -0.04 | 2 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Roche | -0.04
Roche Cobas c500 series | 394 | 7.48 | -0.04 | 2.3 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Roche | 2.26
Roche Cobas c500 series | 350 | 5.67 | -0.13 | 2.4 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Roche | 2.27
Roche Cobas c500 series | 350 | 9.1 | -0.02 | 2.3 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Roche | 2.28
Roche Cobas c500 series | 350 | 5.39 | -0.08 | 2.2 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Roche | 2.12
Roche Cobas c500 series | 350 | 7.23 | -0.02 | 2.1 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Roche | 2.08
Roche Cobas c500 series | 350 | 9.26 | 0 | 2.4 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Roche | 2.4
Roche Cobas c500 series | 409 | 5.1 | -0.11 | 2.9 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Roche | 2.79
Roche Cobas c500 series | 409 | 6.63 | 0.05 | 2.2 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Roche | 2.25
Roche Cobas c500 series | 409 | 8.25 |  | 2.4 | GH5-13 |  | 2021 | CAP | Fresh | Roche | 2.4
Roche Cobas c500 series | 409 | 5.93 | -0.05 | 2.4 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Roche | 2.35
Roche Cobas c500 series | 409 | 10.1 | -0.02 | 2.6 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Roche | 2.58
Roche Cobas c500/700 | 171 | 8.47 | -0.03 | 3 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Roche | 2.97
Roche Cobas c500/700 | 171 | 5.58 | 0.18 | 3.3 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Roche | 3.48
Roche Cobas c500/700 | 171 | 6.52 | 0.12 | 3 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Roche | 3.12
Roche Cobas c500/700 | 12 | 6.29 | -0.01 | 3.4 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Roche | 3.39
Roche Cobas c500/700 | 12 | 7.68 | 0.08 | 2.2 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Roche | 2.28
Roche Cobas c500/700 | 12 | 9.35 | 0.15 | 4.1 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Roche | 4.25
Roche Cobas c500/700 | 248 | 5.62 | 0.02 | 2.5 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Roche | 2.52
Roche Cobas c500/700 | 248 | 9.26 | -0.14 | 2.6 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Roche | 2.46
Roche Cobas c500/700 | 248 | 7.15 | -0.05 | 2.4 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Roche | 2.35
Roche Cobas c500/700 | 262 | 5.5 | 0.1 | 2.5 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Roche | 2.6
Roche Cobas c500/700 | 262 | 8.12 | -0.18 | 2.5 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Roche | 2.32
Roche Cobas c500/700 | 262 | 5.61 | -0.04 | 2.6 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Roche | 2.56
Roche Cobas c501 | 6 | 5.8 |  |  | GH2-01 | 5.4 | 2007 | CAP | Fresh | Roche | 0
Roche Cobas c501 | 6 | 11.4 |  |  | GH2-02 | 11.6 | 2007 | CAP | Fresh | Roche | 0
Roche Cobas c501 | 6 | 7.5 |  |  | GH2-03 | 7.6 | 2007 | CAP | Fresh | Roche | 0
Roche Cobas c501 | 20 | 9 |  | 3.1 | GH2-04 | 9.2 | 2007 | CAP | Fresh | Roche | 3.1
Roche Cobas c501 | 20 | 11 |  | 3.9 | GH2-05 | 11.1 | 2007 | CAP | Fresh | Roche | 3.9
Roche Cobas c501 | 20 | 6.3 |  | 3.4 | GH2-06 | 6.2 | 2007 | CAP | Fresh | Roche | 3.4
Roche Cobas c501 | 45 | 12.3 |  | 3.7 | GH2-01 | 13 | 2008 | CAP | Fresh | Roche | 3.7
Roche Cobas c501 | 45 | 5.4 |  | 3.2 | GH2-02 | 5.3 | 2008 | CAP | Fresh | Roche | 3.2
Roche Cobas c501 | 45 | 8.6 |  | 3.9 | GH2-03 | 8.8 | 2008 | CAP | Fresh | Roche | 3.9
Roche Cobas c501 | 73 | 11.4 |  | 3.6 | GH2-04 | 11.9 | 2008 | CAP | Fresh | Roche | 3.6
Roche Cobas c501 | 73 | 6.7 |  | 3.2 | GH2-05 | 6.8 | 2008 | CAP | Fresh | Roche | 3.2
Roche Cobas c501 | 73 | 5.3 |  | 2.6 | GH2-06 | 5 | 2008 | CAP | Fresh | Roche | 2.6
Roche Cobas c501 | 108 | 5.3 |  | 3.3 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Roche | 3.3
Roche Cobas c501 | 108 | 8.2 |  | 3.6 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Roche | 3.6
Roche Cobas c501 | 108 | 6 |  | 3.4 | GH2-03 | 6 | 2009 | CAP | Fresh | Roche | 3.4
Roche Cobas c501 | 133 | 6.7 | 0.1 | 3 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Roche | 3.1
Roche Cobas c501 | 133 | 9.29 | -0.21 | 3.7 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Roche | 3.49
Roche Cobas c501 | 133 | 7.35 | -0.05 | 3.7 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Roche | 3.65
Roche Cobas c501 | 133 | 6.07 | 0.17 | 2.8 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Roche | 2.97
Roche Cobas c501 | 133 | 9.51 | -0.29 | 3 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Roche | 2.71
Roche Cobas c501 | 133 | 7.4 | 0 | 2.8 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Roche | 2.8
Roche Cobas c501 | 161 | 5.4 | 0.2 | 3.3 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Roche | 3.5
Roche Cobas c501 | 161 | 8.55 | -0.15 | 3.1 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Roche | 2.95
Roche Cobas c501 | 161 | 6.37 | 0.07 | 2.9 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Roche | 2.97
Roche Cobas c513 | 13 | 7.43 | 0.01 | 1.7 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Roche | 1.71
Roche Cobas c513 | 13 | 5.12 | -0.08 | 2.3 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Roche | 2.22
Roche Cobas c513 | 13 | 8.36 | 0.05 | 1.7 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Roche | 1.75
Roche Cobas c513 | 13 | 10.60 | 0.27 | 1.8 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Roche | 2.07
Roche Cobas c513 | 13 | 5.78 | -0.09 | 2.2 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Roche | 2.11
Roche Cobas c513 | 34 | 7.24 | -0.01 | 1.5 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Roche | 1.49
Roche Cobas c513 | 34 | 6.78 | -0.03 | 1.4 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Roche | 1.37
Roche Cobas c513 | 34 | 8.69 | 0.04 | 1.6 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Roche | 1.64
Roche Cobas c513 | 34 | 9.67 | 0.17 | 1.5 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Roche | 1.67
Roche Cobas c513 | 34 | 5.42 | -0.03 | 1.3 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Roche | 1.27
Roche Cobas c513 | 44 | 7.15 | 0 | 2.3 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Roche | 2.3
Roche Cobas c513 | 44 | 5.19 | 0 | 2.2 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Roche | 2.2
Roche Cobas c513 | 44 | 8.46 | 0.04 | 1.8 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Roche | 1.84
Roche Cobas c513 | 44 | 10.03 | 0.24 | 1.8 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Roche | 2.04
Roche Cobas c513 | 44 | 6.10 | -0.02 | 1.6 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Roche | 1.58
Roche Cobas c513 | 47 | 6.29 | -0.02 | 1.5 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Roche | 1.48
Roche Cobas c513 | 47 | 9.32 | 0.21 | 1.7 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Roche | 1.91
Roche Cobas c513 | 47 | 5.28 | -0.03 | 1.4 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Roche | 1.37
Roche Cobas c513 | 47 | 8.45 | 0.1 | 1.8 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Roche | 1.9
Roche Cobas c513 | 47 | 7.70 | 0.08 | 1.5 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Roche | 1.58
Roche Cobas c513 | 55 | 9.39 | 0.02 | 2 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Roche | 2.02
Roche Cobas c513 | 55 | 4.93 | 0.01 | 2.1 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Roche | 2.11
Roche Cobas c513 | 55 | 6.04 | 0 | 1.6 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Roche | 1.6
Roche Cobas c513 | 55 | 5.28 | 0 | 2 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Roche | 2
Roche Cobas c513 | 55 | 8.04 | -0.04 | 1.7 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Roche | 1.66
Roche Cobas c513 | 63 | 5.42 | -0.04 | 1.9 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Roche | 1.86
Roche Cobas c513 | 63 | 5.64 | -0.02 | 1.8 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Roche | 1.78
Roche Cobas c513 | 63 | 9.53 | 0.22 | 1.8 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Roche | 2.02
Roche Cobas c513 | 63 | 5.23 | -0.05 | 1.8 | GH5-04 | 5.28 | 2019 | CAP | Fresh | Roche | 1.75
Roche Cobas c513 | 63 | 7.45 | 0.04 | 1.6 | GH5-05 | 7.41 | 2019 | CAP | Fresh | Roche | 1.64
Roche Cobas c513 | 54 | 6.43 | 0.02 | 1.3 | GH5-06 | 6.41 | 2019 | CAP | Fresh | Roche | 1.32
Roche Cobas c513 | 54 | 8.66 | 0.06 | 1.6 | GH5-07 | 8.6 | 2019 | CAP | Fresh | Roche | 1.66
Roche Cobas c513 | 54 | 5.42 | 0 | 1.4 | GH5-08 | 5.42 | 2019 | CAP | Fresh | Roche | 1.4
Roche Cobas c513 | 54 | 7.40 | 0.02 | 1.7 | GH5-09 | 7.38 | 2019 | CAP | Fresh | Roche | 1.72
Roche Cobas c513 | 54 | 9.90 | 0.15 | 1.8 | GH5-10 | 9.75 | 2019 | CAP | Fresh | Roche | 1.95
Roche Cobas c513 | 59 | 5.21 | -0.04 | 1.6 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Roche | 1.56
Roche Cobas c513 | 59 | 6.41 | 0 | 1.7 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Roche | 1.7
Roche Cobas c513 | 59 | 8.24 | 0.03 | 1.9 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Roche | 1.93
Roche Cobas c513 | 59 | 5.43 | -0.04 | 1.5 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Roche | 1.46
Roche Cobas c513 | 59 | 9.57 | 0.09 | 1.9 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Roche | 1.99
Roche Cobas c513 | 56 | 8.02 | 0.05 | 1.7 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Roche | 1.75
Roche Cobas c513 | 56 | 5.88 | -0.01 | 1.3 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Roche | 1.29
Roche Cobas c513 | 56 | 5.1 | -0.03 | 1.5 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Roche | 1.47
Roche Cobas c513 | 56 | 7.39 | -0.01 | 1.5 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Roche | 1.49
Roche Cobas c513 | 56 | 9.21 | 0.04 | 1.6 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Roche | 1.64
Roche Cobas c513 | 58 | 7.79 | -0.02 | 1.9 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Roche | 1.88
Roche Cobas c513 | 58 | 5.86 | -0.05 | 1.7 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Roche | 1.65
Roche Cobas c513 | 58 | 8.7 | 0.01 | 1.8 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Roche | 1.81
Roche Cobas c513 | 58 | 5.82 | -0.04 | 1.6 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Roche | 1.56
Roche Cobas c513 | 58 | 9.58 | 0.02 | 2 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Roche | 0.02
Roche Cobas c513 | 60 | 5.45 | -0.03 | 1.9 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Roche | 1.87
Roche Cobas c513 | 60 | 8.29 | 0.03 | 1.8 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Roche | 1.83
Roche Cobas c513 | 60 | 5.13 | -0.02 | 1.8 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Roche | 1.78
Roche Cobas c513 | 60 | 10.18 | 0.13 | 2.2 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Roche | 2.33
Roche Cobas c513 | 60 | 5.95 | -0.04 | 1.9 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Roche | 1.86
Roche Cobas c513 | 63 | 8.95 | 0.09 | 1.8 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Roche | 1.89
Roche Cobas c513 | 63 | 6.81 | -0.03 | 1.5 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Roche | 1.47
Roche Cobas c513 | 63 | 9.54 | 0.15 | 1.7 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Roche | 1.85
Roche Cobas c513 | 63 | 6.13 | 0 | 1.4 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Roche | 1.4
Roche Cobas c513 | 63 | 7.51 | -0.01 | 1.5 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Roche | 1.49
Roche Cobas c513 | 61 | 5.82 | 0.02 | 1.7 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Roche | 1.72
Roche Cobas c513 | 61 | 9.26 | 0.14 | 1.6 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Roche | 1.74
Roche Cobas c513 | 61 | 5.5 | 0.03 | 1.3 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Roche | 1.33
Roche Cobas c513 | 61 | 7.32 | 0.07 | 1.3 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Roche | 1.37
Roche Cobas c513 | 61 | 9.38 | 0.12 | 1.6 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Roche | 1.72
Roche Cobas c513 | 64 | 5.27 | 0.06 | 1.5 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Roche | 1.56
Roche Cobas c513 | 64 | 6.73 | 0.05 | 1.5 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Roche | 1.55
Roche Cobas c513 | 64 | 8.37 |  | 2.1 | GH5-13 |  | 2021 | CAP | Fresh | Roche | 2.1
Roche Cobas c513 | 64 | 6.04 | 0.06 | 1.4 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Roche | 1.46
Roche Cobas c513 | 64 | 10.41 | 0.29 | 1.5 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Roche | 1.79
Roche Cobas Integra | 248 | 11.1 |  | 4 | GH2-01 | 10.7 | 2006 | CAP | Fresh | Roche | 4
Roche Cobas Integra | 248 | 5.5 |  | 3.6 | GH2-02 | 5.30 | 2006 | CAP | Fresh | Roche | 3.6
Roche Cobas Integra | 248 | 8.7 |  | 3.7 | GH2-03 | 8.40 | 2006 | CAP | Fresh | Roche | 3.7
Roche Cobas Integra | 253 | 7.1 |  | 3.7 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Roche | 3.7
Roche Cobas Integra | 253 | 7.1 |  | 3.7 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Roche | 3.7
Roche Cobas Integra | 253 | 12.1 |  | 3.9 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Roche | 3.9
Roche Cobas Integra | 253 | 7.1 |  | 3.7 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Roche | 3.7
Roche Cobas Integra | 253 | 7.1 |  | 3.7 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Roche | 3.7
Roche Cobas Integra | 253 | 12.1 |  | 3.9 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Roche | 3.9
Roche Cobas Integra | 239 | 5.5 |  | 3.4 | GH2-01 | 5.4 | 2007 | CAP | Fresh | Roche | 3.4
Roche Cobas Integra | 239 | 11.8 |  | 3.7 | GH2-02 | 11.6 | 2007 | CAP | Fresh | Roche | 3.7
Roche Cobas Integra | 239 | 7.7 |  | 3.6 | GH2-03 | 7.6 | 2007 | CAP | Fresh | Roche | 3.6
Roche Cobas Integra | 177 | 9.2 |  | 4.3 | GH2-04 | 9.2 | 2007 | CAP | Fresh | Roche | 4.3
Roche Cobas Integra | 177 | 11.1 |  | 4 | GH2-05 | 11.1 | 2007 | CAP | Fresh | Roche | 4
Roche Cobas Integra | 177 | 6.3 |  | 3.2 | GH2-06 | 6.2 | 2007 | CAP | Fresh | Roche | 3.2
Roche Cobas Integra | 146 | 12.7 |  | 3.9 | GH2-01 | 13 | 2008 | CAP | Fresh | Roche | 3.9
Roche Cobas Integra | 146 | 5.4 |  | 3 | GH2-02 | 5.3 | 2008 | CAP | Fresh | Roche | 3
Roche Cobas Integra | 146 | 8.9 |  | 3.1 | GH2-03 | 8.8 | 2008 | CAP | Fresh | Roche | 3.1
Roche Cobas Integra | 121 | 11.7 |  | 3.4 | GH2-04 | 11.9 | 2008 | CAP | Fresh | Roche | 3.4
Roche Cobas Integra | 121 | 6.8 |  | 3.3 | GH2-05 | 6.8 | 2008 | CAP | Fresh | Roche | 3.3
Roche Cobas Integra | 121 | 5.2 |  | 3 | GH2-06 | 5 | 2008 | CAP | Fresh | Roche | 3
Roche Cobas Integra 400 | 38 | 6.87 | 0.27 | 3.6 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Roche | 3.87
Roche Cobas Integra 400 | 38 | 9.75 | 0.25 | 2.4 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Roche | 2.65
Roche Cobas Integra 400 | 38 | 7.71 | 0.31 | 3.3 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Roche | 3.61
Roche Cobas Integra 400 | 42 | 6.1 | 0.2 | 2.5 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Roche | 2.7
Roche Cobas Integra 400 | 42 | 9.89 | 0.09 | 2.5 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Roche | 2.59
Roche Cobas Integra 400 | 42 | 7.68 | 0.28 | 2.8 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Roche | 3.08
Roche Cobas Integra 400 | 47 | 5.36 | 0.16 | 2.9 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Roche | 3.06
Roche Cobas Integra 400 | 47 | 8.84 | 0.14 | 2.5 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Roche | 2.64
Roche Cobas Integra 400 | 47 | 6.5 | 0.2 | 2.9 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Roche | 3.1
Roche Cobas Integra 400 | 52 | 8.68 | 0.18 | 3.8 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Roche | 3.98
Roche Cobas Integra 400 | 52 | 5.58 | 0.18 | 3.5 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Roche | 3.68
Roche Cobas Integra 400 | 52 | 6.63 | 0.23 | 3.9 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Roche | 4.13
Roche Cobas Integra 400 | 210 | 6.23 | -0.07 | 2.9 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Roche | 2.83
Roche Cobas Integra 400 | 210 | 7.50 | -0.1 | 2.5 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Roche | 2.4
Roche Cobas Integra 400 | 210 | 9.07 | -0.13 | 2.8 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Roche | 2.67
Roche Cobas Integra 400 | 53 | 5.6 | 0 | 2.8 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Roche | 2.8
Roche Cobas Integra 400 | 53 | 9.5 | 0.1 | 2.8 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Roche | 2.9
Roche Cobas Integra 400 | 53 | 7.33 | 0.13 | 2.4 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Roche | 2.53
Roche Cobas Integra 400 | 51 | 5.47 | 0.07 | 3.2 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Roche | 3.27
Roche Cobas Integra 400 | 51 | 8.4 | 0.1 | 3.8 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Roche | 3.9
Roche Cobas Integra 400 | 51 | 5.65 | 0 | 3.3 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Roche | 3.3
Roche Cobas Integra 400 | 43 | 7.18 | 0.07 | 3.7 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Roche | 3.77
Roche Cobas Integra 400 | 43 | 9.54 | 0.22 | 3.9 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Roche | 4.12
Roche Cobas Integra 400 | 43 | 6.15 | 0.08 | 3.9 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Roche | 3.98
Roche Cobas Integra 400 | 41 | 7.96 | -0.09 | 3.4 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Roche | 3.31
Roche Cobas Integra 400 | 41 | 5.29 | -0.01 | 3.1 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Roche | 3.09
Roche Cobas Integra 400 | 41 | 6.08 | -0.06 | 3 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Roche | 2.94
Roche Cobas Integra 400 | 60 | 6.63 | 0.05 | 4.3 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Roche | 4.35
Roche Cobas Integra 400 | 60 | 8.55 | 0.16 | 4.2 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Roche | 4.36
Roche Cobas Integra 400 | 60 | 5.65 | 0 | 4.1 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Roche | 4.1
Roche Cobas Integra 400 | 50 | 6.46 | -0.03 | 2.9 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Roche | 2.87
Roche Cobas Integra 400 | 50 | 6.95 | -0.02 | 2.9 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Roche | 2.88
Roche Cobas Integra 400 | 50 | 9.73 | 0.08 | 3 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Roche | 3.08
Roche Cobas Integra 400 | 58 | 6.78 | -0.01 | 3.2 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Roche | 3.19
Roche Cobas Integra 400 | 58 | 10.09 | -0.19 | 3.7 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Roche | 3.51
Roche Cobas Integra 400 | 58 | 6.83 | 0.01 | 3.3 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Roche | 3.31
Roche Cobas Integra 400 | 24 | 8.71 | 0.08 | 3.3 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Roche | 3.38
Roche Cobas Integra 400 | 24 | 5.36 | 0.04 | 3.4 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Roche | 3.44
Roche Cobas Integra 400 | 26 | 5.77 | -0.07 | 3.2 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Roche | 3.13
Roche Cobas Integra 400 | 26 | 11.26 | -0.45 | 3.4 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Roche | 2.95
Roche Cobas Integra 400 | 26 | 9.52 | -0.01 | 3.1 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Roche | 3.09
Roche Cobas Integra 400 | 26 | 5.06 | 0.02 | 2.8 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Roche | 2.82
Roche Cobas Integra 400 | 26 | 7.4 | 0.02 | 3.9 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Roche | 3.92
Roche Cobas Integra 400 | 25-64 | 11.21 | -0.48 | 4.5 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Roche | 4.02
Roche Cobas Integra 400 | 25-64 | 5.91 | -0.02 | 3.3 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Roche | 3.28
Roche Cobas Integra 400 | 25-64 | 5.22 | 0.05 | 3.7 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Roche | 3.75
Roche Cobas Integra 400 | 25-64 | 8.28 | 0.14 | 3.5 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Roche | 3.64
Roche Cobas Integra 400 | 25-64 | 9.36 | 0.06 | 3.3 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Roche | 3.36
Roche Cobas Integra 400 | 31-57 | 5.23 | -0.09 | 3.9 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Roche | 3.81
Roche Cobas Integra 400 | 31-57 | 9.16 | 0 | 3.4 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Roche | 3.4
Roche Cobas Integra 400 | 31-57 | 5.25 | -0.06 | 4 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Roche | 3.94
Roche Cobas Integra 400 | 31-57 | 11.85 | -0.18 | 3 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Roche | 2.82
Roche Cobas Integra 400 | 31-57 | 5.88 | -0.06 | 2.5 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Roche | 2.44
Roche Cobas Integra 400 | 34 | 5.23 | -0.04 | 3.3 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Roche | 3.26
Roche Cobas Integra 400 | 34 | 10.50 | -0.09 | 2.5 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Roche | 2.41
Roche Cobas Integra 400 | 34 | 6.06 | -0.14 | 3.6 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Roche | 3.46
Roche Cobas Integra 400 | 34 | 12.06 | -0.17 | 3 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Roche | 2.83
Roche Cobas Integra 400 | 34 | 7.45 | -0.06 | 2.6 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Roche | 2.54
Roche Cobas Integra 400 | 30-60 | 9.16 | 0.05 | 3.1 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Roche | 3.15
Roche Cobas Integra 400 | 30-60 | 5.99 | -0.02 | 2.8 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Roche | 2.78
Roche Cobas Integra 400 | 30-60 | 11.64 | -0.07 | 3.5 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Roche | 3.43
Roche Cobas Integra 400 | 30-60 | 5.16 | 0.14 | 3.6 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Roche | 3.74
Roche Cobas Integra 400 | 30-60 | 7.61 | 0.03 | 2.8 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Roche | 2.83
Roche Cobas Integra 400 | 28-54 | 6.42 | 0.01 | 3.2 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Roche | 3.21
Roche Cobas Integra 400 | 28-54 | 9.62 | 0.09 | 2.9 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Roche | 2.99
Roche Cobas Integra 400 | 28-54 | 5.43 | 0.09 | 3.4 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Roche | 3.49
Roche Cobas Integra 400 | 28-54 | 8.66 | 0.15 | 3.8 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Roche | 3.95
Roche Cobas Integra 400 | 28-54 | 7.39 | 0.14 | 3.6 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Roche | 3.74
Roche Cobas Integra 400 | 25 | 7.46 | 0.04 | 1.7 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Roche | 1.74
Roche Cobas Integra 400 | 25 | 5.27 | 0.07 | 3.5 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Roche | 3.57
Roche Cobas Integra 400 | 25 | 8.42 | 0.11 | 2.2 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Roche | 2.31
Roche Cobas Integra 400 | 25 | 10.55 | 0.22 | 2.3 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Roche | 2.52
Roche Cobas Integra 400 | 25 | 5.93 | 0.06 | 2.5 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Roche | 2.56
Roche Cobas Integra 400 | 50 | 7.22 | -0.03 | 2.8 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Roche | 2.77
Roche Cobas Integra 400 | 50 | 6.73 | -0.08 | 2.7 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Roche | 2.62
Roche Cobas Integra 400 | 50 | 8.77 | 0.12 | 2.8 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Roche | 2.92
Roche Cobas Integra 400 | 50 | 9.94 | 0.44 | 3.5 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Roche | 3.94
Roche Cobas Integra 400 | 50 | 5.51 | 0.06 | 3.4 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Roche | 3.46
Roche Cobas Integra 400 | 47 | 6.96 | -0.19 | 2.2 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Roche | 2.01
Roche Cobas Integra 400 | 47 | 5.19 | 0 | 2.9 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Roche | 2.9
Roche Cobas Integra 400 | 47 | 8.37 | -0.05 | 1.9 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Roche | 1.85
Roche Cobas Integra 400 | 47 | 9.94 | 0.15 | 2.2 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Roche | 2.35
Roche Cobas Integra 400 | 47 | 6.07 | -0.05 | 2.9 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Roche | 2.85
Roche Cobas Integra 400 | 19 | 6.12 | -0.19 | 2.9 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Roche | 2.71
Roche Cobas Integra 400 | 19 | 9.13 | 0.02 | 2.6 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Roche | 2.62
Roche Cobas Integra 400 | 19 | 5.21 | -0.1 | 2.8 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Roche | 2.7
Roche Cobas Integra 400 | 19 | 8.31 | -0.04 | 2.3 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Roche | 2.26
Roche Cobas Integra 400 | 19 | 7.64 | 0.02 | 2.6 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Roche | 2.62
Roche Cobas Integra 400 | 42 | 9.32 | -0.05 | 2.2 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Roche | 2.15
Roche Cobas Integra 400 | 42 | 4.94 | 0.02 | 2.4 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Roche | 2.42
Roche Cobas Integra 400 | 42 | 6.01 | -0.03 | 2.4 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Roche | 2.37
Roche Cobas Integra 400 | 42 | 5.29 | 0.01 | 2.4 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Roche | 2.41
Roche Cobas Integra 400 | 42 | 7.93 | -0.15 | 1.6 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Roche | 1.45
Roche Cobas Integra 400 | 44 | 5.40 | -0.06 | 4.5 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Roche | 4.44
Roche Cobas Integra 400 | 44 | 5.64 | -0.02 | 3.4 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Roche | 3.38
Roche Cobas Integra 400 | 44 | 9.32 | 0.01 | 2.6 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Roche | 2.61
Roche Cobas Integra 400 | 44 | 5.28 | 0 | 3.3 | GH5-04 | 5.28 | 2019 | CAP | Fresh | Roche | 3.3
Roche Cobas Integra 400 | 44 | 7.35 | -0.06 | 3.3 | GH5-05 | 7.41 | 2019 | CAP | Fresh | Roche | 3.24
Roche Cobas Integra 400 | 23 | 6.37 | -0.04 | 2.2 | GH5-06 | 6.41 | 2019 | CAP | Fresh | Roche | 2.16
Roche Cobas Integra 400 | 23 | 8.72 | 0.12 | 2.3 | GH5-07 | 8.6 | 2019 | CAP | Fresh | Roche | 2.42
Roche Cobas Integra 400 | 23 | 5.35 | -0.07 | 3.2 | GH5-08 | 5.42 | 2019 | CAP | Fresh | Roche | 3.13
Roche Cobas Integra 400 | 23 | 7.47 | 0.09 | 2.5 | GH5-09 | 7.38 | 2019 | CAP | Fresh | Roche | 2.59
Roche Cobas Integra 400 | 23 | 9.90 | 0.15 | 2.4 | GH5-10 | 9.75 | 2019 | CAP | Fresh | Roche | 2.55
Roche Cobas Integra 400 | 41 | 5.17 | -0.08 | 2.4 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Roche | 2.32
Roche Cobas Integra 400 | 41 | 6.36 | -0.05 | 1.9 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Roche | 1.85
Roche Cobas Integra 400 | 41 | 8.33 | 0.12 | 2.1 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Roche | 2.22
Roche Cobas Integra 400 | 41 | 5.46 | -0.01 | 1.9 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Roche | 1.89
Roche Cobas Integra 400 | 41 | 9.54 | 0.06 | 1.6 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Roche | 1.66
Roche Cobas Integra 400 | 16 | 8.06 | 0.09 | 2.7 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Roche | 2.79
Roche Cobas Integra 400 | 16 | 5.83 | -0.06 | 2.4 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Roche | 2.34
Roche Cobas Integra 400 | 16 | 5.15 | 0.02 | 2.1 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Roche | 2.12
Roche Cobas Integra 400 | 16 |  |  |  | GH5-04 | 7.4 | 2020 | CAP | Fresh | Roche | 0
Roche Cobas Integra 400 | 16 |  |  |  | GH5-05 | 9.17 | 2020 | CAP | Fresh | Roche | 0
Roche Cobas Integra 400 | 19 | 7.85 | 0.04 | 2.4 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Roche | 2.44
Roche Cobas Integra 400 | 19 | 5.86 | -0.05 | 3 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Roche | -0.05
Roche Cobas Integra 400 | 19 | 8.78 | 0.09 | 2.2 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Roche | 2.29
Roche Cobas Integra 400 | 19 | 5.85 | -0.01 | 3 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Roche | -0.01
Roche Cobas Integra 400 | 19 | 9.66 | 0.1 | 2 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Roche | 0.1
Roche Cobas Integra 400 | 36 | 5.47 | -0.01 | 2.6 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Roche | 2.59
Roche Cobas Integra 400 | 36 | 8.41 | 0.15 | 2.2 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Roche | 2.35
Roche Cobas Integra 400 | 36 | 5.19 | 0.04 | 2.4 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Roche | 2.44
Roche Cobas Integra 400 | 36 | 10.3 | 0.25 | 2.3 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Roche | 2.55
Roche Cobas Integra 400 | 36 | 6.01 | 0.02 | 2.5 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Roche | 2.52
Roche Cobas Integra 400 | 36 | 8.85 | -0.01 | 3.4 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Roche | 3.39
Roche Cobas Integra 400 | 36 | 6.76 | -0.08 | 3.2 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Roche | 3.12
Roche Cobas Integra 400 | 36 | 9.48 | 0.09 | 2.8 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Roche | 2.89
Roche Cobas Integra 400 | 36 | 6.07 | -0.06 | 3 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Roche | -0.06
Roche Cobas Integra 400 | 36 | 7.38 | -0.14 | 4.2 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Roche | 4.06
Roche Cobas Integra 400 | 18 | 5.75 | -0.05 | 3.2 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Roche | 3.15
Roche Cobas Integra 400 | 18 | 9.27 | 0.15 | 2.3 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Roche | 2.45
Roche Cobas Integra 400 | 18 | 5.45 | -0.02 | 2.5 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Roche | 2.48
Roche Cobas Integra 400 | 18 | 7.34 | 0.09 | 2.5 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Roche | 2.59
Roche Cobas Integra 400 | 18 | 9.43 | 0.17 | 2.4 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Roche | 2.57
Roche Cobas Integra 400 | 32 | 5.18 | -0.03 | 2.6 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Roche | 2.57
Roche Cobas Integra 400 | 32 | 6.65 | -0.03 | 2.6 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Roche | 2.57
Roche Cobas Integra 400 | 32 | 8.5 |  | 2.2 | GH5-13 |  | 2021 | CAP | Fresh | Roche | 2.2
Roche Cobas Integra 400 | 32 | 5.98 | 0 | 2.2 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Roche | 2.2
Roche Cobas Integra 400 | 32 | 10.31 | 0.19 | 2.2 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Roche | 2.39
Roche Cobas Integra 400 (non TQ) | 20 | 5.3 |  | 2.9 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Roche | 2.9
Roche Cobas Integra 400 (non TQ) | 20 | 8.7 |  | 3.3 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Roche | 3.3
Roche Cobas Integra 400 (non TQ) | 20 | 6.2 |  | 3.3 | GH2-03 | 6 | 2009 | CAP | Fresh | Roche | 3.3
Roche Cobas Integra 400 Gen.2 | 26 | 5.2 |  | 4.1 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Roche | 4.1
Roche Cobas Integra 400 Gen.2 | 26 | 8.5 |  | 2.4 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Roche | 2.4
Roche Cobas Integra 400 Gen.2 | 26 | 6.1 |  | 3.6 | GH2-03 | 6 | 2009 | CAP | Fresh | Roche | 3.6
Roche Cobas Integra 800 | 139 | 6.78 | 0.18 | 3.2 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Roche | 3.38
Roche Cobas Integra 800 | 139 | 9.59 | 0.09 | 2.8 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Roche | 2.89
Roche Cobas Integra 800 | 139 | 7.55 | 0.15 | 3.3 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Roche | 3.45
Roche Cobas Integra 800 | 115 | 6.07 | 0.17 | 3 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Roche | 3.17
Roche Cobas Integra 800 | 115 | 9.8 | 0 | 2.5 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Roche | 2.5
Roche Cobas Integra 800 | 115 | 7.55 | 0.15 | 2.5 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Roche | 2.65
Roche Cobas Integra 800 | 118 | 5.39 | 0.19 | 3 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Roche | 3.19
Roche Cobas Integra 800 | 118 | 8.74 | 0.04 | 2.5 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Roche | 2.54
Roche Cobas Integra 800 | 118 | 6.41 | 0.11 | 2.6 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Roche | 2.71
Roche Cobas Integra 800 | 111 | 8.55 | 0.05 | 3.3 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Roche | 3.35
Roche Cobas Integra 800 | 111 | 5.61 | 0.21 | 3 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Roche | 3.21
Roche Cobas Integra 800 | 111 | 6.57 | 0.17 | 3.1 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Roche | 3.27
Roche Cobas Integra 800 | 47 | 6.33 | 0.03 | 2.6 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Roche | 2.63
Roche Cobas Integra 800 | 47 | 7.71 | 0.11 | 2.8 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Roche | 2.91
Roche Cobas Integra 800 | 47 | 9.34 | 0.14 | 2.7 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Roche | 2.84
Roche Cobas Integra 800 | 144 | 5.65 | 0.05 | 2.1 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Roche | 2.15
Roche Cobas Integra 800 | 144 | 9.51 | 0.11 | 2.3 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Roche | 2.41
Roche Cobas Integra 800 | 144 | 7.24 | 0.04 | 2.2 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Roche | 2.24
Roche Cobas Integra 800 | 145 | 5.51 | 0.11 | 2.5 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Roche | 2.61
Roche Cobas Integra 800 | 145 | 8.26 | -0.04 | 2 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Roche | 1.96
Roche Cobas Integra 800 | 145 | 5.65 | 0 | 2.4 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Roche | 2.4
Roche Cobas Integra 800 | 118 | 7.03 | -0.08 | 1.8 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Roche | 1.72
Roche Cobas Integra 800 | 118 | 9.35 | 0.03 | 2 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Roche | 2.03
Roche Cobas Integra 800 | 118 | 6.08 | 0.01 | 2.1 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Roche | 2.11
Roche Cobas Integra 800 | 116 | 7.91 | -0.14 | 2.5 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Roche | 2.36
Roche Cobas Integra 800 | 116 | 5.39 | 0.09 | 2.7 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Roche | 2.79
Roche Cobas Integra 800 | 116 | 6.18 | 0.04 | 1.9 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Roche | 1.94
Roche Cobas Integra 800 | 135 | 6.68 | 0.1 | 2.8 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Roche | 2.9
Roche Cobas Integra 800 | 135 | 8.63 | 0.24 | 2.3 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Roche | 2.54
Roche Cobas Integra 800 | 135 | 5.76 | 0.11 | 2.6 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Roche | 2.71
Roche Cobas Integra 800 | 136 | 6.57 | 0.08 | 2.2 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Roche | 2.28
Roche Cobas Integra 800 | 136 | 7.05 | 0.08 | 2.2 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Roche | 2.28
Roche Cobas Integra 800 | 136 | 9.79 | 0.14 | 2.1 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Roche | 2.24
Roche Cobas Integra 800 | 125 | 6.8 | 0.01 | 2.7 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Roche | 2.71
Roche Cobas Integra 800 | 125 | 10.28 | 0 | 2.6 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Roche | 2.6
Roche Cobas Integra 800 | 125 | 6.81 | -0.01 | 2.2 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Roche | 2.19
Roche Cobas Integra 800 | 110 | 8.67 | 0.04 | 2.3 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Roche | 2.34
Roche Cobas Integra 800 | 110 | 5.48 | 0.16 | 2.5 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Roche | 2.66
Roche Cobas Integra 800 | 108 | 5.89 | 0.05 | 2.4 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Roche | 2.45
Roche Cobas Integra 800 | 108 | 11.75 | 0.04 | 2.2 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Roche | 2.24
Roche Cobas Integra 800 | 108 | 9.6 | 0.07 | 2.1 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Roche | 2.17
Roche Cobas Integra 800 | 108 | 5.23 | 0.19 | 2.4 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Roche | 2.59
Roche Cobas Integra 800 | 108 | 7.36 | -0.02 | 2.2 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Roche | 2.18
Roche Cobas Integra 800 | 106-120 | 11.76 | 0.07 | 2.5 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Roche | 2.57
Roche Cobas Integra 800 | 106-120 | 6.02 | 0.09 | 2.5 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Roche | 2.59
Roche Cobas Integra 800 | 106-120 | 5.36 | 0.19 | 2.6 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Roche | 2.79
Roche Cobas Integra 800 | 106-120 | 8.17 | 0.03 | 2.2 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Roche | 2.23
Roche Cobas Integra 800 | 106-120 | 9.24 | -0.06 | 2.5 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Roche | 2.44
Roche Cobas Integra 800 | 105-120 | 5.40 | 0.08 | 2.6 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Roche | 2.68
Roche Cobas Integra 800 | 105-120 | 9.20 | 0.04 | 2.2 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Roche | 2.24
Roche Cobas Integra 800 | 105-120 | 5.41 | 0.1 | 2.2 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Roche | 2.3
Roche Cobas Integra 800 | 105-120 | 12.11 | 0.08 | 2.3 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Roche | 2.38
Roche Cobas Integra 800 | 105-120 | 6.00 | 0.06 | 2.3 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Roche | 2.36
Roche Cobas Integra 800 | 104 | 5.37 | 0.1 | 2.6 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Roche | 2.7
Roche Cobas Integra 800 | 104 | 10.64 | 0.05 | 1.7 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Roche | 1.75
Roche Cobas Integra 800 | 104 | 6.19 | -0.01 | 2.3 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Roche | 2.29
Roche Cobas Integra 800 | 104 | 12.27 | 0.04 | 1.9 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Roche | 1.94
Roche Cobas Integra 800 | 104 | 7.49 | -0.02 | 2.3 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Roche | 2.28
Roche Cobas Integra 800 | 104-118 | 9.14 | 0.03 | 2.3 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Roche | 2.33
Roche Cobas Integra 800 | 104-118 | 6.06 | 0.05 | 2.1 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Roche | 2.15
Roche Cobas Integra 800 | 104-118 | 11.81 | 0.1 | 1.9 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Roche | 2
Roche Cobas Integra 800 | 104-118 | 5.20 | 0.18 | 2.3 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Roche | 2.48
Roche Cobas Integra 800 | 104-118 | 7.61 | 0.03 | 2.1 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Roche | 2.13
Roche Cobas Integra 800 | 77-91 | 6.43 | 0.02 | 2.6 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Roche | 2.62
Roche Cobas Integra 800 | 77-91 | 9.61 | 0.08 | 2.1 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Roche | 2.18
Roche Cobas Integra 800 | 77-91 | 5.48 | 0.14 | 2.5 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Roche | 2.64
Roche Cobas Integra 800 | 77-91 | 8.59 | 0.08 | 1.9 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Roche | 1.98
Roche Cobas Integra 800 | 77-91 | 7.36 | 0.11 | 1.8 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Roche | 1.91
Roche Cobas Integra 800 | 61 | 7.42 | 0 | 2.6 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Roche | 2.6
Roche Cobas Integra 800 | 61 | 5.26 | 0.06 | 2.6 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Roche | 2.66
Roche Cobas Integra 800 | 61 | 8.36 | 0.05 | 2.1 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Roche | 2.15
Roche Cobas Integra 800 | 61 | 10.48 | 0.15 | 2.3 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Roche | 2.45
Roche Cobas Integra 800 | 61 | 5.94 | 0.07 | 2.7 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Roche | 2.77
Roche Cobas Integra 800 | 38 | 7.22 | -0.03 | 1.7 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Roche | 1.67
Roche Cobas Integra 800 | 38 | 6.79 | -0.02 | 2.6 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Roche | 2.58
Roche Cobas Integra 800 | 38 | 8.69 | 0.04 | 2.4 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Roche | 2.44
Roche Cobas Integra 800 | 38 | 9.60 | 0.1 | 2.8 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Roche | 2.9
Roche Cobas Integra 800 | 38 | 5.55 | 0.1 | 2.7 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Roche | 2.8
Roche Cobas Integra 800 | 14 | 7.01 | -0.14 | 1.7 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Roche | 1.56
Roche Cobas Integra 800 | 14 | 5.29 | 0.1 | 2.8 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Roche | 2.9
Roche Cobas Integra 800 | 14 | 8.34 | -0.08 | 1.6 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Roche | 1.52
Roche Cobas Integra 800 | 14 | 9.85 | 0.06 | 1.9 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Roche | 1.96
Roche Cobas Integra 800 | 14 | 6.12 | 0 | 1.7 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Roche | 1.7
Roche Cobas Integra 800 (non TQ) | 11 | 5.2 |  | 4.9 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Roche | 4.9
Roche Cobas Integra 800 (non TQ) | 11 | 8.5 |  | 3.1 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Roche | 3.1
Roche Cobas Integra 800 (non TQ) | 11 | 6.1 |  | 3.4 | GH2-03 | 6 | 2009 | CAP | Fresh | Roche | 3.4
Roche Cobas Integra 800 Gen.2 | 116 | 5.3 |  | 3.9 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Roche | 3.9
Roche Cobas Integra 800 Gen.2 | 116 | 8.5 |  | 3.6 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Roche | 3.6
Roche Cobas Integra 800 Gen.2 | 116 | 6.1 |  | 3.3 | GH2-03 | 6 | 2009 | CAP | Fresh | Roche | 3.3
Roche Cobas Integra Gen.2 | 20 | 6.9 |  | 3.1 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Roche | 3.1
Roche Cobas Integra Gen.2 | 20 | 7 |  | 2.8 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Roche | 2.8
Roche Cobas Integra Gen.2 | 20 | 11.6 |  | 4.1 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Roche | 4.1
Roche Cobas Integra Gen.2 | 72 | 5.5 |  | 3.3 | GH2-01 | 5.4 | 2007 | CAP | Fresh | Roche | 3.3
Roche Cobas Integra Gen.2 | 72 | 11.7 |  | 2.8 | GH2-02 | 11.6 | 2007 | CAP | Fresh | Roche | 2.8
Roche Cobas Integra Gen.2 | 72 | 7.5 |  | 2.3 | GH2-03 | 7.6 | 2007 | CAP | Fresh | Roche | 2.3
Roche Cobas Integra Gen.2 | 81 | 9 |  | 2.6 | GH2-04 | 9.2 | 2007 | CAP | Fresh | Roche | 2.6
Roche Cobas Integra Gen.2 | 81 | 11 |  | 3.2 | GH2-05 | 11.1 | 2007 | CAP | Fresh | Roche | 3.2
Roche Cobas Integra Gen.2 | 81 | 6.2 |  | 2.7 | GH2-06 | 6.2 | 2007 | CAP | Fresh | Roche | 2.7
Roche Cobas Integra Gen.2 | 90 | 12.5 |  | 3.6 | GH2-01 | 13 | 2008 | CAP | Fresh | Roche | 3.6
Roche Cobas Integra Gen.2 | 90 | 5.4 |  | 3.2 | GH2-02 | 5.3 | 2008 | CAP | Fresh | Roche | 3.2
Roche Cobas Integra Gen.2 | 90 | 8.8 |  | 2.9 | GH2-03 | 8.8 | 2008 | CAP | Fresh | Roche | 2.9
Roche Cobas Integra Gen.2 | 101 | 11.6 |  | 3.4 | GH2-04 | 11.9 | 2008 | CAP | Fresh | Roche | 3.4
Roche Cobas Integra Gen.2 | 101 | 6.8 |  | 3.4 | GH2-05 | 6.8 | 2008 | CAP | Fresh | Roche | 3.4
Roche Cobas Integra Gen.2 | 101 | 5.2 |  | 2.8 | GH2-06 | 5 | 2008 | CAP | Fresh | Roche | 2.8
Roche/Hitachi (Tina Quant II) | 62 | 10.6 |  | 3.3 | GH2-01 | 10.7 | 2006 | CAP | Fresh | Roche/Hitachi | 3.3
Roche/Hitachi (Tina Quant II) | 62 | 5.4 |  | 4.5 | GH2-02 | 5.30 | 2006 | CAP | Fresh | Roche/Hitachi | 4.5
Roche/Hitachi (Tina Quant II) | 62 | 8 |  | 4.7 | GH2-03 | 8.40 | 2006 | CAP | Fresh | Roche/Hitachi | 4.7
Roche/Hitachi (Tina Quant II) | 53 | 6.8 |  | 5.1 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Roche/Hitachi | 5.1
Roche/Hitachi (Tina Quant II) | 53 | 6.8 |  | 4.8 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Roche/Hitachi | 4.8
Roche/Hitachi (Tina Quant II) | 53 | 11.5 |  | 3.8 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Roche/Hitachi | 3.8
Roche/Hitachi (Tina Quant II) | 53 | 6.8 |  | 5.1 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Roche/Hitachi | 5.1
Roche/Hitachi (Tina Quant II) | 53 | 6.8 |  | 4.8 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Roche/Hitachi | 4.8
Roche/Hitachi (Tina Quant II) | 53 | 11.5 |  | 3.8 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Roche/Hitachi | 3.8
Roche/Hitachi (Tina Quant II) | 63 | 5.5 |  | 5.2 | GH2-01 | 5.4 | 2007 | CAP | Fresh | Roche/Hitachi | 5.2
Roche/Hitachi (Tina Quant II) | 63 | 11.8 |  | 4.2 | GH2-02 | 11.6 | 2007 | CAP | Fresh | Roche/Hitachi | 4.2
Roche/Hitachi (Tina Quant II) | 63 | 7.5 |  | 3.5 | GH2-03 | 7.6 | 2007 | CAP | Fresh | Roche/Hitachi | 3.5
Roche/Hitachi (Tina Quant II) | 44 | 9.1 |  | 3.5 | GH2-04 | 9.2 | 2007 | CAP | Fresh | Roche/Hitachi | 3.5
Roche/Hitachi (Tina Quant II) | 44 | 11.2 |  | 4.5 | GH2-05 | 11.1 | 2007 | CAP | Fresh | Roche/Hitachi | 4.5
Roche/Hitachi (Tina Quant II) | 44 | 6.3 |  | 5.1 | GH2-06 | 6.2 | 2007 | CAP | Fresh | Roche/Hitachi | 5.1
Roche/Hitachi (Tina Quant II) | 41 | 13.2 |  | 4.5 | GH2-01 | 13 | 2008 | CAP | Fresh | Roche/Hitachi | 4.5
Roche/Hitachi (Tina Quant II) | 41 | 5.4 |  | 3.6 | GH2-02 | 5.3 | 2008 | CAP | Fresh | Roche/Hitachi | 3.6
Roche/Hitachi (Tina Quant II) | 41 | 8.6 |  | 4.2 | GH2-03 | 8.8 | 2008 | CAP | Fresh | Roche/Hitachi | 4.2
Roche/Hitachi (Tina Quant II) | 41 | 11.7 |  | 3.1 | GH2-04 | 11.9 | 2008 | CAP | Fresh | Roche/Hitachi | 3.1
Roche/Hitachi (Tina Quant II) | 41 | 6.8 |  | 4.6 | GH2-05 | 6.8 | 2008 | CAP | Fresh | Roche/Hitachi | 4.6
Roche/Hitachi (Tina Quant II) | 41 | 5.2 |  | 4.5 | GH2-06 | 5 | 2008 | CAP | Fresh | Roche/Hitachi | 4.5
Roche/Hitachi Modular P | 20 | 5.2 |  | 3.2 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Roche/Hitachi | 3.2
Roche/Hitachi Modular P | 20 | 8.3 |  | 3.2 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Roche/Hitachi | 3.2
Roche/Hitachi Modular P | 20 | 6 |  | 4.7 | GH2-03 | 6 | 2009 | CAP | Fresh | Roche/Hitachi | 4.7
Roche/Hitachi Modular P | 15 | 6.61 | 0.01 | 3.5 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Roche/Hitachi | 3.51
Roche/Hitachi Modular P | 15 | 9.18 | -0.32 | 3.3 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Roche/Hitachi | 2.98
Roche/Hitachi Modular P | 15 | 7.27 | -0.13 | 4.4 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Roche/Hitachi | 4.27
Roche/Hitachi Modular P | 14 | 5.95 | 0.05 | 2.9 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Roche/Hitachi | 2.95
Roche/Hitachi Modular P | 14 | 9.63 | -0.17 | 4.1 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Roche/Hitachi | 3.93
Roche/Hitachi Modular P | 14 | 7.24 | -0.16 | 2.1 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Roche/Hitachi | 1.94
Roche/Hitachi Modular P | 17 | 5.29 | 0.09 | 4.3 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Roche/Hitachi | 4.39
Roche/Hitachi Modular P | 17 | 8.4 | -0.3 | 3.3 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Roche/Hitachi | 3
Roche/Hitachi Modular P | 17 | 6.22 | -0.08 | 3.6 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Roche/Hitachi | 3.52
Roche/Hitachi Modular P | 14 | 8.26 | -0.24 | 5.5 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Roche/Hitachi | 5.26
Roche/Hitachi Modular P | 14 | 5.41 | 0.01 | 3.9 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Roche/Hitachi | 3.91
Roche/Hitachi Modular P | 14 | 6.31 | -0.09 | 4 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Roche/Hitachi | 3.91
Roche/Hitachi Modular P | 127 | 6.31 | 0.01 | 1.8 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Roche/Hitachi | 1.81
Roche/Hitachi Modular P | 127 | 7.68 | 0.08 | 2.1 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Roche/Hitachi | 2.18
Roche/Hitachi Modular P | 127 | 9.37 | 0.17 | 2.2 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Roche/Hitachi | 2.37
Roche/Hitachi Modular P | 14 | 5.61 | 0.01 | 2.9 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Roche/Hitachi | 2.91
Roche/Hitachi Modular P | 14 | 9.29 | -0.11 | 3.8 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Roche/Hitachi | 3.69
Roche/Hitachi Modular P | 14 | 7.21 | 0.01 | 2.4 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Roche/Hitachi | 2.41
Sebia Capillarys 2 | 53 | 7.10 | -0.15 | 1.5 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Sebia | 1.35
Sebia Capillarys 2 | 53 | 6.70 | -0.11 | 1.9 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Sebia | 1.79
Sebia Capillarys 2 | 53 | 8.50 | -0.15 | 1.3 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Sebia | 1.15
Sebia Capillarys 2 | 53 | 9.33 | -0.17 | 1.6 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Sebia | 1.43
Sebia Capillarys 2 | 53 | 5.34 | -0.11 | 2 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Sebia | 1.89
Sebia Capillarys 2 | 53 | 7.09 | -0.06 | 1.6 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Sebia | 1.54
Sebia Capillarys 2 | 53 | 5.12 | -0.07 | 2 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Sebia | 1.93
Sebia Capillarys 2 | 53 | 8.30 | -0.12 | 1.4 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Sebia | 1.28
Sebia Capillarys 2 | 53 | 9.54 | -0.25 | 1.6 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Sebia | 1.35
Sebia Capillarys 2 | 53 | 6.03 | -0.09 | 1.7 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Sebia | 1.61
Sebia Capillarys 2 | 52 | 6.25 | -0.06 | 1.9 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Sebia | 1.84
Sebia Capillarys 2 | 52 | 8.94 | -0.17 | 1.3 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Sebia | 1.13
Sebia Capillarys 2 | 52 | 5.23 | -0.08 | 1.9 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Sebia | 1.82
Sebia Capillarys 2 | 52 | 8.33 | -0.02 | 1.8 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Sebia | 1.78
Sebia Capillarys 2 | 52 | 7.61 | -0.01 | 1.4 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Sebia | 1.39
Sebia Capillarys 2 Flex Piercing | 73 | 5.59 | -0.07 | 2.8 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Sebia | 2.73
Sebia Capillarys 2 Flex Piercing | 73 | 9.28 | -0.03 | 1.3 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Sebia | 1.27
Sebia Capillarys 2 Flex Piercing | 73 | 5.18 | -0.1 | 2.4 | GH5-04 | 5.28 | 2019 | CAP | Fresh | Sebia | 2.3
Sebia Capillarys 2 Flex Piercing | 73 | 7.33 | -0.08 | 1.4 | GH5-05 | 7.41 | 2019 | CAP | Fresh | Sebia | 1.32
Sebia Capillarys 2 Flex Piercing | 71 | 9.21 | -0.16 | 1.6 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Sebia | 1.44
Sebia Capillarys 2 Flex Piercing | 71 | 4.93 | 0.01 | 2.2 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Sebia | 2.21
Sebia Capillarys 2 Flex Piercing | 71 | 6.07 | 0.03 | 1.5 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Sebia | 1.53
Sebia Capillarys 2 Flex Piercing | 71 | 5.26 | -0.02 | 1.9 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Sebia | 1.88
Sebia Capillarys 2 Flex Piercing | 71 | 8.03 | -0.05 | 2.1 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Sebia | 2.05
Sebia Capillarys 2 Flex Piercing | 63 | 6.39 | -0.02 | 1.9 | GH5-06 | 6.41 | 2019 | CAP | Fresh | Sebia | 1.88
Sebia Capillarys 2 Flex Piercing | 63 | 8.58 | -0.02 | 1.3 | GH5-07 | 8.6 | 2019 | CAP | Fresh | Sebia | 1.28
Sebia Capillarys 2 Flex Piercing | 63 | 5.36 | -0.06 | 1.9 | GH5-08 | 5.42 | 2019 | CAP | Fresh | Sebia | 1.84
Sebia Capillarys 2 Flex Piercing | 63 | 7.38 | 0 | 1.3 | GH5-09 | 7.38 | 2019 | CAP | Fresh | Sebia | 1.3
Sebia Capillarys 2 Flex Piercing | 63 | 9.52 | -0.23 | 1.5 | GH5-10 | 9.75 | 2019 | CAP | Fresh | Sebia | 1.27
Sebia Capillarys 2 Flex Piercing | 67 | 5.18 | -0.07 | 1.8 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Sebia | 1.73
Sebia Capillarys 2 Flex Piercing | 67 | 6.33 | -0.08 | 1.6 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Sebia | 1.52
Sebia Capillarys 2 Flex Piercing | 67 | 8.12 | -0.09 | 1.4 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Sebia | 1.31
Sebia Capillarys 2 Flex Piercing | 67 | 5.43 | -0.04 | 1.7 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Sebia | 1.66
Sebia Capillarys 2 Flex Piercing | 67 | 9.46 | -0.02 | 1.5 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Sebia | 1.48
Sebia Capillarys 2 Flex Piercing | 62 | 7.91 | -0.06 | 2 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Sebia | -0.06
Sebia Capillarys 2 Flex Piercing | 62 | 5.87 | -0.02 | 2 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Sebia | -0.02
Sebia Capillarys 2 Flex Piercing | 62 | 5.1 | -0.03 | 2 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Sebia | -0.03
Sebia Capillarys 2 Flex Piercing | 62 | 7.24 | -0.16 | 1.9 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Sebia | 1.74
Sebia Capillarys 2 Flex Piercing | 62 | 9.14 | -0.03 | 1.3 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Sebia | 1.27
Sebia Capillarys 2 Flex Piercing | 61 | 7.77 | -0.04 | 1.5 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Sebia | 1.46
Sebia Capillarys 2 Flex Piercing | 61 | 5.82 | -0.09 | 1.8 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Sebia | 1.71
Sebia Capillarys 2 Flex Piercing | 61 | 8.6 | -0.09 | 1.5 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Sebia | 1.41
Sebia Capillarys 2 Flex Piercing | 61 | 5.82 | -0.04 | 1.6 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Sebia | 1.56
Sebia Capillarys 2 Flex Piercing | 61 | 9.38 | -0.18 | 1.2 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Sebia | 1.02
Sebia Capillarys 2 Flex Piercing | 61 | 5.4 | -0.08 | 2.2 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Sebia | 2.12
Sebia Capillarys 2 Flex Piercing | 61 | 8.2 | -0.06 | 1.5 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Sebia | 1.44
Sebia Capillarys 2 Flex Piercing | 61 | 5.09 | -0.06 | 2.1 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Sebia | 2.04
Sebia Capillarys 2 Flex Piercing | 61 | 9.95 | -0.1 | 1.2 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Sebia | 1.1
Sebia Capillarys 2 Flex Piercing | 61 | 5.94 | -0.05 | 1.2 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Sebia | 1.15
Sebia Capillarys 2 Flex Piercing | 60 | 8.79 | -0.07 | 1.7 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Sebia | 1.63
Sebia Capillarys 2 Flex Piercing | 60 | 6.69 | -0.15 | 1.9 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Sebia | 1.75
Sebia Capillarys 2 Flex Piercing | 60 | 9.24 | -0.15 | 1.8 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Sebia | 1.65
Sebia Capillarys 2 Flex Piercing | 60 | 5.95 | -0.18 | 2 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Sebia | -0.18
Sebia Capillarys 2 Flex Piercing | 60 | 7.37 | -0.15 | 1.6 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Sebia | 1.45
Sebia Capillarys 2 Flex Piercing | 53 | 5.72 | -0.08 | 2.1 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Sebia | 2.02
Sebia Capillarys 2 Flex Piercing | 53 | 8.93 | -0.19 | 1.5 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Sebia | 1.31
Sebia Capillarys 2 Flex Piercing | 53 | 5.43 | -0.04 | 2.3 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Sebia | 2.26
Sebia Capillarys 2 Flex Piercing | 53 | 7.07 | -0.18 | 1.3 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Sebia | 1.12
Sebia Capillarys 2 Flex Piercing | 53 | 9.13 | -0.13 | 1.4 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Sebia | 1.27
Sebia Capillarys 2 Flex Piercing | 55 | 5.13 | -0.08 | 2.3 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Sebia | 2.22
Sebia Capillarys 2 Flex Piercing | 55 | 6.7 | 0.02 | 1.7 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Sebia | 1.72
Sebia Capillarys 2 Flex Piercing | 55 | 8.05 |  | 1.6 | GH5-13 |  | 2021 | CAP | Fresh | Sebia | 1.6
Sebia Capillarys 2 Flex Piercing | 55 | 5.98 | 0 | 1.7 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Sebia | 1.7
Sebia Capillarys 2 Flex Piercing | 55 | 10.03 | -0.09 | 1.6 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Sebia | 1.51
Sebia Capillarys 2 Flex Piercing | 73 | 5.26 | -0.2 | 3.1 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Sebia | 2.9
Sebia Capillarys 2 Flex Piercing | 36 | 7.29 | -0.13 | 1.7 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Sebia | 1.57
Sebia Capillarys 2 Flex Piercing | 36 | 5.05 | -0.15 | 2.2 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Sebia | 2.05
Sebia Capillarys 2 Flex Piercing | 36 | 8.28 | -0.03 | 1.4 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Sebia | 1.37
Sebia Capillarys 2 Flex Piercing | 36 | 10.26 | -0.07 | 1.3 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Sebia | 1.23
Sebia Capillarys 2 Flex Piercing | 36 | 5.74 | -0.13 | 2.1 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Sebia | 1.97
Sebia Capillarys 2 Flex Piercing | 32-44 | 6.34 | -0.07 | 2.2 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Sebia | 2.13
Sebia Capillarys 2 Flex Piercing | 32-44 | 9.44 | -0.09 | 1.2 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Sebia | 1.11
Sebia Capillarys 2 Flex Piercing | 32-44 | 5.24 | -0.1 | 1.7 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Sebia | 1.6
Sebia Capillarys 2 Flex Piercing | 32-44 | 8.52 | 0.01 | 1.4 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Sebia | 1.41
Sebia Capillarys 2 Flex Piercing | 32-44 | 7.14 | -0.11 | 2.3 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Sebia | 2.19
Sebia Capillarys 2/ Minicap | 15 | 6.43 | -0.15 | 2 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Sebia | 1.85
Sebia Capillarys 2/ Minicap | 15 | 8.32 | -0.07 | 1.3 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Sebia | 1.23
Sebia Capillarys 2/ Minicap | 15 | 5.43 | -0.22 | 1.1 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Sebia | 0.88
Sebia Capillarys 2/ Minicap | 12 | 6.38 | -0.11 | 1.4 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Sebia | 1.29
Sebia Capillarys 2/ Minicap | 12 | 6.86 | -0.11 | 1 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Sebia | 0.89
Sebia Capillarys 2/ Minicap | 12 | 9.55 | -0.1 | 0.5 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Sebia | 0.4
Sebia Capillarys 2/ Minicap | 20 | 6.62 | -0.17 | 1.6 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Sebia | 1.43
Sebia Capillarys 2/ Minicap | 20 | 10.04 | -0.24 | 2.5 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Sebia | 2.26
Sebia Capillarys 2/ Minicap | 20 | 6.6 | -0.22 | 2.1 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Sebia | 1.88
Sebia Capillarys 2/ Minicap | 12 | 8.58 | -0.05 | 1.7 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Sebia | 1.65
Sebia Capillarys 2/ Minicap | 12 | 5.14 | -0.18 | 2.7 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Sebia | 2.52
Sebia Capillarys 2/ Minicap | 17 | 5.66 | -0.18 | 2.7 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Sebia | 2.52
Sebia Capillarys 2/ Minicap | 17 | 11.6 | -0.11 | 1.8 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Sebia | 1.69
Sebia Capillarys 2/ Minicap | 17 | 9.41 | -0.12 | 1.5 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Sebia | 1.38
Sebia Capillarys 2/ Minicap | 17 | 4.82 | -0.22 | 2.8 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Sebia | 2.58
Sebia Capillarys 2/ Minicap | 17 | 7.18 | -0.2 | 1.9 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Sebia | 1.7
Sebia Capillarys 2/ Minicap | 17-28 | 11.59 | -0.1 | 2.2 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Sebia | 2.1
Sebia Capillarys 2/ Minicap | 17-28 | 5.83 | -0.1 | 1.8 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Sebia | 1.7
Sebia Capillarys 2/ Minicap | 17-28 | 5.03 | -0.14 | 1.6 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Sebia | 1.46
Sebia Capillarys 2/ Minicap | 17-28 | 8.08 | -0.06 | 1.6 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Sebia | 1.54
Sebia Capillarys 2/ Minicap | 17-28 | 9.14 | -0.16 | 1.7 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Sebia | 1.54
Sebia Capillarys 2/ Minicap | 19-31 | 5.16 | -0.16 | 2.2 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Sebia | 2.04
Sebia Capillarys 2/ Minicap | 19-31 | 9.17 | 0.01 | 1.7 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Sebia | 1.71
Sebia Capillarys 2/ Minicap | 19-31 | 5.15 | -0.16 | 2.5 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Sebia | 2.34
Sebia Capillarys 2/ Minicap | 19-31 | 11.94 | -0.09 | 1.9 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Sebia | 1.81
Sebia Capillarys 2/ Minicap | 19-31 | 5.83 | -0.11 | 1.6 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Sebia | 1.49
Sebia Capillarys 2/ Minicap | 25 | 5.14 | -0.13 | 2.6 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Sebia | 2.47
Sebia Capillarys 2/ Minicap | 25 | 10.52 | -0.07 | 1.3 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Sebia | 1.23
Sebia Capillarys 2/ Minicap | 25 | 6.00 | -0.2 | 1.7 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Sebia | 1.5
Sebia Capillarys 2/ Minicap | 25 | 12.14 | -0.09 | 1.6 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Sebia | 1.51
Sebia Capillarys 2/ Minicap | 25 | 7.44 | -0.07 | 1.3 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Sebia | 1.23
Sebia Capillarys 2/ Minicap | 28-43 | 9.02 | -0.09 | 1.5 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Sebia | 1.41
Sebia Capillarys 2/ Minicap | 28-43 | 5.89 | -0.12 | 2.3 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Sebia | 2.18
Sebia Capillarys 2/ Minicap | 28-43 | 11.53 | -0.18 | 1.8 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Sebia | 1.62
Sebia Capillarys 2/ Minicap | 28-43 | 4.90 | -0.12 | 2 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Sebia | 1.88
Sebia Capillarys 2/ Minicap | 28-43 | 7.56 | -0.02 | 2 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Sebia | 1.98
Sebia Capillarys 3 (CAPI 3) Tera | 14 | 7.93 | -0.04 | 2.8 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Sebia | 2.76
Sebia Capillarys 3 (CAPI 3) Tera | 14 | 5.94 | 0.05 | 1.9 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Sebia | 1.95
Sebia Capillarys 3 (CAPI 3) Tera | 14 | 5.1 | -0.03 | 1.5 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Sebia | 1.47
Sebia Capillarys 3 (CAPI 3) Tera | 14 | 7.25 | -0.15 | 2.2 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Sebia | 2.05
Sebia Capillarys 3 (CAPI 3) Tera | 14 | 9.16 | -0.01 | 1.4 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Sebia | 1.39
Sebia Capillarys 3 (CAPI 3) Tera | 16 | 7.74 | -0.07 | 1.4 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Sebia | 1.33
Sebia Capillarys 3 (CAPI 3) Tera | 16 | 5.81 | -0.1 | 1.8 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Sebia | 1.7
Sebia Capillarys 3 (CAPI 3) Tera | 16 | 8.59 | -0.1 | 0.8 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Sebia | 0.7
Sebia Capillarys 3 (CAPI 3) Tera | 16 | 5.78 | -0.08 | 1.3 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Sebia | 1.22
Sebia Capillarys 3 (CAPI 3) Tera | 16 | 9.39 | -0.17 | 1.3 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Sebia | 1.13
Sebia Capillarys 3 (CAPI 3) Tera/Octa | 22 | 8.82 | -0.04 | 1.4 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Sebia | 1.36
Sebia Capillarys 3 (CAPI 3) Tera/Octa | 22 | 6.72 | -0.12 | 1.8 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Sebia | 1.68
Sebia Capillarys 3 (CAPI 3) Tera/Octa | 22 | 9.28 | -0.11 | 1.6 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Sebia | 1.49
Sebia Capillarys 3 (CAPI 3) Tera/Octa | 22 | 5.95 | -0.18 | 1.7 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Sebia | 1.52
Sebia Capillarys 3 (CAPI 3) Tera/Octa | 22 | 7.39 | -0.13 | 1.1 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Sebia | 0.97
Sebia Capillarys 3 (CAPI 3) Tera/Octa | 22 | 5.39 | -0.09 | 1.5 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Sebia | 1.41
Sebia Capillarys 3 (CAPI 3) Tera/Octa | 22 | 8.2 | -0.06 | 1 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Sebia | -0.06
Sebia Capillarys 3 (CAPI 3) Tera/Octa | 22 | 5.09 | -0.06 | 1.6 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Sebia | 1.54
Sebia Capillarys 3 (CAPI 3) Tera/Octa | 22 | 9.91 | -0.14 | 1.4 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Sebia | 1.26
Sebia Capillarys 3 (CAPI 3) Tera/Octa | 22 | 5.97 | -0.02 | 1.3 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Sebia | 1.28
Sebia Capillarys 3 (CAPI 3) Tera/Octa | 22 | 5.74 | -0.06 | 1.6 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Sebia | 1.54
Sebia Capillarys 3 (CAPI 3) Tera/Octa | 22 | 8.99 | -0.13 | 1.2 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Sebia | 1.07
Sebia Capillarys 3 (CAPI 3) Tera/Octa | 22 | 5.43 | -0.04 | 1.5 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Sebia | 1.46
Sebia Capillarys 3 (CAPI 3) Tera/Octa | 22 | 7.06 | -0.19 | 1.1 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Sebia | 0.91
Sebia Capillarys 3 (CAPI 3) Tera/Octa | 22 | 9.17 | -0.09 | 1 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Sebia | -0.09
Sebia Capillarys 3 (CAPI 3) Tera/Octa | 26 | 5.2 | -0.01 | 2.3 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Sebia | 2.29
Sebia Capillarys 3 (CAPI 3) Tera/Octa | 26 | 6.69 | 0.01 | 1.6 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Sebia | 1.61
Sebia Capillarys 3 (CAPI 3) Tera/Octa | 26 | 8.06 |  | 1.5 | GH5-13 |  | 2021 | CAP | Fresh | Sebia | 1.5
Sebia Capillarys 3 (CAPI 3) Tera/Octa | 26 | 6.01 | 0.03 | 1.3 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Sebia | 1.33
Sebia Capillarys 3 (CAPI 3) Tera/Octa | 26 | 10.1 | -0.02 | 1.2 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Sebia | 1.18
Siemens Advia | 47 | 5.86 | -0.04 | 5.7 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Siemens | 5.66
Siemens Advia | 47 | 9.73 | -0.07 | 4.3 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Siemens | 4.23
Siemens Advia | 47 | 7.49 | 0.09 | 4.5 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Siemens | 4.59
Siemens Advia | 52 | 5.3 | 0.1 | 4.2 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Siemens | 4.3
Siemens Advia | 52 | 8.74 | 0.04 | 2.5 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Siemens | 2.54
Siemens Advia | 52 | 6.43 | 0.13 | 4.1 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Siemens | 4.23
Siemens Advia | 48 | 8.51 | 0.01 | 3.3 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Siemens | 3.31
Siemens Advia | 48 | 5.49 | 0.09 | 3.8 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Siemens | 3.89
Siemens Advia | 48 | 6.54 | 0.14 | 3.9 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Siemens | 4.04
Siemens Advia | 32 | 6.52 | -0.06 | 4.3 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Siemens | 4.24
Siemens Advia | 32 | 8.33 | -0.06 | 4.2 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Siemens | 4.14
Siemens Advia | 32 | 5.5 | -0.15 | 6 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Siemens | 5.85
Siemens Advia | 40 | 6.78 | 0.29 | 3.8 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Siemens | 4.09
Siemens Advia | 40 | 7.27 | 0.3 | 3.1 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Siemens | 3.4
Siemens Advia | 40 | 9.72 | 0.07 | 3.7 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Siemens | 3.77
Siemens Advia | 29 | 6.93 | 0.14 | 4 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Siemens | 4.14
Siemens Advia | 29 | 10.09 | -0.19 | 4.9 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Siemens | 4.71
Siemens Advia | 29 | 6.88 | 0.06 | 2.9 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Siemens | 2.96
Siemens Advia | 25 | 8.64 | 0.01 | 2.9 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Siemens | 2.91
Siemens Advia | 25 | 5.36 | 0.04 | 3.7 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Siemens | 3.74
Siemens Advia | 24 | 6.02 | 0.18 | 5.1 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Siemens | 5.28
Siemens Advia | 24 | 11.72 | 0.01 | 4.5 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Siemens | 4.51
Siemens Advia | 24 | 9.71 | 0.18 | 5.2 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Siemens | 5.38
Siemens Advia | 24 | 5.14 | 0.1 | 2.7 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Siemens | 2.8
Siemens Advia | 24 | 7.74 | 0.36 | 5.5 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Siemens | 5.86
Siemens Advia | 22-24 | 11.53 | -0.16 | 3.4 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Siemens | 3.24
Siemens Advia | 22-24 | 5.98 | 0.05 | 4.1 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Siemens | 4.15
Siemens Advia | 22-24 | 5.16 | -0.01 | 4.1 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Siemens | 4.09
Siemens Advia | 22-24 | 8.17 | 0.03 | 3.5 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Siemens | 3.53
Siemens Advia | 22-24 | 8.78 | -0.52 | 3.9 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Siemens | 3.38
Siemens Advia | 25-27 | 5.42 | 0.1 | 4.2 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Siemens | 4.3
Siemens Advia | 25-27 | 9.31 | 0.15 | 3 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Siemens | 3.15
Siemens Advia | 25-27 | 5.39 | 0.08 | 4.1 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Siemens | 4.18
Siemens Advia | 25-27 | 11.79 | -0.24 | 3.3 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Siemens | 3.06
Siemens Advia | 25-27 | 6.03 | 0.09 | 3.7 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Siemens | 3.79
Siemens Advia | 25 | 5.27 | 0 | 2.8 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Siemens | 2.8
Siemens Advia | 25 | 10.50 | -0.09 | 3.2 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Siemens | 3.11
Siemens Advia | 25 | 6.12 | -0.08 | 3.7 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Siemens | 3.62
Siemens Advia | 25 | 11.88 | -0.35 | 3.4 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Siemens | 3.05
Siemens Advia | 25 | 7.56 | 0.05 | 3.5 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Siemens | 3.55
Siemens Advia | 21-24 | 9.08 | -0.03 | 2.3 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Siemens | 2.27
Siemens Advia | 21-24 | 6.09 | 0.08 | 2.5 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Siemens | 2.58
Siemens Advia | 21-24 | 11.42 | -0.29 | 3 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Siemens | 2.71
Siemens Advia | 21-24 | 4.98 | -0.04 | 3.8 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Siemens | 3.76
Siemens Advia | 21-24 | 7.60 | 0.02 | 2.6 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Siemens | 2.62
Siemens Advia | 18-19 | 6.57 | 0.16 | 2.7 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Siemens | 2.86
Siemens Advia | 18-19 | 9.49 | -0.04 | 2.8 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Siemens | 2.76
Siemens Advia | 18-19 | 5.46 | 0.12 | 2.3 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Siemens | 2.42
Siemens Advia | 18-19 | 8.63 | 0.12 | 2.1 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Siemens | 2.22
Siemens Advia | 18-19 | 7.65 | 0.4 | 3.2 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Siemens | 3.6
Siemens Advia | 15 | 7.19 | 0.04 | 2.8 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Siemens | 2.84
Siemens Advia | 15 | 5.31 | 0.12 | 3.9 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Siemens | 4.02
Siemens Advia | 15 | 8.32 | -0.1 | 3.1 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Siemens | 3
Siemens Advia | 15 | 9.60 | -0.19 | 3.3 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Siemens | 3.11
Siemens Advia | 15 | 6.16 | 0.04 | 2.4 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Siemens | 2.44
Siemens Advia | 12 | 6.43 | 0.12 | 2.1 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Siemens | 2.22
Siemens Advia | 12 | 9.17 | 0.06 | 3 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Siemens | 3.06
Siemens Advia | 12 | 5.35 | 0.04 | 2.8 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Siemens | 2.84
Siemens Advia | 12 | 8.29 | -0.06 | 3.1 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Siemens | 3.04
Siemens Advia | 12 | 7.77 | 0.15 | 3.7 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Siemens | 3.85
Siemens Advia | 16 | 7.38 | -0.04 | 1.8 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Siemens | 1.76
Siemens Advia | 16 | 5.16 | -0.04 | 2.6 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Siemens | 2.56
Siemens Advia | 16 | 8.24 | -0.07 | 2.9 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Siemens | 2.83
Siemens Advia | 16 | 9.95 | -0.38 | 2.6 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Siemens | 2.22
Siemens Advia | 16 | 5.89 | 0.02 | 2.5 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Siemens | 2.52
Siemens Advia | 16 | 7.29 | 0.04 | 2.8 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Siemens | 2.84
Siemens Advia | 16 | 6.79 | -0.02 | 2.8 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Siemens | 2.78
Siemens Advia | 16 | 8.59 | -0.06 | 2.3 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Siemens | 2.24
Siemens Advia | 16 | 9.31 | -0.19 | 2.3 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Siemens | 2.11
Siemens Advia | 16 | 5.45 | 0 | 2.8 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Siemens | 2.8
Siemens Advia 1650 | 16 | 5.1 |  | 4.1 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Siemens | 4.1
Siemens Advia 1650 | 16 | 8.3 |  | 3 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Siemens | 3
Siemens Advia 1650 | 16 | 6.1 |  | 2.7 | GH2-03 | 6 | 2009 | CAP | Fresh | Siemens | 2.7
Siemens Advia 1650/2400 | 34 | 13.6 |  | 7.1 | GH2-01 | 13 | 2008 | CAP | Fresh | Siemens | 7.1
Siemens Advia 1650/2400 | 34 | 5.6 |  | 4.2 | GH2-02 | 5.3 | 2008 | CAP | Fresh | Siemens | 4.2
Siemens Advia 1650/2400 | 34 | 9 |  | 3 | GH2-03 | 8.8 | 2008 | CAP | Fresh | Siemens | 3
Siemens Advia 1650/2400 | 42 | 12 |  | 5.1 | GH2-04 | 11.9 | 2008 | CAP | Fresh | Siemens | 5.1
Siemens Advia 1650/2400 | 42 | 7 |  | 3.6 | GH2-05 | 6.8 | 2008 | CAP | Fresh | Siemens | 3.6
Siemens Advia 1650/2400 | 42 | 5.3 |  | 4.2 | GH2-06 | 5 | 2008 | CAP | Fresh | Siemens | 4.2
Siemens Advia 1800 | 14 | 5.1 |  | 5.4 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Siemens | 5.4
Siemens Advia 1800 | 14 | 8.4 |  | 4.1 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Siemens | 4.1
Siemens Advia 1800 | 14 | 6.1 |  | 5.1 | GH2-03 | 6 | 2009 | CAP | Fresh | Siemens | 5.1
Siemens Advia A1c_3 Reagent | 59 | 5.56 | 0.16 | 5.7 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Siemens | 5.86
Siemens Advia A1c_3 Reagent | 59 | 8.54 | 0.24 | 5.6 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Siemens | 5.84
Siemens Advia A1c_3 Reagent | 59 | 5.9 | 0.25 | 6.9 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Siemens | 7.15
Siemens Advia Chemistry (new) | 21 | 6.41 | -0.19 | 7.6 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Siemens | 7.41
Siemens Advia Chemistry (new) | 21 | 9 | -0.5 | 3.4 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Siemens | 2.9
Siemens Advia Chemistry (new) | 21 | 6.98 | -0.42 | 9.7 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Siemens | 9.28
Siemens Advia Chemistry (original) | 26 | 6.79 | 0.19 | 3.8 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Siemens | 3.99
Siemens Advia Chemistry (original) | 26 | 9.31 | -0.19 | 4.1 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Siemens | 3.91
Siemens Advia Chemistry (original) | 26 | 7.54 | 0.14 | 4.1 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Siemens | 4.24
Siemens Advia Chemistry Systems | 43 | 7.19 | 0.08 | 4.7 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Siemens | 4.78
Siemens Advia Chemistry Systems | 43 | 9.28 | -0.04 | 3.8 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Siemens | 3.76
Siemens Advia Chemistry Systems | 43 | 6.33 | 0.26 | 4.4 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Siemens | 4.66
Siemens Advia Chemistry Systems | 44 | 8.22 | 0.17 | 4 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Siemens | 4.17
Siemens Advia Chemistry Systems | 44 | 5.5 | 0.2 | 3.6 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Siemens | 3.8
Siemens Advia Chemistry Systems | 44 | 6.42 | 0.28 | 3.8 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Siemens | 4.08
Siemens Advia Chemistry Systems | 13 | 9.25 | -0.12 | 3.7 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Siemens | 3.58
Siemens Advia Chemistry Systems | 13 | 5.02 | 0.1 | 3.3 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Siemens | 3.4
Siemens Advia Chemistry Systems | 13 | 6.22 | 0.18 | 3.6 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Siemens | 3.78
Siemens Advia Chemistry Systems | 13 | 5.40 | 0.12 | 2.2 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Siemens | 2.32
Siemens Advia Chemistry Systems | 13 | 7.93 | -0.15 | 3.8 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Siemens | 3.65
Siemens Advia New Reagent | 21 | 5.63 | -0.27 | 4 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Siemens | 3.73
Siemens Advia New Reagent | 21 | 9.61 | -0.19 | 4 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Siemens | 3.81
Siemens Advia New Reagent | 21 | 7.31 | -0.09 | 4.1 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Siemens | 4.01
Siemens Advia New Reagent | 24 | 5.22 | 0.02 | 4.2 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Siemens | 4.22
Siemens Advia New Reagent | 24 | 8.73 | 0.03 | 2.4 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Siemens | 2.43
Siemens Advia New Reagent | 24 | 6.33 | 0.03 | 3.9 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Siemens | 3.93
Siemens Advia New Reagent | 25 | 8.48 | -0.02 | 3.3 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Siemens | 3.28
Siemens Advia New Reagent | 25 | 5.42 | 0.02 | 3.9 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Siemens | 3.92
Siemens Advia New Reagent | 25 | 6.47 | 0.07 | 4.2 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Siemens | 4.27
Siemens Advia New Reagent | 10 | 6.32 | 0.02 | 4.6 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Siemens | 4.62
Siemens Advia New Reagent | 10 | 7.49 | -0.11 | 4 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Siemens | 3.89
Siemens Advia New Reagent | 10 | 9.13 | -0.07 | 5.4 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Siemens | 5.33
Siemens Advia New Reagent | 41 | 5.35 | -0.25 | 5.1 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Siemens | 4.85
Siemens Advia New Reagent | 41 | 9.01 | -0.39 | 6.7 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Siemens | 6.31
Siemens Advia New Reagent | 41 | 7.02 | -0.18 | 5.7 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Siemens | 5.52
Siemens Advia Original Reagent | 13 | 6.03 | 0.13 | 2.7 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Siemens | 2.83
Siemens Advia Original Reagent | 13 | 9.65 | -0.15 | 4.1 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Siemens | 3.95
Siemens Advia Original Reagent | 13 | 7.51 | 0.11 | 3.5 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Siemens | 3.61
Siemens Advia Original Reagent | 13 | 5.42 | 0.22 | 3.3 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Siemens | 3.52
Siemens Advia Original Reagent | 13 | 8.83 | 0.13 | 2.5 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Siemens | 2.63
Siemens Advia Original Reagent | 13 | 6.55 | 0.25 | 2.7 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Siemens | 2.95
Siemens Advia Original Reagent | 21 | 8.54 | 0.04 | 2.8 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Siemens | 2.84
Siemens Advia Original Reagent | 21 | 5.57 | 0.17 | 3.1 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Siemens | 3.27
Siemens Advia Original Reagent | 21 | 6.62 | 0.22 | 3.3 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Siemens | 3.52
Siemens Advia Original Reagent | 38 | 6.29 | -0.01 | 3.8 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Siemens | 3.79
Siemens Advia Original Reagent | 38 | 7.54 | -0.06 | 4 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Siemens | 3.94
Siemens Advia Original Reagent | 38 | 9.00 | -0.2 | 4.7 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Siemens | 4.5
Siemens Advia Original Reagent | 20 | 5.8 | 0.2 | 3.3 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Siemens | 3.5
Siemens Advia Original Reagent | 20 | 9.69 | 0.29 | 5.1 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Siemens | 5.39
Siemens Advia Original Reagent | 20 | 7.53 | 0.33 | 4.2 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Siemens | 4.53
Siemens Atellica CH- Enzymatic | 25 | 7.86 | -0.11 | 2.4 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Siemens | 2.29
Siemens Atellica CH- Enzymatic | 25 | 5.68 | 0.21 | 3.2 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Siemens | 3.41
Siemens Atellica CH- Enzymatic | 25 | 4.92 | -0.21 | 2.9 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Siemens | 2.69
Siemens Atellica CH- Enzymatic | 25 | 7.15 | -0.25 | 1.6 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Siemens | 1.35
Siemens Atellica CH- Enzymatic | 25 | 8.96 | -0.21 | 2.3 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Siemens | 2.09
Siemens Atellica CH- Enzymatic | 34 | 7.63 | -0.18 | 3.1 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Siemens | 2.92
Siemens Atellica CH- Enzymatic | 34 | 5.65 | -0.26 | 2.9 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Siemens | 2.64
Siemens Atellica CH- Enzymatic | 34 | 8.48 | -0.21 | 2.9 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Siemens | 2.69
Siemens Atellica CH- Enzymatic | 34 | 5.64 | -0.22 | 2.2 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Siemens | 1.98
Siemens Atellica CH- Enzymatic | 34 | 9.36 | -0.2 | 3.2 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Siemens | 3
Siemens Atellica CH- Enzymatic reagent | 75 | 8.82 | -0.04 | 1.9 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Siemens | 1.86
Siemens Atellica CH- Enzymatic reagent | 75 | 6.67 | -0.17 | 2 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Siemens | -0.17
Siemens Atellica CH- Enzymatic reagent | 75 | 9.36 | -0.03 | 2 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Siemens | -0.03
Siemens Atellica CH- Enzymatic reagent | 75 | 5.91 | -0.22 | 1.8 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Siemens | 1.58
Siemens Atellica CH- Enzymatic reagent | 75 | 7.41 | -0.11 | 1.8 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Siemens | 1.69
Siemens Atellica CH- Enzymatic reagent | 58 | 5.45 | -0.03 | 2.8 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Siemens | 2.77
Siemens Atellica CH- Enzymatic reagent | 58 | 8.34 | 0.08 | 1.9 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Siemens | 1.98
Siemens Atellica CH- Enzymatic reagent | 58 | 5.11 | -0.04 | 3 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Siemens | -0.04
Siemens Atellica CH- Enzymatic reagent | 58 | 10.23 | 0.18 | 1.7 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Siemens | 1.88
Siemens Atellica CH- Enzymatic reagent | 58 | 5.97 | -0.02 | 2.7 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Siemens | 2.68
Siemens Atellica CH- Enzymatic reagent | 71 | 5.57 | -0.23 | 1.7 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Siemens | 1.47
Siemens Atellica CH- Enzymatic reagent | 71 | 9.04 | -0.08 | 1 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Siemens | -0.08
Siemens Atellica CH- Enzymatic reagent | 71 | 5.28 | -0.19 | 2 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Siemens | -0.19
Siemens Atellica CH- Enzymatic reagent | 71 | 7.15 | -0.1 | 1.3 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Siemens | 1.2
Siemens Atellica CH- Enzymatic reagent | 71 | 9.18 | -0.08 | 1.3 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Siemens | 1.22
Siemens Atellica CH- Enzymatic reagent | 83 | 5.07 | -0.14 | 2 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Siemens | -0.14
Siemens Atellica CH- Enzymatic reagent | 83 | 6.55 | -0.13 | 1.4 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Siemens | 1.27
Siemens Atellica CH- Enzymatic reagent | 83 | 8.21 |  | 1.8 | GH5-13 |  | 2021 | CAP | Fresh | Siemens | 1.8
Siemens Atellica CH- Enzymatic reagent | 83 | 5.9 | -0.08 | 1.8 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Siemens | 1.72
Siemens Atellica CH- Enzymatic reagent | 83 | 10.27 | 0.15 | 1.2 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Siemens | 1.35
Siemens Atellica CH- Immunoassay | 17 | 8.1 | 0.13 | 2.2 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Siemens | 2.33
Siemens Atellica CH- Immunoassay | 17 | 6 | 0.11 | 2.9 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Siemens | 3.01
Siemens Atellica CH- Immunoassay | 17 | 5.13 | 0 | 3.2 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Siemens | 3.2
Siemens Atellica CH- Immunoassay | 17 | 7.59 | 0.19 | 3 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Siemens | 0.19
Siemens Atellica CH- Immunoassay | 17 | 9.23 | 0.06 | 2.4 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Siemens | 2.46
Siemens Atellica CH- Immunoassay | 11 | 7.86 | 0.05 | 3.6 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Siemens | 3.65
Siemens Atellica CH- Immunoassay | 11 | 5.85 | -0.06 | 3.2 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Siemens | 3.14
Siemens Atellica CH- Immunoassay | 11 | 8.71 | 0.02 | 2.6 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Siemens | 2.62
Siemens Atellica CH- Immunoassay | 11 | 5.92 | 0.06 | 4.2 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Siemens | 4.26
Siemens Atellica CH- Immunoassay | 11 | 9.65 | 0.09 | 3.4 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Siemens | 3.49
Siemens DCA 2000 | 187 | 13.7 |  | 4 | GH2-01 | 13 | 2008 | CAP | Fresh | Siemens | 4
Siemens DCA 2000 | 187 | 5.2 |  | 2.8 | GH2-02 | 5.3 | 2008 | CAP | Fresh | Siemens | 2.8
Siemens DCA 2000 | 187 | 8.5 |  | 2.9 | GH2-03 | 8.8 | 2008 | CAP | Fresh | Siemens | 2.9
Siemens DCA 2000 | 178 | 12 |  | 4.9 | GH2-04 | 11.9 | 2008 | CAP | Fresh | Siemens | 4.9
Siemens DCA 2000 | 178 | 6.8 |  | 2.6 | GH2-05 | 6.8 | 2008 | CAP | Fresh | Siemens | 2.6
Siemens DCA 2000 | 178 | 5.1 |  | 2.6 | GH2-06 | 5 | 2008 | CAP | Fresh | Siemens | 2.6
Siemens DCA 2000/200+ | 134 | 5.1 |  | 2.9 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Siemens | 2.9
Siemens DCA 2000/200+ | 134 | 8.4 |  | 3.3 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Siemens | 3.3
Siemens DCA 2000/200+ | 134 | 6 |  | 3.3 | GH2-03 | 6 | 2009 | CAP | Fresh | Siemens | 3.3
Siemens DCA 2000/200+ | 106 | 6.81 | 0.21 | 2.5 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Siemens | 2.71
Siemens DCA 2000/200+ | 106 | 9.48 | -0.02 | 3 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Siemens | 2.98
Siemens DCA 2000/200+ | 106 | 7.56 | 0.16 | 3.1 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Siemens | 3.26
Siemens DCA 2000/200+ | 84 | 6.07 | 0.17 | 2.8 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Siemens | 2.97
Siemens DCA 2000/200+ | 84 | 9.66 | -0.14 | 3.2 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Siemens | 3.06
Siemens DCA 2000/200+ | 84 | 7.5 | 0.1 | 2.4 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Siemens | 2.5
Siemens DCA 2000/200+ | 75 | 5.29 | 0.09 | 2.6 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Siemens | 2.69
Siemens DCA 2000/200+ | 75 | 8.71 | 0.01 | 3.2 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Siemens | 3.21
Siemens DCA 2000/200+ | 75 | 6.45 | 0.15 | 2.6 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Siemens | 2.75
Siemens DCA 2000/2000+ | 66 | 8.50 | 0 | 3.1 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Siemens | 3.1
Siemens DCA 2000/2000+ | 66 | 5.47 | 0.07 | 2.9 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Siemens | 2.97
Siemens DCA 2000/2000+ | 66 | 6.54 | 0.14 | 2.7 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Siemens | 2.84
Siemens DCA 2000/2000+ | 19 | 6.58 | 0.28 | 1.8 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Siemens | 2.08
Siemens DCA 2000/2000+ | 19 | 7.95 | 0.35 | 2.6 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Siemens | 2.95
Siemens DCA 2000/2000+ | 19 | 9.52 | 0.32 | 2.4 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Siemens | 2.72
Siemens DCA 2000/2000+ | 71 | 5.61 | 0.01 | 2.7 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Siemens | 2.71
Siemens DCA 2000/2000+ | 71 | 9.25 | -0.15 | 3.2 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Siemens | 3.05
Siemens DCA 2000/2000+ | 71 | 7.19 | -0.01 | 3.3 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Siemens | 3.29
Siemens DCA 2000/2000+ | 58 | 5.46 | 0.06 | 2.8 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Siemens | 2.86
Siemens DCA 2000/2000+ | 58 | 8.28 | -0.02 | 3 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Siemens | 2.98
Siemens DCA 2000/2000+ | 58 | 5.77 | 0.12 | 3.1 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Siemens | 3.22
Siemens DCA 2000/2000+ | 39 | 7.08 | -0.03 | 3.6 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Siemens | 3.57
Siemens DCA 2000/2000+ | 39 | 9.41 | 0.09 | 3.1 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Siemens | 3.19
Siemens DCA 2000/2000+ | 39 | 6.08 | 0.01 | 3.1 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Siemens | 3.11
Siemens DCA 2000/2000+ | 37 | 7.92 | -0.13 | 2.7 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Siemens | 2.57
Siemens DCA 2000/2000+ | 37 | 5.24 | -0.06 | 2.5 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Siemens | 2.44
Siemens DCA 2000/2000+ | 37 | 6.09 | -0.05 | 2.3 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Siemens | 2.25
Siemens DCA 2000/2000+ | 29 | 6.51 | -0.07 | 3.2 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Siemens | 3.13
Siemens DCA 2000/2000+ | 29 | 8.21 | -0.18 | 3.3 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Siemens | 3.12
Siemens DCA 2000/2000+ | 29 | 5.52 | -0.13 | 3.4 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Siemens | 3.27
Siemens DCA 2000/2000+ | 30 | 6.6 | 0.11 | 3 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Siemens | 3.11
Siemens DCA 2000/2000+ | 30 | 7.08 | 0.11 | 2.7 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Siemens | 2.81
Siemens DCA 2000/2000+ | 30 | 9.68 | 0.03 | 3.5 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Siemens | 3.53
Siemens DCA 2000/2000+ | 26 | 6.67 | -0.12 | 2.7 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Siemens | 2.58
Siemens DCA 2000/2000+ | 26 | 10.21 | -0.07 | 3.2 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Siemens | 3.13
Siemens DCA 2000/2000+ | 26 | 6.67 | -0.15 | 2.9 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Siemens | 2.75
Siemens DCA 2000/2000+ | <10 |  |  |  | GH5-04 | 8.63 | 2015 | CAP | Fresh | Siemens | 0
Siemens DCA 2000/2000+ | <10 |  |  |  | GH5-05 | 5.32 | 2015 | CAP | Fresh | Siemens | 0
Siemens DCA 2000/2000+ | 6-20 | 12.25 | 0.56 | 6.9 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Siemens | 7.46
Siemens DCA 2000/2000+ | 6-20 | 5.83 | -0.1 | 2.5 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Siemens | 2.4
Siemens DCA 2000/2000+ | 6-20 | 5.16 | -0.01 | 2.7 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Siemens | 2.69
Siemens DCA 2000/2000+ | 6-20 |  |  |  | GH5-14 | 8.14 | 2015 | CAP | Fresh | Siemens | 0
Siemens DCA 2000/2000+ | 6-20 |  |  |  | GH5-15 | 9.3 | 2015 | CAP | Fresh | Siemens | 0
Siemens DCA 2000/2000+ | 8-26 | 5.33 | 0.01 | 3.3 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Siemens | 3.31
Siemens DCA 2000/2000+ | 8-26 | 9.20 | 0.04 | 2.9 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Siemens | 2.94
Siemens DCA 2000/2000+ | 8-26 | 5.29 | -0.02 | 2.2 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Siemens | 2.18
Siemens DCA 2000/2000+ | 8-26 |  |  |  | GH5-04 | 12.03 | 2016 | CAP | Fresh | Siemens | 0
Siemens DCA 2000/2000+ | 8-26 |  |  |  | GH5-05 | 5.94 | 2016 | CAP | Fresh | Siemens | 0
Siemens DCA 2000/2000+ | 5-21 | 9.22 | 0.11 | 3.4 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Siemens | 3.51
Siemens DCA 2000/2000+ | 5-21 | 6.12 | 0.11 | 4 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Siemens | 4.11
Siemens DCA 2000/2000+ | 5-21 | 12.17 | 0.46 | 3.6 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Siemens | 4.06
Siemens DCA 2000/2000+ | 5-21 |  |  |  | GH5-14 | 5.02 | 2016 | CAP | Fresh | Siemens | 0
Siemens DCA 2000/2000+ | 5-21 |  |  |  | GH5-15 | 7.58 | 2016 | CAP | Fresh | Siemens | 0
Siemens DCA 2000/2000+ | 0-14 | 6.54 | 0.13 | 3 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Siemens | 0.13
Siemens DCA 2000/2000+ | 0-14 | 9.59 | 0.06 | 3.7 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Siemens | 3.76
Siemens DCA 2000/2000+ | 0-14 | 5.4 | 0.06 | 2.4 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Siemens | 2.46
Siemens DCA 2000/2000+ | 0-14 |  |  |  | GH5-04 | 8.51 | 2017 | CAP | Fresh | Siemens | 0
Siemens DCA 2000/2000+ | 0-14 |  |  |  | GH5-05 | 7.25 | 2017 | CAP | Fresh | Siemens | 0
Siemens DCA 2000/2000+ | 10 | 7.43 | 0.18 | 3 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Siemens | 0.18
Siemens DCA 2000/2000+ | 10 | 7.08 | 0.27 | 3.2 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Siemens | 3.47
Siemens DCA 2000/2000+ | 10 | 8.96 | 0.31 | 3.1 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Siemens | 3.41
Siemens DCA 2000/2000+ | 10 |  |  |  | GH5-14 | 9.50 | 2017 | CAP | Fresh | Siemens | 0
Siemens DCA 2000/2000+ | 10 |  |  |  | GH5-15 | 5.45 | 2017 | CAP | Fresh | Siemens | 0
Siemens DCA 2000/2000+ | 11 | 7.17 | 0.02 | 4.2 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Siemens | 4.22
Siemens DCA 2000/2000+ | 11 | 5.27 | 0.08 | 4.5 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Siemens | 4.58
Siemens DCA 2000/2000+ | 11 | 8.47 | 0.05 | 2.9 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Siemens | 2.95
Siemens DCA 2000/2000+ | 11 | - | - |  | GH5-04 | 9.79 | 2018 | CAP | Fresh | Siemens | 0
Siemens DCA 2000/2000+ | 11 | - | - |  | GH5-05 | 6.12 | 2018 | CAP | Fresh | Siemens | 0
Siemens DCA Vantage | 59 | 5.1 |  | 2.3 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Siemens | 2.3
Siemens DCA Vantage | 59 | 8.4 |  | 2.6 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Siemens | 2.6
Siemens DCA Vantage | 59 | 6 |  | 2.9 | GH2-03 | 6 | 2009 | CAP | Fresh | Siemens | 2.9
Siemens DCA Vantage | 97 | 6.78 | 0.18 | 2.8 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Siemens | 2.98
Siemens DCA Vantage | 97 | 9.39 | -0.11 | 3.4 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Siemens | 3.29
Siemens DCA Vantage | 97 | 7.55 | 0.15 | 2.5 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Siemens | 2.65
Siemens DCA Vantage | 118 | 6.07 | 0.17 | 2.2 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Siemens | 2.37
Siemens DCA Vantage | 118 | 9.62 | -0.18 | 3.2 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Siemens | 3.02
Siemens DCA Vantage | 118 | 7.44 | 0.04 | 2.6 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Siemens | 2.64
Siemens DCA Vantage | 115 | 5.33 | 0.13 | 2.7 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Siemens | 2.83
Siemens DCA Vantage | 115 | 8.63 | -0.07 | 2.8 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Siemens | 2.73
Siemens DCA Vantage | 115 | 6.4 | 0.1 | 2.2 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Siemens | 2.3
Siemens DCA Vantage | 162 | 8.39 | -0.11 | 2.7 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Siemens | 2.59
Siemens DCA Vantage | 162 | 5.48 | 0.08 | 2.4 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Siemens | 2.48
Siemens DCA Vantage | 162 | 6.52 | 0.12 | 2.5 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Siemens | 2.62
Siemens DCA Vantage | 59 | 6.26 | -0.04 | 3.1 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Siemens | 3.06
Siemens DCA Vantage | 59 | 7.66 | 0.06 | 3.5 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Siemens | 3.56
Siemens DCA Vantage | 59 | 9.17 | -0.03 | 2.9 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Siemens | 2.87
Siemens DCA Vantage | 253 | 5.56 | -0.04 | 2.7 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Siemens | 2.66
Siemens DCA Vantage | 253 | 9.18 | -0.22 | 3.1 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Siemens | 2.88
Siemens DCA Vantage | 253 | 7.14 | -0.06 | 2.7 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Siemens | 2.64
Siemens DCA Vantage | 261 | 5.44 | 0.04 | 2.6 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Siemens | 2.64
Siemens DCA Vantage | 261 | 8.17 | -0.13 | 2.6 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Siemens | 2.47
Siemens DCA Vantage | 261 | 5.78 | 0.13 | 2.8 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Siemens | 2.93
Siemens DCA Vantage | 235 | 6.99 | -0.12 | 2.8 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Siemens | 2.68
Siemens DCA Vantage | 235 | 9.26 | -0.06 | 3 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Siemens | 2.94
Siemens DCA Vantage | 235 | 6.05 | -0.02 | 2.6 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Siemens | 2.58
Siemens DCA Vantage | 258 | 7.84 | -0.21 | 2.7 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Siemens | 2.49
Siemens DCA Vantage | 258 | 5.21 | -0.09 | 2.3 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Siemens | 2.21
Siemens DCA Vantage | 258 | 6.03 | -0.11 | 2.5 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Siemens | 2.39
Siemens DCA Vantage | 307 | 6.44 | -0.14 | 2.8 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Siemens | 2.66
Siemens DCA Vantage | 307 | 8.14 | -0.25 | 3 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Siemens | 2.75
Siemens DCA Vantage | 307 | 5.45 | -0.2 | 3.2 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Siemens | 3
Siemens DCA Vantage | 305 | 6.46 | -0.03 | 2.6 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Siemens | 2.57
Siemens DCA Vantage | 305 | 6.95 | -0.02 | 2.8 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Siemens | 2.78
Siemens DCA Vantage | 305 | 9.61 | -0.04 | 3.2 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Siemens | 3.16
Siemens DCA Vantage | 340 | 6.6 | -0.19 | 2.7 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Siemens | 2.51
Siemens DCA Vantage | 340 | 9.9 | -0.38 | 3.3 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Siemens | 2.92
Siemens DCA Vantage | 340 | 6.61 | -0.21 | 2.7 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Siemens | 2.49
Siemens DCA Vantage | 145 | 8.31 | -0.32 | 2.8 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Siemens | 2.48
Siemens DCA Vantage | 145 | 5.22 | -0.1 | 2.7 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Siemens | 2.6
Siemens DCA Vantage | 168 | 5.74 | -0.1 | 2.3 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Siemens | 2.2
Siemens DCA Vantage | 168 | 11.91 | 0.2 | 5.2 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Siemens | 5.4
Siemens DCA Vantage | 168 | 9.2 | -0.33 | 3.2 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Siemens | 2.87
Siemens DCA Vantage | 168 | 4.97 | -0.07 | 2.7 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Siemens | 2.63
Siemens DCA Vantage | 168 | 7.13 | -0.25 | 2.6 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Siemens | 2.35
Siemens DCA Vantage | 171-360 | 11.94 | 0.25 | 6.1 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Siemens | 6.35
Siemens DCA Vantage | 171-360 | 5.77 | -0.16 | 2.5 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Siemens | 2.34
Siemens DCA Vantage | 171-360 | 5.1 | -0.07 | 2.7 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Siemens | 2.63
Siemens DCA Vantage | 171-360 | 7.89 | -0.25 | 2.6 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Siemens | 2.35
Siemens DCA Vantage | 171-360 | 8.99 | -0.31 | 3.5 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Siemens | 3.19
Siemens DCA Vantage | 189-473 | 5.25 | -0.07 | 2.2 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Siemens | 2.13
Siemens DCA Vantage | 189-473 | 9.07 | -0.09 | 2.8 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Siemens | 2.71
Siemens DCA Vantage | 189-473 | 5.27 | -0.04 | 2.5 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Siemens | 2.46
Siemens DCA Vantage | 189-473 | 12.26 | 0.23 | 4.7 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Siemens | 4.93
Siemens DCA Vantage | 189-473 | 5.86 | -0.08 | 2 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Siemens | 1.92
Siemens DCA Vantage | 186 | 5.20 | -0.07 | 2.2 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Siemens | 2.13
Siemens DCA Vantage | 186 | 10.44 | -0.15 | 4 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Siemens | 3.85
Siemens DCA Vantage | 186 | 6.06 | -0.14 | 2.2 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Siemens | 2.06
Siemens DCA Vantage | 186 | 12.57 | 0.34 | 5.6 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Siemens | 5.94
Siemens DCA Vantage | 186 | 7.30 | -0.21 | 2.5 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Siemens | 2.29
Siemens DCA Vantage | 176-460 | 9.05 | -0.06 | 2.6 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Siemens | 2.54
Siemens DCA Vantage | 176-460 | 5.96 | -0.05 | 2.2 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Siemens | 2.15
Siemens DCA Vantage | 176-460 | 11.92 | 0.21 | 4.3 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Siemens | 4.51
Siemens DCA Vantage | 176-460 | 5.00 | -0.02 | 2.5 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Siemens | 2.48
Siemens DCA Vantage | 176-460 | 7.55 | -0.03 | 2.6 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Siemens | 2.57
Siemens DCA Vantage | 150-483 | 6.45 | 0.04 | 2.3 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Siemens | 2.34
Siemens DCA Vantage | 150-483 | 9.46 | -0.07 | 2.8 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Siemens | 2.73
Siemens DCA Vantage | 150-483 | 5.37 | 0.03 | 2.5 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Siemens | 2.53
Siemens DCA Vantage | 150-483 | 8.49 | -0.02 | 2.4 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Siemens | 2.38
Siemens DCA Vantage | 150-483 | 7.31 | 0.06 | 2 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Siemens | 2.06
Siemens DCA Vantage | 148 | 7.60 | 0.18 | 2.8 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Siemens | 2.98
Siemens DCA Vantage | 148 | 5.23 | 0.03 | 2.7 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Siemens | 2.73
Siemens DCA Vantage | 148 | 8.48 | 0.17 | 2.7 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Siemens | 2.87
Siemens DCA Vantage | 148 | 10.61 | 0.28 | 3.6 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Siemens | 3.88
Siemens DCA Vantage | 148 | 6.02 | 0.15 | 2.6 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Siemens | 2.75
Siemens DCA Vantage | 435 | 7.31 | 0.06 | 2.8 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Siemens | 2.86
Siemens DCA Vantage | 435 | 6.93 | 0.12 | 2.8 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Siemens | 2.92
Siemens DCA Vantage | 435 | 8.73 | 0.08 | 2.9 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Siemens | 2.98
Siemens DCA Vantage | 435 | 9.74 | 0.24 | 3.5 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Siemens | 3.74
Siemens DCA Vantage | 435 | 5.48 | 0.03 | 2.3 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Siemens | 2.33
Siemens DCA Vantage | 429 | 7.17 | 0.02 | 2.4 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Siemens | 2.42
Siemens DCA Vantage | 429 | 5.29 | 0.1 | 2.7 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Siemens | 2.8
Siemens DCA Vantage | 429 | 8.43 | 0.01 | 2.8 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Siemens | 2.81
Siemens DCA Vantage | 429 | 9.87 | 0.08 | 3.3 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Siemens | 3.38
Siemens DCA Vantage | 429 | 6.16 | 0.04 | 2.6 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Siemens | 2.64
Siemens DCA Vantage | 129 | 6.37 | 0.06 | 2.8 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Siemens | 2.86
Siemens DCA Vantage | 129 | 9.24 | 0.13 | 3.4 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Siemens | 3.53
Siemens DCA Vantage | 129 | 5.37 | 0.06 | 3.2 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Siemens | 3.26
Siemens DCA Vantage | 129 | 8.48 | 0.13 | 2.7 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Siemens | 2.83
Siemens DCA Vantage | 129 | 7.73 | 0.11 | 2.7 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Siemens | 2.81
Siemens DCA Vantage | 410 | 9.29 | -0.08 | 3.1 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Siemens | 3.02
Siemens DCA Vantage | 410 | 5.03 | 0.11 | 2.5 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Siemens | 2.61
Siemens DCA Vantage | 410 | 6.11 | 0.07 | 2.5 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Siemens | 2.57
Siemens DCA Vantage | 410 | 5.29 | 0.01 | 2.2 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Siemens | 2.21
Siemens DCA Vantage | 410 | 8.08 | 0 | 2.3 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Siemens | 2.3
Siemens DCA Vantage | 392 | 5.46 | 0 | 3 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Siemens | 3
Siemens DCA Vantage | 392 | 5.98 | 0.32 | 2.9 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Siemens | 3.22
Siemens DCA Vantage | 392 | 9.36 | 0.05 | 3.3 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Siemens | 3.35
Siemens DCA Vantage | 392 | 5.22 | -0.06 | 2.6 | GH5-04 | 5.28 | 2019 | CAP | Fresh | Siemens | 2.54
Siemens DCA Vantage | 392 | 7.40 | -0.01 | 2.5 | GH5-05 | 7.41 | 2019 | CAP | Fresh | Siemens | 2.49
Siemens DCA Vantage | 132 | 6.37 | -0.04 | 2.4 | GH5-06 | 6.41 | 2019 | CAP | Fresh | Siemens | 2.36
Siemens DCA Vantage | 132 | 8.52 | -0.08 | 3 | GH5-07 | 8.6 | 2019 | CAP | Fresh | Siemens | 2.92
Siemens DCA Vantage | 132 | 5.37 | -0.05 | 2.6 | GH5-08 | 5.42 | 2019 | CAP | Fresh | Siemens | 2.55
Siemens DCA Vantage | 132 | 7.41 | 0.03 | 2.6 | GH5-09 | 7.38 | 2019 | CAP | Fresh | Siemens | 2.63
Siemens DCA Vantage | 132 | 9.74 | -0.01 | 3.2 | GH5-10 | 9.75 | 2019 | CAP | Fresh | Siemens | 3.19
Siemens DCA Vantage | 395 | 5.16 | -0.09 | 2.4 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Siemens | 2.31
Siemens DCA Vantage | 395 | 6.38 | -0.03 | 2.4 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Siemens | 2.37
Siemens DCA Vantage | 395 | 8.23 | 0.02 | 2.4 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Siemens | 2.42
Siemens DCA Vantage | 395 | 5.43 | -0.04 | 2.6 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Siemens | 2.56
Siemens DCA Vantage | 395 | 9.48 | 0 | 3.3 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Siemens | 3.3
Siemens DCA Vantage | 305 | 8.13 | 0.16 | 3.1 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Siemens | 3.26
Siemens DCA Vantage | 305 | 5.91 | 0.02 | 2.5 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Siemens | 2.52
Siemens DCA Vantage | 305 | 5.11 | -0.02 | 2.4 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Siemens | 2.38
Siemens DCA Vantage | 305 | 7.5 | 0.1 | 2.4 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Siemens | 2.5
Siemens DCA Vantage | 305 | 9.07 | -0.1 | 2.7 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Siemens | 2.6
Siemens DCA Vantage | 72 | 7.74 | -0.07 | 2.8 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Siemens | 2.73
Siemens DCA Vantage | 72 | 5.88 | -0.03 | 2.3 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Siemens | 2.27
Siemens DCA Vantage | 72 | 8.63 | -0.06 | 2.6 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Siemens | 2.54
Siemens DCA Vantage | 72 | 5.8 | -0.06 | 2.6 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Siemens | 2.54
Siemens DCA Vantage | 72 | 9.45 | -0.11 | 3.6 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Siemens | 3.49
Siemens DCA Vantage | 331 | 5.43 | -0.05 | 2.5 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Siemens | 2.45
Siemens DCA Vantage | 331 | 8.31 | 0.05 | 2.7 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Siemens | 2.75
Siemens DCA Vantage | 331 | 5.09 | -0.06 | 2.6 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Siemens | 2.54
Siemens DCA Vantage | 331 | 10.19 | 0.14 | 3.9 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Siemens | 4.04
Siemens DCA Vantage | 331 | 5.93 | -0.06 | 2.6 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Siemens | 2.54
Siemens DCA Vantage | 321 | 8.88 | 0.02 | 2.8 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Siemens | 2.82
Siemens DCA Vantage | 321 | 6.84 | 0 | 2.3 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Siemens | 2.3
Siemens DCA Vantage | 321 | 9.46 | 0.07 | 3.1 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Siemens | 3.17
Siemens DCA Vantage | 321 | 6.09 | -0.04 | 2.5 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Siemens | 2.46
Siemens DCA Vantage | 321 | 7.47 | -0.05 | 2.4 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Siemens | 2.35
Siemens DCA Vantage | 80 | 5.76 | -0.04 | 2.2 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Siemens | 2.16
Siemens DCA Vantage | 80 | 9.22 | 0.1 | 2.8 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Siemens | 2.9
Siemens DCA Vantage | 80 | 5.41 | -0.06 | 2.2 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Siemens | 2.14
Siemens DCA Vantage | 80 | 7.33 | 0.08 | 2.2 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Siemens | 2.28
Siemens DCA Vantage | 80 | 9.31 | 0.05 | 2.5 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Siemens | 2.55
Siemens DCA Vantage | 306 | 5.13 | -0.08 | 2.8 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Siemens | 2.72
Siemens DCA Vantage | 306 | 6.64 | -0.04 | 2.6 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Siemens | 2.56
Siemens DCA Vantage | 306 | 8.26 |  | 3.3 | GH5-13 |  | 2021 | CAP | Fresh | Siemens | 3.3
Siemens DCA Vantage | 306 | 5.97 | -0.01 | 2.4 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Siemens | 2.39
Siemens DCA Vantage | 306 | 10.11 | -0.01 | 3.7 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Siemens | 3.69
Siemens Dimension ExL | 214 | 5.62 | 0.16 | 4.1 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Siemens | 4.26
Siemens Dimension ExL | 214 | 5.94 | 0.28 | 3.7 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Siemens | 3.98
Siemens Dimension ExL | 214 | 9.03 | -0.28 | 2.2 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Siemens | 1.92
Siemens Dimension ExL | 214 | 5.41 | 0.13 | 4 | GH5-04 | 5.28 | 2019 | CAP | Fresh | Siemens | 4.13
Siemens Dimension ExL | 214 | 7.43 | 0.02 | 2.9 | GH5-05 | 7.41 | 2019 | CAP | Fresh | Siemens | 2.92
Siemens Dimension ExL | 12 | 5.4 |  | 3.7 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Siemens | 3.7
Siemens Dimension ExL | 12 | 8.4 |  | 3.6 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Siemens | 3.6
Siemens Dimension ExL | 12 | 6.2 |  | 3.1 | GH2-03 | 6 | 2009 | CAP | Fresh | Siemens | 3.1
Siemens Dimension ExL | 25 | 6.62 | 0.02 | 4.6 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Siemens | 4.62
Siemens Dimension ExL | 25 | 9.32 | -0.18 | 5 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Siemens | 4.82
Siemens Dimension ExL | 25 | 7.28 | -0.12 | 5.2 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Siemens | 5.08
Siemens Dimension ExL | 39 | 6.03 | 0.13 | 2.2 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Siemens | 2.33
Siemens Dimension ExL | 39 | 9.6 | -0.2 | 4 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Siemens | 3.8
Siemens Dimension ExL | 39 | 7.19 | -0.21 | 2.4 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Siemens | 2.19
Siemens Dimension ExL | 51 | 5.48 | 0.28 | 3.5 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Siemens | 3.78
Siemens Dimension ExL | 51 | 8.42 | -0.28 | 3.7 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Siemens | 3.42
Siemens Dimension ExL | 51 | 6.33 | 0.03 | 3.1 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Siemens | 3.13
Siemens Dimension ExL | 59 | 8.55 | 0.05 | 3 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Siemens | 3.05
Siemens Dimension ExL | 59 | 5.72 | 0.32 | 2.5 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Siemens | 2.82
Siemens Dimension ExL | 59 | 6.55 | 0.15 | 2.9 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Siemens | 3.05
Siemens Dimension ExL | 131 | 7.14 | 0.03 | 2.2 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Siemens | 2.23
Siemens Dimension ExL | 131 | 9.44 | 0.12 | 2.6 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Siemens | 2.72
Siemens Dimension ExL | 131 | 6.29 | 0.22 | 2.9 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Siemens | 3.12
Siemens Dimension ExL | 173 | 8.14 | 0.09 | 3 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Siemens | 3.09
Siemens Dimension ExL | 173 | 5.46 | 0.16 | 3.7 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Siemens | 3.86
Siemens Dimension ExL | 173 | 6.39 | 0.25 | 3.2 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Siemens | 3.45
Siemens Dimension ExL | 185 | 6.76 | 0.18 | 2.6 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Siemens | 2.78
Siemens Dimension ExL | 185 | 8.35 | -0.04 | 2.9 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Siemens | 2.86
Siemens Dimension ExL | 185 | 5.78 | 0.13 | 3.5 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Siemens | 3.63
Siemens Dimension ExL | 188 | 6.85 | 0.36 | 3.3 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Siemens | 3.66
Siemens Dimension ExL | 188 | 7.25 | 0.28 | 3.4 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Siemens | 3.68
Siemens Dimension ExL | 188 | 9.64 | -0.01 | 2.4 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Siemens | 2.39
Siemens Dimension ExL | 193 | 6.89 | 0.1 | 3.5 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Siemens | 3.6
Siemens Dimension ExL | 193 | 10.23 | -0.05 | 3.4 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Siemens | 3.35
Siemens Dimension ExL | 193 | 6.92 | 0.1 | 3.2 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Siemens | 3.3
Siemens Dimension ExL | 128 | 8.65 | 0.02 | 2.9 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Siemens | 2.92
Siemens Dimension ExL | 128 | 5.43 | 0.11 | 3.7 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Siemens | 3.81
Siemens Dimension ExL | 126 | 5.94 | 0.1 | 3.7 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Siemens | 3.8
Siemens Dimension ExL | 126 | 11.35 | -0.36 | 3 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Siemens | 2.64
Siemens Dimension ExL | 126 | 9.34 | -0.19 | 3.2 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Siemens | 3.01
Siemens Dimension ExL | 126 | 5.16 | 0.12 | 4 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Siemens | 4.12
Siemens Dimension ExL | 126 | 7.25 | -0.13 | 4 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Siemens | 3.87
Siemens Dimension ExL | 134-205 | 11.26 | -0.43 | 3.1 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Siemens | 2.67
Siemens Dimension ExL | 134-205 | 6.03 | 0.1 | 3.2 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Siemens | 3.3
Siemens Dimension ExL | 134-205 | 5.3 | 0.13 | 3.7 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Siemens | 3.83
Siemens Dimension ExL | 134-205 | 7.97 | -0.17 | 2.5 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Siemens | 2.33
Siemens Dimension ExL | 134-205 | 9.02 | -0.28 | 2.2 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Siemens | 1.92
Siemens Dimension ExL | 132-205 | 5.41 | 0.09 | 3.1 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Siemens | 3.19
Siemens Dimension ExL | 132-205 | 9.12 | -0.04 | 2.7 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Siemens | 2.66
Siemens Dimension ExL | 132-205 | 5.41 | 0.1 | 3.2 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Siemens | 3.3
Siemens Dimension ExL | 132-205 | 11.80 | -0.23 | 3.1 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Siemens | 2.87
Siemens Dimension ExL | 132-205 | 5.98 | 0.04 | 3.4 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Siemens | 3.44
Siemens Dimension ExL | 134 | 5.31 | 0.04 | 2.9 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Siemens | 2.94
Siemens Dimension ExL | 134 | 10.24 | -0.35 | 2.2 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Siemens | 1.85
Siemens Dimension ExL | 134 | 6.17 | -0.03 | 3 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Siemens | 2.97
Siemens Dimension ExL | 134 | 11.64 | -0.59 | 2.7 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Siemens | 2.11
Siemens Dimension ExL | 134 | 7.44 | -0.07 | 2.6 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Siemens | 2.53
Siemens Dimension ExL | 142-209 | 9.02 | -0.09 | 2.5 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Siemens | 2.41
Siemens Dimension ExL | 142-209 | 6.06 | 0.05 | 2.3 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Siemens | 2.35
Siemens Dimension ExL | 142-209 | 11.39 | -0.32 | 2.3 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Siemens | 1.98
Siemens Dimension ExL | 142-209 | 5.10 | 0.08 | 3.4 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Siemens | 3.48
Siemens Dimension ExL | 142-209 | 7.53 | -0.05 | 3.1 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Siemens | 3.05
Siemens Dimension ExL | 140-209 | 6.62 | 0.21 | 3.6 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Siemens | 3.81
Siemens Dimension ExL | 140-209 | 9.58 | 0.05 | 2.7 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Siemens | 2.75
Siemens Dimension ExL | 140-209 | 5.46 | 0.12 | 3.7 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Siemens | 3.82
Siemens Dimension ExL | 140-209 | 8.65 | 0.14 | 2.8 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Siemens | 2.94
Siemens Dimension ExL | 140-209 | 7.35 | 0.1 | 2.7 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Siemens | 2.8
Siemens Dimension ExL | 207 | 7.19 | 0.04 | 2.9 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Siemens | 2.94
Siemens Dimension ExL | 207 | 5.26 | 0.07 | 3.6 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Siemens | 3.67
Siemens Dimension ExL | 207 | 8.16 | -0.26 | 2.6 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Siemens | 2.34
Siemens Dimension ExL | 207 | 9.47 | -0.32 | 2.6 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Siemens | 2.28
Siemens Dimension ExL | 207 | 6.16 | 0.04 | 3.4 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Siemens | 3.44
Siemens Dimension ExL | 148 | 6.39 | 0.08 | 3.1 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Siemens | 3.18
Siemens Dimension ExL | 148 | 8.89 | -0.22 | 2.6 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Siemens | 2.38
Siemens Dimension ExL | 148 | 5.40 | 0.09 | 2.7 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Siemens | 2.79
Siemens Dimension ExL | 148 | 8.37 | 0.02 | 2.5 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Siemens | 2.52
Siemens Dimension ExL | 148 | 7.47 | -0.15 | 2.5 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Siemens | 2.35
Siemens Dimension ExL | 210 | 9.20 | -0.17 | 2.1 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Siemens | 1.93
Siemens Dimension ExL | 210 | 5.10 | 0.18 | 3.1 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Siemens | 3.28
Siemens Dimension ExL | 210 | 6.21 | 0.17 | 2.7 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Siemens | 2.87
Siemens Dimension ExL | 210 | 5.46 | 0.18 | 2.8 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Siemens | 2.98
Siemens Dimension ExL | 210 | 8.08 | 0 | 2.6 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Siemens | 2.6
Siemens Dimension ExL | 149 | 6.48 | 0.07 | 2.8 | GH5-06 | 6.41 | 2019 | CAP | Fresh | Siemens | 2.87
Siemens Dimension ExL | 149 | 8.50 | -0.1 | 2.7 | GH5-07 | 8.6 | 2019 | CAP | Fresh | Siemens | 2.6
Siemens Dimension ExL | 149 | 5.51 | 0.09 | 2.5 | GH5-08 | 5.42 | 2019 | CAP | Fresh | Siemens | 2.59
Siemens Dimension ExL | 149 | 7.39 | 0.01 | 2.9 | GH5-09 | 7.38 | 2019 | CAP | Fresh | Siemens | 2.91
Siemens Dimension ExL | 149 | 9.58 | -0.17 | 2.6 | GH5-10 | 9.75 | 2019 | CAP | Fresh | Siemens | 2.43
Siemens Dimension ExL | 197 | 5.31 | 0.06 | 2.8 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Siemens | 2.86
Siemens Dimension ExL | 197 | 6.57 | 0.16 | 2.5 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Siemens | 2.66
Siemens Dimension ExL | 197 | 8.08 | -0.13 | 2.2 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Siemens | 2.07
Siemens Dimension ExL | 197 | 5.54 | 0.07 | 2.8 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Siemens | 2.87
Siemens Dimension ExL | 197 | 9.28 | -0.2 | 2.3 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Siemens | 2.1
Siemens Dimension ExL | 189 | 8.03 | 0.06 | 2 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Siemens | 0.06
Siemens Dimension ExL | 189 | 6.07 | 0.18 | 2.4 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Siemens | 2.58
Siemens Dimension ExL | 189 | 5.27 | 0.14 | 2.9 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Siemens | 3.04
Siemens Dimension ExL | 189 | 7.44 | 0.04 | 2.2 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Siemens | 2.24
Siemens Dimension ExL | 189 | 9.02 | -0.15 | 2 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Siemens | -0.15
Siemens Dimension ExL | 144 | 7.65 | -0.16 | 2.3 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Siemens | 2.14
Siemens Dimension ExL | 144 | 6 | 0.09 | 2.9 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Siemens | 2.99
Siemens Dimension ExL | 144 | 8.48 | -0.21 | 2 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Siemens | -0.21
Siemens Dimension ExL | 144 | 6.04 | 0.18 | 3.2 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Siemens | 3.38
Siemens Dimension ExL | 144 | 9.31 | -0.25 | 2.4 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Siemens | 2.15
Siemens Dimension ExL | 189 | 5.54 | 0.06 | 2.5 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Siemens | 2.56
Siemens Dimension ExL | 189 | 8.06 | -0.2 | 2 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Siemens | -0.2
Siemens Dimension ExL | 189 | 5.19 | 0.04 | 2.9 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Siemens | 2.94
Siemens Dimension ExL | 189 | 9.83 | -0.22 | 2.1 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Siemens | 1.88
Siemens Dimension ExL | 189 | 6.04 | 0.05 | 2.4 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Siemens | 2.45
Siemens Dimension ExL | 185 | 8.86 | 0 | 2.2 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Siemens | 2.2
Siemens Dimension ExL | 185 | 6.88 | 0.04 | 2.2 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Siemens | 2.24
Siemens Dimension ExL | 185 | 9.24 | -0.15 | 2.5 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Siemens | 2.35
Siemens Dimension ExL | 185 | 6.17 | 0.04 | 2.2 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Siemens | 2.24
Siemens Dimension ExL | 185 | 7.52 | 0 | 2.4 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Siemens | 2.4
Siemens Dimension ExL | 136 | 5.91 | 0.11 | 3.1 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Siemens | 3.21
Siemens Dimension ExL | 136 | 8.97 | -0.15 | 2 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Siemens | -0.15
Siemens Dimension ExL | 136 | 5.54 | 0.07 | 2.2 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Siemens | 2.27
Siemens Dimension ExL | 136 | 7.31 | 0.06 | 2.4 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Siemens | 2.46
Siemens Dimension ExL | 136 | 9.15 | -0.11 | 2.1 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Siemens | 1.99
Siemens Dimension ExL | 185 | 5.27 | 0.06 | 2.5 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Siemens | 2.56
Siemens Dimension ExL | 185 | 6.7 | 0.02 | 1.9 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Siemens | 1.92
Siemens Dimension ExL | 185 | 8.18 |  | 1.8 | GH5-13 |  | 2021 | CAP | Fresh | Siemens | 1.8
Siemens Dimension ExL | 185 | 6.1 | 0.12 | 1.9 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Siemens | 2.02
Siemens Dimension ExL | 185 | 9.92 | -0.2 | 1.9 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Siemens | 1.7
Siemens Dimension ExL | 144 | 7.61 | 0.19 | 3 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Siemens | 0.19
Siemens Dimension ExL | 144 | 5.30 | 0.1 | 3.1 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Siemens | 3.2
Siemens Dimension ExL | 144 | 8.28 | -0.03 | 2.4 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Siemens | 2.37
Siemens Dimension ExL | 144 | 10.48 | 0.15 | 2.4 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Siemens | 2.55
Siemens Dimension ExL | 144 | 6.03 | 0.16 | 3.6 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Siemens | 3.76
Siemens Dimension ExL | 206 | 7.29 | 0.04 | 2.5 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Siemens | 2.54
Siemens Dimension ExL | 206 | 6.96 | 0.15 | 3.1 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Siemens | 3.25
Siemens Dimension ExL | 206 | 8.53 | -0.12 | 2.9 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Siemens | 2.78
Siemens Dimension ExL | 206 | 9.49 | -0.01 | 2.5 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Siemens | 2.49
Siemens Dimension ExL | 206 | 5.51 | 0.06 | 3.1 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Siemens | 3.16
Siemens Dimension ExL new reagent | 196 | 6.25 | -0.05 | 2.5 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Siemens | 2.45
Siemens Dimension ExL new reagent | 196 | 7.53 | -0.07 | 3.1 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Siemens | 3.03
Siemens Dimension ExL new reagent | 196 | 9.02 | -0.18 | 3.2 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Siemens | 3.02
Siemens Dimension ExL new reagent | 62 | 5.77 | 0.17 | 3.3 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Siemens | 3.47
Siemens Dimension ExL new reagent | 62 | 9.57 | 0.17 | 3.3 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Siemens | 3.47
Siemens Dimension ExL new reagent | 62 | 7.34 | 0.14 | 2.7 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Siemens | 2.84
Siemens Dimension ExL new reagent | 113 | 5.67 | 0.27 | 3.3 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Siemens | 3.57
Siemens Dimension ExL new reagent | 113 | 8.08 | -0.22 | 3.1 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Siemens | 2.88
Siemens Dimension ExL new reagent | 113 | 5.9 | 0.25 | 3.7 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Siemens | 3.95
Siemens Dimension ExL original reagent | 44 | 8.58 | 0.08 | 3.1 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Siemens | 3.18
Siemens Dimension ExL original reagent | 44 | 5.71 | 0.31 | 2.3 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Siemens | 2.61
Siemens Dimension ExL original reagent | 44 | 6.55 | 0.15 | 2.6 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Siemens | 2.75
Siemens Dimension ExL original reagent | 29 | 5.75 | 0.15 | 2.6 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Siemens | 2.75
Siemens Dimension ExL original reagent | 29 | 9.22 | -0.18 | 3 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Siemens | 2.82
Siemens Dimension ExL original reagent | 29 | 7.08 | -0.12 | 4.2 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Siemens | 4.08
Siemens Dimension ExL original reagent | 32 | 6.39 | 0.09 | 3.9 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Siemens | 3.99
Siemens Dimension ExL original reagent | 32 | 7.67 | 0.07 | 2.4 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Siemens | 2.47
Siemens Dimension ExL original reagent | 32 | 9.26 | 0.06 | 3.2 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Siemens | 3.26
Siemens Dimension RxL | 10 | 5.71 | 0.25 | 3 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Siemens | 3.25
Siemens Dimension RxL | 10 | 5.98 | 0.32 | 2.7 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Siemens | 3.02
Siemens Dimension RxL | 10 | 9.06 | -0.25 | 2 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Siemens | 1.75
Siemens Dimension RxL | 10 |  |  |  | GH5-04 | 5.28 | 2019 | CAP | Fresh | Siemens | 0
Siemens Dimension RxL | 10 |  |  |  | GH5-05 | 7.41 | 2019 | CAP | Fresh | Siemens | 0
Siemens Dimension RxL | 346 | 5.4 |  | 3.7 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Siemens | 3.7
Siemens Dimension RxL | 346 | 8.4 |  | 4.1 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Siemens | 4.1
Siemens Dimension RxL | 346 | 6.1 |  | 3.9 | GH2-03 | 6 | 2009 | CAP | Fresh | Siemens | 3.9
Siemens Dimension RxL | 349 | 6.67 | 0.07 | 3.3 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Siemens | 3.37
Siemens Dimension RxL | 349 | 9.41 | -0.09 | 4.4 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Siemens | 4.31
Siemens Dimension RxL | 349 | 7.29 | -0.11 | 3.7 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Siemens | 3.59
Siemens Dimension RxL | 294 | 6.03 | 0.13 | 3.1 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Siemens | 3.23
Siemens Dimension RxL | 294 | 9.53 | -0.27 | 4.2 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Siemens | 3.93
Siemens Dimension RxL | 294 | 7.21 | -0.19 | 3.1 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Siemens | 2.91
Siemens Dimension RxL | 252 | 5.45 | 0.25 | 3.1 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Siemens | 3.35
Siemens Dimension RxL | 252 | 8.43 | -0.27 | 3.4 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Siemens | 3.13
Siemens Dimension RxL | 252 | 6.34 | 0.04 | 3.1 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Siemens | 3.14
Siemens Dimension RxL | 236 | 8.47 | -0.03 | 2.8 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Siemens | 2.77
Siemens Dimension RxL | 236 | 5.69 | 0.29 | 2.7 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Siemens | 2.99
Siemens Dimension RxL | 236 | 6.53 | 0.13 | 2.7 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Siemens | 2.83
Siemens Dimension RxL | 118 | 7.17 | 0.06 | 2.2 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Siemens | 2.26
Siemens Dimension RxL | 118 | 9.38 | 0.06 | 2.4 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Siemens | 2.46
Siemens Dimension RxL | 118 | 6.33 | 0.26 | 2.7 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Siemens | 2.96
Siemens Dimension RxL | 86 | 8.16 | 0.11 | 3 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Siemens | 3.11
Siemens Dimension RxL | 86 | 5.47 | 0.17 | 3.7 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Siemens | 3.87
Siemens Dimension RxL | 86 | 6.4 | 0.26 | 3 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Siemens | 3.26
Siemens Dimension RxL | 67 | 6.77 | 0.19 | 3 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Siemens | 3.19
Siemens Dimension RxL | 67 | 8.34 | -0.05 | 2.5 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Siemens | 2.45
Siemens Dimension RxL | 67 | 5.8 | 0.15 | 3.1 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Siemens | 3.25
Siemens Dimension RxL | 86 | 6.88 | 0.39 | 3.7 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Siemens | 4.09
Siemens Dimension RxL | 86 | 7.27 | 0.3 | 3.1 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Siemens | 3.4
Siemens Dimension RxL | 86 | 9.67 | 0.02 | 2.4 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Siemens | 2.42
Siemens Dimension RxL | 38 | 6.95 | 0.16 | 2.2 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Siemens | 2.36
Siemens Dimension RxL | 38 | 10.32 | 0.04 | 2.5 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Siemens | 2.54
Siemens Dimension RxL | 38 | 6.91 | 0.09 | 2.5 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Siemens | 2.59
Siemens Dimension RxL | 23 | 8.71 | 0.08 | 2.7 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Siemens | 2.78
Siemens Dimension RxL | 23 | 5.43 | 0.11 | 3.7 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Siemens | 3.81
Siemens Dimension RxL | 25 | 5.94 | 0.1 | 3.8 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Siemens | 3.9
Siemens Dimension RxL | 25 | 11.4 | -0.31 | 2.7 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Siemens | 2.39
Siemens Dimension RxL | 25 | 9.35 | -0.18 | 2.1 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Siemens | 1.92
Siemens Dimension RxL | 25 | 5.2 | 0.16 | 3.5 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Siemens | 3.66
Siemens Dimension RxL | 25 | 7.24 | -0.14 | 2.2 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Siemens | 2.06
Siemens Dimension RxL | 24-38 | 11.23 | -0.46 | 3.4 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Siemens | 2.94
Siemens Dimension RxL | 24-38 | 6.04 | 0.11 | 2.9 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Siemens | 3.01
Siemens Dimension RxL | 24-38 | 5.28 | 0.11 | 3.2 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Siemens | 3.31
Siemens Dimension RxL | 24-38 | 7.93 | -0.21 | 2.7 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Siemens | 2.49
Siemens Dimension RxL | 24-38 | 9.04 | -0.26 | 1.6 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Siemens | 1.34
Siemens Dimension RxL | 16-23 | 5.34 | 0.02 | 2.9 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Siemens | 2.92
Siemens Dimension RxL | 16-23 | 9.16 | 0 | 1.9 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Siemens | 1.9
Siemens Dimension RxL | 16-23 | 5.37 | 0.06 | 3.5 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Siemens | 3.56
Siemens Dimension RxL | 16-23 | 11.89 | -0.14 | 2.1 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Siemens | 1.96
Siemens Dimension RxL | 16-23 | 5.96 | 0.02 | 2.6 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Siemens | 2.62
Siemens Dimension RxL | 16 | 5.31 | 0.04 | 2.4 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Siemens | 2.44
Siemens Dimension RxL | 16 | 10.28 | -0.31 | 1.6 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Siemens | 1.29
Siemens Dimension RxL | 16 | 6.19 | -0.01 | 1.4 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Siemens | 1.39
Siemens Dimension RxL | 16 | 11.76 | -0.47 | 2.6 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Siemens | 2.13
Siemens Dimension RxL | 16 | 7.47 | -0.04 | 2.2 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Siemens | 2.16
Siemens Dimension RxL | 14-22 | 9.02 | -0.09 | 1.6 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Siemens | 1.51
Siemens Dimension RxL | 14-22 | 6.05 | 0.04 | 1.7 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Siemens | 1.74
Siemens Dimension RxL | 14-22 | 11.42 | -0.29 | 2.1 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Siemens | 1.81
Siemens Dimension RxL | 14-22 | 5.13 | 0.11 | 2.8 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Siemens | 2.91
Siemens Dimension RxL | 14-22 | 7.58 | 0 | 2.4 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Siemens | 2.4
Siemens Dimension RxL | 13-18 | 6.56 | 0.15 | 3.7 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Siemens | 3.85
Siemens Dimension RxL | 13-18 | 9.47 | -0.06 | 2.5 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Siemens | 2.44
Siemens Dimension RxL | 13-18 | 5.42 | 0.08 | 3.7 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Siemens | 3.78
Siemens Dimension RxL | 13-18 | 8.57 | 0.06 | 1.5 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Siemens | 1.56
Siemens Dimension RxL | 13-18 | 7.26 | 0.01 | 2 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Siemens | 2.01
Siemens Dimension RxL | 14 | 7.22 | 0.07 | 1.7 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Siemens | 1.77
Siemens Dimension RxL | 14 | 5.24 | 0.05 | 2.4 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Siemens | 2.45
Siemens Dimension RxL | 14 | 8.25 | -0.17 | 1.3 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Siemens | 1.13
Siemens Dimension RxL | 14 | 9.39 | -0.4 | 3 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Siemens | 2.6
Siemens Dimension RxL | 14 | 6.00 | -0.12 | 4.2 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Siemens | 4.08
Siemens Dimension RxL | 12 | 6.46 | 0.15 | 3 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Siemens | 3.15
Siemens Dimension RxL | 12 | 8.92 | -0.19 | 2.8 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Siemens | 2.61
Siemens Dimension RxL | 12 | 5.46 | 0.15 | 4.3 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Siemens | 4.45
Siemens Dimension RxL | 12 | 8.31 | -0.04 | 4.1 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Siemens | 4.06
Siemens Dimension RxL | 12 | 7.45 | -0.17 | 3.9 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Siemens | 3.73
Siemens Dimension RxL | 11 | 9.18 | -0.19 | 2.7 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Siemens | 2.51
Siemens Dimension RxL | 11 | 5.13 | 0.21 | 3.5 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Siemens | 3.71
Siemens Dimension RxL | 11 | 6.23 | 0.19 | 2.2 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Siemens | 2.39
Siemens Dimension RxL | 11 | 5.40 | 0.12 | 4.6 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Siemens | 4.72
Siemens Dimension RxL | 11 | 8.19 | 0.11 | 2.7 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Siemens | 2.81
Siemens Dimension RxL | 13 | 7.62 | 0.2 | 3.7 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Siemens | 3.9
Siemens Dimension RxL | 13 | 5.25 | 0.05 | 5.2 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Siemens | 5.25
Siemens Dimension RxL | 13 | 8.31 | 0 | 3.2 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Siemens | 3.2
Siemens Dimension RxL | 13 | 10.47 | 0.14 | 3 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Siemens | 3.14
Siemens Dimension RxL | 13 | 5.98 | 0.11 | 5.4 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Siemens | 5.51
Siemens Dimension RxL | 15 | 7.29 | 0.04 | 4.3 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Siemens | 4.34
Siemens Dimension RxL | 15 | 7.05 | 0.24 | 3 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Siemens | 3.24
Siemens Dimension RxL | 15 | 8.59 | -0.06 | 1.7 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Siemens | 1.64
Siemens Dimension RxL | 15 | 9.44 | -0.06 | 2.5 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Siemens | 2.44
Siemens Dimension RxL | 15 | 5.54 | 0.09 | 4.2 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Siemens | 4.29
Siemens Dimension RxL new reagent | 34 | 6.32 | 0.02 | 4.2 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Siemens | 4.22
Siemens Dimension RxL new reagent | 34 | 7.52 | -0.08 | 3.5 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Siemens | 3.42
Siemens Dimension RxL new reagent | 34 | 9.20 | 0 | 2.8 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Siemens | 2.8
Siemens Dimension RxL new reagent | 114 | 5.67 | 0.07 | 3 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Siemens | 3.07
Siemens Dimension RxL new reagent | 114 | 9.62 | 0.22 | 2.7 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Siemens | 2.92
Siemens Dimension RxL new reagent | 114 | 7.29 | 0.09 | 2.8 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Siemens | 2.89
Siemens Dimension RxL new reagent | 131 | 5.7 | 0.3 | 3.4 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Siemens | 3.7
Siemens Dimension RxL new reagent | 131 | 8.18 | -0.12 | 3.6 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Siemens | 3.48
Siemens Dimension RxL new reagent | 131 | 5.92 | 0.27 | 3 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Siemens | 3.27
Siemens Dimension RxL original reagent | 178 | 8.48 | -0.02 | 2.7 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Siemens | 2.68
Siemens Dimension RxL original reagent | 178 | 5.69 | 0.29 | 2.7 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Siemens | 2.99
Siemens Dimension RxL original reagent | 178 | 6.53 | 0.13 | 2.7 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Siemens | 2.83
Siemens Dimension RxL original reagent | 76 | 5.72 | 0.12 | 3 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Siemens | 3.12
Siemens Dimension RxL original reagent | 76 | 9.22 | -0.18 | 3.1 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Siemens | 2.92
Siemens Dimension RxL original reagent | 76 | 7.07 | -0.13 | 2.5 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Siemens | 2.37
Siemens Dimension RxL original reagent | 11 | 5.55 | 0.15 | 2.8 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Siemens | 2.95
Siemens Dimension RxL original reagent | 11 | 8.09 | -0.21 | 2.6 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Siemens | 2.39
Siemens Dimension RxL original reagent | 11 | 5.82 | 0.17 | 2.5 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Siemens | 2.67
Siemens Dimension RxL original reagent | 76 | 6.42 | 0.12 | 2.5 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Siemens | 2.62
Siemens Dimension RxL original reagent | 76 | 7.66 | 0.06 | 2 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Siemens | 2.06
Siemens Dimension RxL original reagent | 76 | 9.26 | 0.06 | 2.4 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Siemens | 2.46
Siemens Dimension Vista | 290 | 5.50 | 0.04 | 3.5 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Siemens | 3.54
Siemens Dimension Vista | 290 | 5.84 | 0.18 | 3.3 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Siemens | 3.48
Siemens Dimension Vista | 290 | 8.91 | -0.4 | 2.3 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Siemens | 1.9
Siemens Dimension Vista | 290 | 5.30 | 0.02 | 3.4 | GH5-04 | 5.28 | 2019 | CAP | Fresh | Siemens | 3.42
Siemens Dimension Vista | 290 | 7.55 | 0.14 | 2.5 | GH5-05 | 7.41 | 2019 | CAP | Fresh | Siemens | 2.64
Siemens Dimension Vista | 22 | 4.9 |  | 4 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Siemens | 4
Siemens Dimension Vista | 22 | 8.6 |  | 2.3 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Siemens | 2.3
Siemens Dimension Vista | 22 | 6.1 |  | 5.6 | GH2-03 | 6 | 2009 | CAP | Fresh | Siemens | 5.6
Siemens Dimension Vista | 36 | 6.84 | 0.24 | 6.5 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Siemens | 6.74
Siemens Dimension Vista | 36 | 9.21 | -0.29 | 2.4 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Siemens | 2.11
Siemens Dimension Vista | 36 | 7.45 | 0.05 | 4.4 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Siemens | 4.45
Siemens Dimension Vista | 61 | 5.65 | -0.25 | 2.7 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Siemens | 2.45
Siemens Dimension Vista | 61 | 9.23 | -0.57 | 1.7 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Siemens | 1.13
Siemens Dimension Vista | 61 | 7.64 | 0.24 | 3.8 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Siemens | 4.04
Siemens Dimension Vista | 113 | 5.12 | -0.08 | 3.2 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Siemens | 3.12
Siemens Dimension Vista | 113 | 8.75 | 0.05 | 2.1 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Siemens | 2.15
Siemens Dimension Vista | 113 | 6.21 | -0.09 | 3.3 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Siemens | 3.21
Siemens Dimension Vista | 125 | 8.75 | 0.25 | 2.3 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Siemens | 2.55
Siemens Dimension Vista | 125 | 5.39 | -0.01 | 2.8 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Siemens | 2.79
Siemens Dimension Vista | 125 | 6.66 | 0.26 | 4.2 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Siemens | 4.46
Siemens Dimension Vista | 204 | 7.34 | 0.23 | 2.8 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Siemens | 3.03
Siemens Dimension Vista | 204 | 9.49 | 0.17 | 2.9 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Siemens | 3.07
Siemens Dimension Vista | 204 | 6.25 | 0.18 | 3.3 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Siemens | 3.48
Siemens Dimension Vista | 236 | 8.18 | 0.13 | 3.4 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Siemens | 3.53
Siemens Dimension Vista | 236 | 5.25 | -0.05 | 3.6 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Siemens | 3.55
Siemens Dimension Vista | 236 | 6.21 | 0.07 | 3 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Siemens | 3.07
Siemens Dimension Vista | 269 | 6.81 | 0.23 | 3.1 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Siemens | 3.33
Siemens Dimension Vista | 269 | 8.4 | 0.01 | 2.5 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Siemens | 2.51
Siemens Dimension Vista | 269 | 5.73 | 0.08 | 2.9 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Siemens | 2.98
Siemens Dimension Vista | 270 | 6.84 | 0.35 | 3.3 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Siemens | 3.65
Siemens Dimension Vista | 270 | 7.4 | 0.43 | 2.9 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Siemens | 3.33
Siemens Dimension Vista | 270 | 9.83 | 0.18 | 3.6 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Siemens | 3.78
Siemens Dimension Vista | 259 | 6.98 | 0.19 | 2.5 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Siemens | 2.69
Siemens Dimension Vista | 259 | 10.16 | -0.12 | 2.6 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Siemens | 2.48
Siemens Dimension Vista | 259 | 6.97 | 0.15 | 2.6 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Siemens | 2.75
Siemens Dimension Vista | 232 | 8.63 | 0 | 2.8 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Siemens | 2.8
Siemens Dimension Vista | 232 | 5.45 | 0.13 | 2.6 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Siemens | 2.73
Siemens Dimension Vista | 251 | 5.95 | 0.11 | 2.8 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Siemens | 2.91
Siemens Dimension Vista | 251 | 11.84 | 0.13 | 1.9 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Siemens | 2.03
Siemens Dimension Vista | 251 | 9.34 | -0.19 | 2.2 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Siemens | 2.01
Siemens Dimension Vista | 251 | 5.09 | 0.05 | 3 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Siemens | 3.05
Siemens Dimension Vista | 251 | 7.44 | 0.06 | 2.5 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Siemens | 2.56
Siemens Dimension Vista | 248-283 | 11.49 | -0.2 | 3.1 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Siemens | 2.9
Siemens Dimension Vista | 248-283 | 6.04 | 0.11 | 2.7 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Siemens | 2.81
Siemens Dimension Vista | 248-283 | 5.22 | 0.05 | 3.2 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Siemens | 3.25
Siemens Dimension Vista | 248-283 | 8.19 | 0.05 | 2.6 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Siemens | 2.65
Siemens Dimension Vista | 248-283 | 9.12 | -0.18 | 2.4 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Siemens | 2.22
Siemens Dimension Vista | 260-285 | 5.40 | 0.08 | 2.7 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Siemens | 2.78
Siemens Dimension Vista | 260-285 | 9.05 | -0.11 | 2.1 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Siemens | 1.99
Siemens Dimension Vista | 260-285 | 5.40 | 0.09 | 2.7 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Siemens | 2.79
Siemens Dimension Vista | 260-285 | 12.08 | 0.05 | 1.9 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Siemens | 1.95
Siemens Dimension Vista | 260-285 | 6.01 | 0.07 | 2.6 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Siemens | 2.67
Siemens Dimension Vista | 259 | 5.36 | 0.09 | 2.3 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Siemens | 2.39
Siemens Dimension Vista | 259 | 10.37 | -0.22 | 1.8 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Siemens | 1.58
Siemens Dimension Vista | 259 | 6.23 | 0.03 | 2.3 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Siemens | 2.33
Siemens Dimension Vista | 259 | 12.18 | -0.05 | 2 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Siemens | 1.95
Siemens Dimension Vista | 259 | 7.60 | 0.09 | 2 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Siemens | 2.09
Siemens Dimension Vista | 260-291 | 8.86 | -0.25 | 2.5 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Siemens | 2.25
Siemens Dimension Vista | 260-291 | 6.06 | 0.05 | 2.8 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Siemens | 2.85
Siemens Dimension Vista | 260-291 | 11.39 | -0.32 | 2.2 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Siemens | 1.88
Siemens Dimension Vista | 260-291 | 5.01 | -0.01 | 3.1 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Siemens | 3.09
Siemens Dimension Vista | 260-291 | 7.66 | 0.08 | 2.7 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Siemens | 2.78
Siemens Dimension Vista | 274-299 | 6.61 | 0.2 | 2.9 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Siemens | 3.1
Siemens Dimension Vista | 274-299 | 9.35 | -0.18 | 2.7 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Siemens | 2.52
Siemens Dimension Vista | 274-299 | 5.44 | 0.1 | 3.1 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Siemens | 3.2
Siemens Dimension Vista | 274-299 | 8.61 | 0.1 | 2.8 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Siemens | 2.9
Siemens Dimension Vista | 274-299 | 7.47 | 0.22 | 2.5 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Siemens | 2.72
Siemens Dimension Vista | 303 | 7.42 | 0.27 | 2.6 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Siemens | 2.87
Siemens Dimension Vista | 303 | 5.14 | -0.05 | 3.8 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Siemens | 3.75
Siemens Dimension Vista | 303 | 8.48 | 0.06 | 2.4 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Siemens | 2.46
Siemens Dimension Vista | 303 | 9.51 | -0.28 | 2.4 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Siemens | 2.12
Siemens Dimension Vista | 303 | 6.23 | 0.11 | 2.8 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Siemens | 2.91
Siemens Dimension Vista | 280 | 6.49 | 0.18 | 3 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Siemens | 3.18
Siemens Dimension Vista | 280 | 8.86 | -0.25 | 2.6 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Siemens | 2.35
Siemens Dimension Vista | 280 | 5.36 | 0.05 | 2.9 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Siemens | 2.95
Siemens Dimension Vista | 280 | 8.47 | 0.12 | 3 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Siemens | 3.12
Siemens Dimension Vista | 280 | 7.63 | 0.01 | 3 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Siemens | 3.01
Siemens Dimension Vista | 309 | 9.14 | -0.23 | 4.1 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Siemens | 3.87
Siemens Dimension Vista | 309 | 4.98 | 0.06 | 3.6 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Siemens | 3.66
Siemens Dimension Vista | 309 | 6.20 | 0.16 | 3.3 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Siemens | 3.46
Siemens Dimension Vista | 309 | 5.38 | 0.1 | 3.2 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Siemens | 3.3
Siemens Dimension Vista | 309 | 8.23 | 0.15 | 4.2 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Siemens | 4.35
Siemens Dimension Vista | 276 | 6.58 | 0.17 | 2.6 | GH5-06 | 6.41 | 2019 | CAP | Fresh | Siemens | 2.77
Siemens Dimension Vista | 276 | 8.51 | -0.09 | 2.4 | GH5-07 | 8.6 | 2019 | CAP | Fresh | Siemens | 2.31
Siemens Dimension Vista | 276 | 5.49 | 0.07 | 2.9 | GH5-08 | 5.42 | 2019 | CAP | Fresh | Siemens | 2.97
Siemens Dimension Vista | 276 | 7.54 | 0.16 | 2.3 | GH5-09 | 7.38 | 2019 | CAP | Fresh | Siemens | 2.46
Siemens Dimension Vista | 276 | 9.35 | -0.4 | 2 | GH5-10 | 9.75 | 2019 | CAP | Fresh | Siemens | 1.6
Siemens Dimension Vista | 279 | 5.23 | -0.02 | 2.9 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Siemens | 2.88
Siemens Dimension Vista | 279 | 6.49 | 0.08 | 2.7 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Siemens | 2.78
Siemens Dimension Vista | 279 | 8.03 | -0.18 | 2.1 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Siemens | 1.92
Siemens Dimension Vista | 279 | 5.47 | 0 | 2.6 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Siemens | 2.6
Siemens Dimension Vista | 279 | 9.14 | -0.34 | 2.3 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Siemens | 1.96
Siemens Dimension Vista | 243 | 8.03 | 0.06 | 3.4 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Siemens | 3.46
Siemens Dimension Vista | 243 | 5.98 | 0.09 | 3 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Siemens | 0.09
Siemens Dimension Vista | 243 | 5.13 | 0 | 2.9 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Siemens | 2.9
Siemens Dimension Vista | 243 | 7.41 | 0.01 | 3.7 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Siemens | 3.71
Siemens Dimension Vista | 243 | 8.98 | -0.19 | 2.1 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Siemens | 1.91
Siemens Dimension Vista | 245 | 7.68 | -0.13 | 3 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Siemens | -0.13
Siemens Dimension Vista | 245 | 5.94 | 0.03 | 3.3 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Siemens | 3.33
Siemens Dimension Vista | 245 | 8.47 | -0.22 | 2.2 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Siemens | 1.98
Siemens Dimension Vista | 245 | 5.9 | 0.04 | 3.3 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Siemens | 3.34
Siemens Dimension Vista | 245 | 9.3 | -0.26 | 2.1 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Siemens | 1.84
Siemens Dimension Vista | 232 | 5.47 | -0.01 | 3.2 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Siemens | 3.19
Siemens Dimension Vista | 232 | 8.07 | -0.19 | 3.2 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Siemens | 3.01
Siemens Dimension Vista | 232 | 5.15 | 0 | 2.8 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Siemens | 2.8
Siemens Dimension Vista | 232 | 9.86 | -0.19 | 2 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Siemens | -0.19
Siemens Dimension Vista | 232 | 5.95 | -0.04 | 3.6 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Siemens | 3.56
Siemens Dimension Vista | 228 | 8.49 | -0.37 | 2.2 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Siemens | 1.83
Siemens Dimension Vista | 228 | 6.6 | -0.24 | 1.7 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Siemens | 1.46
Siemens Dimension Vista | 228 | 9.12 | -0.27 | 1.8 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Siemens | 1.53
Siemens Dimension Vista | 228 | 5.99 | -0.14 | 1.5 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Siemens | 1.36
Siemens Dimension Vista | 228 | 7.11 | -0.41 | 1.8 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Siemens | 1.39
Siemens Dimension Vista | 216 | 5.69 | -0.11 | 2.1 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Siemens | 1.99
Siemens Dimension Vista | 216 | 8.83 | -0.29 | 1.8 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Siemens | 1.51
Siemens Dimension Vista | 216 | 5.32 | -0.15 | 1.9 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Siemens | 1.75
Siemens Dimension Vista | 216 | 6.89 | -0.36 | 1.6 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Siemens | 1.24
Siemens Dimension Vista | 216 | 8.93 | -0.33 | 2.1 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Siemens | 1.77
Siemens Dimension Vista | 207 | 5.13 | -0.08 | 1.9 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Siemens | 1.82
Siemens Dimension Vista | 207 | 6.5 | -0.18 | 1.6 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Siemens | 1.42
Siemens Dimension Vista | 207 | 7.77 |  | 2 | GH5-13 |  | 2021 | CAP | Fresh | Siemens | 0
Siemens Dimension Vista | 207 | 5.86 | -0.12 | 1.8 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Siemens | 1.68
Siemens Dimension Vista | 207 | 9.99 | -0.13 | 1.7 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Siemens | 1.57
Siemens Dimension Vista | 282 | 7.70 | 0.28 | 2.6 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Siemens | 2.88
Siemens Dimension Vista | 282 | 5.25 | 0.05 | 2.9 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Siemens | 2.95
Siemens Dimension Vista | 282 | 8.25 | -0.06 | 2.4 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Siemens | 2.34
Siemens Dimension Vista | 282 | 9.97 | -0.36 | 2 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Siemens | 1.64
Siemens Dimension Vista | 282 | 6.04 | 0.17 | 2.8 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Siemens | 2.97
Siemens Dimension Vista | 300 | 7.45 | 0.2 | 2.3 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Siemens | 2.5
Siemens Dimension Vista | 300 | 7.01 | 0.2 | 2.5 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Siemens | 2.7
Siemens Dimension Vista | 300 | 8.46 | -0.19 | 2.1 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Siemens | 1.91
Siemens Dimension Vista | 300 | 9.28 | -0.22 | 2.1 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Siemens | 1.88
Siemens Dimension Vista | 300 | 5.54 | 0.09 | 2.8 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Siemens | 2.89
Siemens Dimension Vista new reagent | 102 | 6.33 | 0.03 | 3.2 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Siemens | 3.23
Siemens Dimension Vista new reagent | 121 | 6.59 | 0.29 | 2.8 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Siemens | 3.09
Siemens Dimension Vista new reagent | 102 | 7.48 | -0.12 | 2.8 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Siemens | 2.68
Siemens Dimension Vista new reagent | 121 | 7.80 | 0.2 | 2.9 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Siemens | 3.1
Siemens Dimension Vista new reagent | 102 | 9.09 | -0.11 | 3 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Siemens | 2.89
Siemens Dimension Vista new reagent | 121 | 9.00 | -0.2 | 2 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Siemens | 1.8
Siemens Dimension Vista new reagent | 181 | 5.63 | 0.03 | 4.5 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Siemens | 4.53
Siemens Dimension Vista new reagent | 181 | 9.11 | -0.29 | 4 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Siemens | 3.71
Siemens Dimension Vista new reagent | 181 | 7.33 | 0.13 | 4 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Siemens | 4.13
Siemens Dimension Vista new reagent | 192 | 5.72 | 0.32 | 4 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Siemens | 4.32
Siemens Dimension Vista new reagent | 192 | 8.61 | 0.31 | 2.8 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Siemens | 3.11
Siemens Dimension Vista new reagent | 192 | 5.95 | 0.3 | 3.4 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Siemens | 3.7
Siemens Dimension Vista original reagent | 25 | 6.48 | 0.18 | 3.9 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Siemens | 4.08
Siemens Dimension Vista original reagent | 25 | 8.03 | 0.43 | 2 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Siemens | 2.43
Siemens Dimension Vista original reagent | 25 | 9.15 | -0.05 | 1.3 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Siemens | 1.25
Siemens Dimension Vista original reagent | 105 | 8.75 | 0.25 | 2.2 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Siemens | 2.45
Siemens Dimension Vista original reagent | 105 | 5.39 | -0.01 | 2.8 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Siemens | 2.79
Siemens Dimension Vista original reagent | 105 | 6.64 | 0.24 | 4 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Siemens | 4.24
Siemens Dimension Xpand | 15 | 5.58 | 0.12 | 2.4 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Siemens | 2.52
Siemens Dimension Xpand | 15 | 5.86 | 0.2 | 3.3 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Siemens | 3.5
Siemens Dimension Xpand | 15 | 9.00 | -0.31 | 3.3 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Siemens | 2.99
Siemens Dimension Xpand | 15 |  |  |  | GH5-04 | 5.28 | 2019 | CAP | Fresh | Siemens | 0
Siemens Dimension Xpand | 15 |  |  |  | GH5-05 | 7.41 | 2019 | CAP | Fresh | Siemens | 0
Siemens Dimension Xpand | 139 | 5.3 |  | 3.8 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Siemens | 3.8
Siemens Dimension Xpand | 139 | 8.3 |  | 4.1 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Siemens | 4.1
Siemens Dimension Xpand | 139 | 6 |  | 3.9 | GH2-03 | 6 | 2009 | CAP | Fresh | Siemens | 3.9
Siemens Dimension Xpand | 141 | 6.6 | 0 | 3.3 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Siemens | 3.3
Siemens Dimension Xpand | 141 | 9.37 | -0.13 | 4.9 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Siemens | 4.77
Siemens Dimension Xpand | 141 | 7.23 | -0.17 | 3.7 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Siemens | 3.53
Siemens Dimension Xpand | 126 | 5.95 | 0.05 | 3.3 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Siemens | 3.35
Siemens Dimension Xpand | 126 | 9.44 | -0.36 | 4 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Siemens | 3.64
Siemens Dimension Xpand | 126 | 7.13 | -0.27 | 3.1 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Siemens | 2.83
Siemens Dimension Xpand | 120 | 5.39 | 0.19 | 3.1 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Siemens | 3.29
Siemens Dimension Xpand | 120 | 8.37 | -0.33 | 3.7 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Siemens | 3.37
Siemens Dimension Xpand | 120 | 6.28 | -0.02 | 3.3 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Siemens | 3.28
Siemens Dimension Xpand | 108 | 8.43 | -0.07 | 2.9 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Siemens | 2.83
Siemens Dimension Xpand | 108 | 5.67 | 0.27 | 3.1 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Siemens | 3.37
Siemens Dimension Xpand | 108 | 6.46 | 0.06 | 3.4 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Siemens | 3.46
Siemens Dimension Xpand | 74 | 7.11 | 0 | 2.3 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Siemens | 2.3
Siemens Dimension Xpand | 74 | 9.41 | 0.09 | 2.1 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Siemens | 2.19
Siemens Dimension Xpand | 74 | 6.22 | 0.15 | 3.4 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Siemens | 3.55
Siemens Dimension Xpand | 68 | 8.12 | 0.07 | 3.1 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Siemens | 3.17
Siemens Dimension Xpand | 68 | 5.42 | 0.12 | 3.2 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Siemens | 3.32
Siemens Dimension Xpand | 68 | 6.34 | 0.2 | 3.2 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Siemens | 3.4
Siemens Dimension Xpand | 59 | 6.76 | 0.18 | 3 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Siemens | 3.18
Siemens Dimension Xpand | 59 | 8.32 | -0.07 | 3.1 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Siemens | 3.03
Siemens Dimension Xpand | 59 | 5.81 | 0.16 | 3.4 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Siemens | 3.56
Siemens Dimension Xpand | 65 | 6.82 | 0.33 | 3.1 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Siemens | 3.43
Siemens Dimension Xpand | 65 | 7.23 | 0.26 | 3.1 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Siemens | 3.36
Siemens Dimension Xpand | 65 | 9.69 | 0.04 | 1.8 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Siemens | 1.84
Siemens Dimension Xpand | 47 | 6.94 | 0.15 | 2.8 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Siemens | 2.95
Siemens Dimension Xpand | 47 | 10.28 | 0 | 3.5 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Siemens | 3.5
Siemens Dimension Xpand | 47 | 6.93 | 0.11 | 3 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Siemens | 3.11
Siemens Dimension Xpand | 21 | 8.67 | 0.04 | 2.9 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Siemens | 2.94
Siemens Dimension Xpand | 21 | 5.5 | 0.18 | 3.8 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Siemens | 3.98
Siemens Dimension Xpand | 22 | 5.87 | 0.03 | 2.8 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Siemens | 2.83
Siemens Dimension Xpand | 22 | 11.41 | -0.3 | 2.3 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Siemens | 2
Siemens Dimension Xpand | 22 | 9.4 | -0.13 | 3.7 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Siemens | 3.57
Siemens Dimension Xpand | 22 | 5.08 | 0.04 | 3.3 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Siemens | 3.34
Siemens Dimension Xpand | 22 | 7.26 | -0.12 | 2.8 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Siemens | 2.68
Siemens Dimension Xpand | 21-42 | 11.35 | -0.34 | 2.9 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Siemens | 2.56
Siemens Dimension Xpand | 21-42 | 6.03 | 0.1 | 4.3 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Siemens | 4.4
Siemens Dimension Xpand | 21-42 | 5.28 | 0.11 | 3.8 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Siemens | 3.91
Siemens Dimension Xpand | 21-42 | 8 | -0.14 | 2 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Siemens | 1.86
Siemens Dimension Xpand | 21-42 | 9 | -0.3 | 2.5 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Siemens | 2.2
Siemens Dimension Xpand | 25-47 | 5.37 | 0.05 | 3.5 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Siemens | 3.55
Siemens Dimension Xpand | 25-47 | 9.11 | -0.05 | 2.9 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Siemens | 2.85
Siemens Dimension Xpand | 25-47 | 5.38 | 0.07 | 3.1 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Siemens | 3.17
Siemens Dimension Xpand | 25-47 | 11.94 | -0.09 | 2.5 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Siemens | 2.41
Siemens Dimension Xpand | 25-47 | 5.99 | 0.05 | 2.7 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Siemens | 2.75
Siemens Dimension Xpand | 19 | 5.30 | 0.03 | 2.9 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Siemens | 2.93
Siemens Dimension Xpand | 19 | 10.20 | -0.39 | 3.6 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Siemens | 3.21
Siemens Dimension Xpand | 19 | 6.13 | -0.07 | 2.3 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Siemens | 2.23
Siemens Dimension Xpand | 19 | 11.69 | -0.54 | 3.1 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Siemens | 2.56
Siemens Dimension Xpand | 19 | 7.42 | -0.09 | 3.3 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Siemens | 3.21
Siemens Dimension Xpand | 17-36 | 9.04 | -0.07 | 2.5 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Siemens | 2.43
Siemens Dimension Xpand | 17-36 | 5.98 | -0.03 | 4 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Siemens | 3.97
Siemens Dimension Xpand | 17-36 | 11.43 | -0.28 | 2.5 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Siemens | 2.22
Siemens Dimension Xpand | 17-36 | 5.00 | -0.02 | 3.5 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Siemens | 3.48
Siemens Dimension Xpand | 17-36 | 7.52 | -0.06 | 2.4 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Siemens | 2.34
Siemens Dimension Xpand | 13-29 | 6.59 | 0.18 | 2 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Siemens | 0.18
Siemens Dimension Xpand | 13-29 | 9.6 | 0.07 | 2.4 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Siemens | 2.47
Siemens Dimension Xpand | 13-29 | 5.43 | 0.09 | 2.8 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Siemens | 2.89
Siemens Dimension Xpand | 13-29 | 8.73 | 0.22 | 1.1 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Siemens | 1.32
Siemens Dimension Xpand | 13-29 | 7.35 | 0.1 | 1.3 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Siemens | 1.4
Siemens Dimension Xpand | 22 | 7.28 | 0.13 | 2.6 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Siemens | 2.73
Siemens Dimension Xpand | 22 | 5.30 | 0.11 | 3.7 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Siemens | 3.81
Siemens Dimension Xpand | 22 | 8.21 | -0.21 | 3.5 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Siemens | 3.29
Siemens Dimension Xpand | 22 | 9.53 | -0.26 | 2.7 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Siemens | 2.44
Siemens Dimension Xpand | 22 | 6.20 | 0.08 | 4.4 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Siemens | 4.48
Siemens Dimension Xpand | 20 | 9.18 | -0.19 | 1.6 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Siemens | 1.41
Siemens Dimension Xpand | 20 | 5.07 | 0.15 | 3.3 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Siemens | 3.45
Siemens Dimension Xpand | 20 | 6.20 | 0.16 | 2.8 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Siemens | 2.96
Siemens Dimension Xpand | 20 | 5.40 | 0.12 | 4.1 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Siemens | 4.22
Siemens Dimension Xpand | 20 | 8.05 | -0.03 | 3.3 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Siemens | 3.27
Siemens Dimension Xpand | 14 | 5.21 | -0.04 | 2.7 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Siemens | 2.66
Siemens Dimension Xpand | 14 | 6.57 | 0.16 | 2.1 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Siemens | 2.26
Siemens Dimension Xpand | 14 | 8.06 | -0.15 | 1.9 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Siemens | 1.75
Siemens Dimension Xpand | 14 |  |  |  | GH5-14 | 5.47 | 2019 | CAP | Fresh | Siemens | 0
Siemens Dimension Xpand | 14 |  |  |  | GH5-15 | 9.48 | 2019 | CAP | Fresh | Siemens | 0
Siemens Dimension Xpand | 11 | 7.98 | 0.01 | 2 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Siemens | 0.01
Siemens Dimension Xpand | 11 | 6.02 | 0.13 | 1.5 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Siemens | 1.63
Siemens Dimension Xpand | 11 | 5.26 | 0.13 | 2.9 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Siemens | 3.03
Siemens Dimension Xpand | 11 |  |  |  | GH5-04 | 7.4 | 2020 | CAP | Fresh | Siemens | 0
Siemens Dimension Xpand | 11 |  |  |  | GH5-05 | 9.17 | 2020 | CAP | Fresh | Siemens | 0
Siemens Dimension Xpand | 10 | 5.54 | 0.06 | 1.7 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Siemens | 1.76
Siemens Dimension Xpand | 10 | 8.06 | -0.2 | 2 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Siemens | -0.2
Siemens Dimension Xpand | 10 | 5.17 | 0.02 | 1.7 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Siemens | 1.72
Siemens Dimension Xpand | 10 |  |  |  | GH5-14 | 10.05 | 2020 | CAP | Fresh | Siemens | 0
Siemens Dimension Xpand | 10 |  |  |  | GH5-15 | 5.99 | 2020 | CAP | Fresh | Siemens | 0
Siemens Dimension Xpand | 11 | 7.71 | 0.29 | 2.6 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Siemens | 2.89
Siemens Dimension Xpand | 11 | 5.33 | 0.13 | 2.2 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Siemens | 2.33
Siemens Dimension Xpand | 11 | 8.37 | 0.06 | 2.5 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Siemens | 2.56
Siemens Dimension Xpand | 11 | 10.46 | 0.13 | 2 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Siemens | 2.13
Siemens Dimension Xpand | 11 | 6.05 | 0.18 | 3.4 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Siemens | 3.58
Siemens Dimension Xpand | 26 | 7.27 | 0.02 | 2.8 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Siemens | 2.82
Siemens Dimension Xpand | 26 | 6.90 | 0.09 | 3.1 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Siemens | 3.19
Siemens Dimension Xpand | 26 | 8.49 | -0.16 | 3.1 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Siemens | 2.94
Siemens Dimension Xpand | 26 | 9.43 | -0.07 | 3.4 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Siemens | 3.33
Siemens Dimension Xpand | 26 | 5.56 | 0.11 | 3.2 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Siemens | 3.31
Siemens Dimension Xpand new reagent | 33 | 6.39 | 0.09 | 3.4 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Siemens | 3.49
Siemens Dimension Xpand new reagent | 33 | 7.56 | -0.04 | 3 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Siemens | 2.96
Siemens Dimension Xpand new reagent | 33 | 9.24 | 0.04 | 2.1 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Siemens | 2.14
Siemens Dimension Xpand new reagent | 65 | 5.63 | 0.03 | 3.8 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Siemens | 3.83
Siemens Dimension Xpand new reagent | 65 | 9.58 | 0.18 | 3.5 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Siemens | 3.68
Siemens Dimension Xpand new reagent | 65 | 7.27 | 0.07 | 3.3 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Siemens | 3.37
Siemens Dimension Xpand new reagent | 69 | 5.66 | 0.26 | 3.3 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Siemens | 3.56
Siemens Dimension Xpand new reagent | 69 | 8.1 | -0.2 | 2.8 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Siemens | 2.6
Siemens Dimension Xpand new reagent | 69 | 5.9 | 0.25 | 3.5 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Siemens | 3.75
Siemens Dimension Xpand original reagent | 37 | 5.66 | 0.06 | 2.6 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Siemens | 2.66
Siemens Dimension Xpand original reagent | 37 | 9.19 | -0.21 | 3.4 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Siemens | 3.19
Siemens Dimension Xpand original reagent | 37 | 7.06 | -0.14 | 2.8 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Siemens | 2.66
Siemens Dimension Xpand original reagent | 13 | 5.65 | 0.25 | 2.6 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Siemens | 2.85
Siemens Dimension Xpand original reagent | 13 | 8.13 | -0.17 | 2.1 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Siemens | 1.93
Siemens Dimension Xpand original reagent | 13 | 5.76 | 0.11 | 2.6 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Siemens | 2.71
Siemens Dimension Xpand original reagent | 47 | 6.30 | 0 | 2.7 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Siemens | 2.7
Siemens Dimension Xpand original reagent | 47 | 7.46 | -0.14 | 2.7 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Siemens | 2.56
Siemens Dimension Xpand original reagent | 47 | 9.15 | -0.05 | 2.7 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Siemens | 2.65
Siemens Dimension Xpand original reagent | 69 | 8.43 | -0.07 | 2.9 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Siemens | 2.83
Siemens Dimension Xpand original reagent | 69 | 5.66 | 0.26 | 3 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Siemens | 3.26
Siemens Dimension Xpand original reagent | 69 | 6.42 | 0.02 | 3.6 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Siemens | 3.62
Synchron | 210 | 5 |  | 3.7 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Synchron | 3.7
Synchron | 210 | 8.1 |  | 3.5 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Synchron | 3.5
Synchron | 210 | 5.8 |  | 3.5 | GH2-03 | 6 | 2009 | CAP | Fresh | Synchron | 3.5
Synchron | 269 | 6.43 | -0.17 | 3.6 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Synchron | 3.43
Synchron | 269 | 9.18 | -0.32 | 3.4 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Synchron | 3.08
Synchron | 269 | 7.06 | -0.34 | 3.6 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Synchron | 3.26
Synchron | 241 | 5.82 | -0.08 | 3.1 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Synchron | 3.02
Synchron | 241 | 9.6 | -0.2 | 3.3 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Synchron | 3.1
Synchron | 241 | 7.11 | -0.29 | 3 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Synchron | 2.71
Synchron | 291 | 5.32 | 0.12 | 4.6 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Synchron | 4.72
Synchron | 291 | 8.51 | -0.19 | 4.4 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Synchron | 4.21
Synchron | 291 | 6.23 | -0.07 | 4.1 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Synchron | 4.03
Synchron | 263 | 8.52 | 0.02 | 3.1 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Synchron | 3.12
Synchron | 263 | 5.45 | 0.05 | 3.9 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Synchron | 3.95
Synchron | 263 | 6.38 | -0.02 | 3.4 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Synchron | 3.38
Synchron | 270 | 6.22 | -0.08 | 3.3 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Synchron | 3.22
Synchron | 270 | 7.55 | -0.05 | 3.2 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Synchron | 3.15
Synchron | 270 | 9.37 | 0.17 | 3.4 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Synchron | 3.57
Synchron | 300 | 5.61 | 0.01 | 3.8 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Synchron | 3.81
Synchron | 300 | 9.53 | 0.13 | 3.9 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Synchron | 4.03
Synchron | 300 | 7.11 | -0.09 | 3.5 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Synchron | 3.41
Synchron | 268 | 5.39 | -0.01 | 3.3 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Synchron | 3.29
Synchron | 268 | 8.24 | -0.06 | 3.3 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Synchron | 3.24
Synchron | 268 | 5.48 | -0.17 | 3.1 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Synchron | 2.93
Tosoh A1c 2.2 Plus | 193 | 11.6 |  | 2.9 | GH2-01 | 10.7 | 2006 | CAP | Fresh | Tosoh | 2.9
Tosoh A1c 2.2 Plus | 193 | 5.4 |  | 2.9 | GH2-02 | 5.30 | 2006 | CAP | Fresh | Tosoh | 2.9
Tosoh A1c 2.2 Plus | 193 | 8.8 |  | 2.6 | GH2-03 | 8.40 | 2006 | CAP | Fresh | Tosoh | 2.6
Tosoh A1c 2.2 Plus | 181 | 7.3 |  | 2.9 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Tosoh | 2.9
Tosoh A1c 2.2 Plus | 181 | 7.3 |  | 2.9 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Tosoh | 2.9
Tosoh A1c 2.2 Plus | 181 | 12.6 |  | 3.1 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Tosoh | 3.1
Tosoh A1c 2.2 Plus | 181 | 7.3 |  | 2.9 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Tosoh | 2.9
Tosoh A1c 2.2 Plus | 181 | 7.3 |  | 2.9 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Tosoh | 2.9
Tosoh A1c 2.2 Plus | 181 | 12.6 |  | 3.1 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Tosoh | 3.1
Tosoh A1c 2.2 Plus | 186 | 5.4 |  | 3.5 | GH2-01 | 5.4 | 2007 | CAP | Fresh | Tosoh | 3.5
Tosoh A1c 2.2 Plus | 186 | 12.4 |  | 2.7 | GH2-02 | 11.6 | 2007 | CAP | Fresh | Tosoh | 2.7
Tosoh A1c 2.2 Plus | 186 | 8 |  | 2.9 | GH2-03 | 7.6 | 2007 | CAP | Fresh | Tosoh | 2.9
Tosoh A1c 2.2 Plus | 140 | 9.8 |  | 2.5 | GH2-04 | 9.2 | 2007 | CAP | Fresh | Tosoh | 2.5
Tosoh A1c 2.2 Plus | 140 | 12 |  | 2.4 | GH2-05 | 11.1 | 2007 | CAP | Fresh | Tosoh | 2.4
Tosoh A1c 2.2 Plus | 140 | 6.5 |  | 3.2 | GH2-06 | 6.2 | 2007 | CAP | Fresh | Tosoh | 3.2
Tosoh A1c 2.2 Plus | 130 | 13.9 |  | 3.9 | GH2-01 | 13 | 2008 | CAP | Fresh | Tosoh | 3.9
Tosoh A1c 2.2 Plus | 130 | 5.5 |  | 4.5 | GH2-02 | 5.3 | 2008 | CAP | Fresh | Tosoh | 4.5
Tosoh A1c 2.2 Plus | 130 | 9.4 |  | 3.4 | GH2-03 | 8.8 | 2008 | CAP | Fresh | Tosoh | 3.4
Tosoh A1c 2.2 Plus | 113 | 12.3 |  | 3.3 | GH2-04 | 11.9 | 2008 | CAP | Fresh | Tosoh | 3.3
Tosoh A1c 2.2 Plus | 113 | 7.2 |  | 2.8 | GH2-05 | 6.8 | 2008 | CAP | Fresh | Tosoh | 2.8
Tosoh A1c 2.2 Plus | 113 | 5.3 |  | 3.1 | GH2-06 | 5 | 2008 | CAP | Fresh | Tosoh | 3.1
Tosoh A1c 2.2 Plus | 91 | 5.3 |  | 3.5 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Tosoh | 3.5
Tosoh A1c 2.2 Plus | 91 | 8.7 |  | 2.6 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Tosoh | 2.6
Tosoh A1c 2.2 Plus | 91 | 6.3 |  | 2.8 | GH2-03 | 6 | 2009 | CAP | Fresh | Tosoh | 2.8
Tosoh A1c 2.2 Plus | 71 | 6.87 | 0.27 | 3.1 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Tosoh | 3.37
Tosoh A1c 2.2 Plus | 71 | 9.63 | 0.13 | 2.8 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Tosoh | 2.93
Tosoh A1c 2.2 Plus | 71 | 7.73 | 0.33 | 3 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Tosoh | 3.33
Tosoh A1c 2.2 Plus | 46 | 6.21 | 0.31 | 2.8 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Tosoh | 3.11
Tosoh A1c 2.2 Plus | 46 | 9.96 | 0.16 | 2.9 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Tosoh | 3.06
Tosoh A1c 2.2 Plus | 46 | 7.66 | 0.26 | 2.8 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Tosoh | 3.06
Tosoh A1c 2.2 Plus | 14 | 5.37 | 0.17 | 2 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Tosoh | 2.17
Tosoh A1c 2.2 Plus | 14 | 8.88 | 0.18 | 2.9 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Tosoh | 3.08
Tosoh A1c 2.2 Plus | 14 | 6.55 | 0.25 | 2.8 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Tosoh | 3.05
Tosoh G11 Automated HPLC | 11 | 5.20 | -0.05 | 2.6 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Tosoh | 2.55
Tosoh G11 Automated HPLC | 11 | 6.49 | 0.08 | 2.1 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Tosoh | 2.18
Tosoh G11 Automated HPLC | 11 | 8.27 | 0.06 | 1.9 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Tosoh | 1.96
Tosoh G11 Automated HPLC | 11 |  |  |  | GH5-14 | 5.47 | 2019 | CAP | Fresh | Tosoh | 0
Tosoh G11 Automated HPLC | 11 |  |  |  | GH5-15 | 9.48 | 2019 | CAP | Fresh | Tosoh | 0
Tosoh G11 Automated HPLC | 12 | 5.55 | 0.07 | 1.2 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Tosoh | 1.27
Tosoh G11 Automated HPLC | 12 | 8.32 | 0.06 | 0.9 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Tosoh | 0.96
Tosoh G11 Automated HPLC | 12 | 5.14 | -0.01 | 1.3 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Tosoh | 1.29
Tosoh G11 Automated HPLC | 12 |  |  |  | GH5-14 | 10.05 | 2020 | CAP | Fresh | Tosoh | 0
Tosoh G11 Automated HPLC | 12 |  |  |  | GH5-15 | 5.99 | 2020 | CAP | Fresh | Tosoh | 0
Tosoh G11 Automated HPLC | 11 | 9.07 | 0.21 | 0.9 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Tosoh | 1.11
Tosoh G11 Automated HPLC | 11 | 6.95 | 0.11 | 1.5 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Tosoh | 1.61
Tosoh G11 Automated HPLC | 11 | 9.54 | 0.15 | 1 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Tosoh | 0.15
Tosoh G11 Automated HPLC | 11 |  |  |  | GH5-04 | 6.13 | 2021 | CAP | Fresh | Tosoh | 0
Tosoh G11 Automated HPLC | 11 |  |  |  | GH5-05 | 7.52 | 2021 | CAP | Fresh | Tosoh | 0
Tosoh G11 Automated HPLC | 10 | 5.21 | 0 | 1.1 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Tosoh | 1.1
Tosoh G11 Automated HPLC | 10 | 6.75 | 0.07 | 1.4 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Tosoh | 1.47
Tosoh G11 Automated HPLC | 10 | 8.36 |  | 2.3 | GH5-13 |  | 2021 | CAP | Fresh | Tosoh | 2.3
Tosoh G11 Automated HPLC | 10 |  |  |  | GH5-14 | 5.98 | 2021 | CAP | Fresh | Tosoh | 0
Tosoh G11 Automated HPLC | 10 |  |  |  | GH5-15 | 10.12 | 2021 | CAP | Fresh | Tosoh | 0
Tosoh G7 Auto HPLC | 196 | 11.3 |  | 2.1 | GH2-01 | 10.7 | 2006 | CAP | Fresh | Tosoh | 2.1
Tosoh G7 Auto HPLC | 196 | 5.2 |  | 2.3 | GH2-02 | 5.30 | 2006 | CAP | Fresh | Tosoh | 2.3
Tosoh G7 Auto HPLC | 196 | 8.6 |  | 2.5 | GH2-03 | 8.40 | 2006 | CAP | Fresh | Tosoh | 2.5
Tosoh G7 Auto HPLC | 200 | 7.1 |  | 1.9 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Tosoh | 1.9
Tosoh G7 Auto HPLC | 200 | 7.1 |  | 1.8 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Tosoh | 1.8
Tosoh G7 Auto HPLC | 200 | 12.3 |  | 1.5 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Tosoh | 1.5
Tosoh G7 Auto HPLC | 200 | 7.1 |  | 1.9 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Tosoh | 1.9
Tosoh G7 Auto HPLC | 200 | 7.1 |  | 1.8 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Tosoh | 1.8
Tosoh G7 Auto HPLC | 200 | 12.3 |  | 1.5 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Tosoh | 1.5
Tosoh G7 Auto HPLC | 261 | 5.4 |  | 2.4 | GH2-01 | 5.4 | 2007 | CAP | Fresh | Tosoh | 2.4
Tosoh G7 Auto HPLC | 261 | 12.3 |  | 1.8 | GH2-02 | 11.6 | 2007 | CAP | Fresh | Tosoh | 1.8
Tosoh G7 Auto HPLC | 261 | 7.9 |  | 2 | GH2-03 | 7.6 | 2007 | CAP | Fresh | Tosoh | 2
Tosoh G7 Auto HPLC | 246 | 9.6 |  | 2 | GH2-04 | 9.2 | 2007 | CAP | Fresh | Tosoh | 2
Tosoh G7 Auto HPLC | 246 | 11.8 |  | 1.7 | GH2-05 | 11.1 | 2007 | CAP | Fresh | Tosoh | 1.7
Tosoh G7 Auto HPLC | 246 | 6.4 |  | 2.2 | GH2-06 | 6.2 | 2007 | CAP | Fresh | Tosoh | 2.2
Tosoh G7 Auto HPLC | 263 | 13.7 |  | 3.8 | GH2-01 | 13 | 2008 | CAP | Fresh | Tosoh | 3.8
Tosoh G7 Auto HPLC | 263 | 5.4 |  | 1.9 | GH2-02 | 5.3 | 2008 | CAP | Fresh | Tosoh | 1.9
Tosoh G7 Auto HPLC | 263 | 9.2 |  | 2.6 | GH2-03 | 8.8 | 2008 | CAP | Fresh | Tosoh | 2.6
Tosoh G7 Auto HPLC | 294 | 12 |  | 2.1 | GH2-04 | 11.9 | 2008 | CAP | Fresh | Tosoh | 2.1
Tosoh G7 Auto HPLC | 294 | 7 |  | 1.8 | GH2-05 | 6.8 | 2008 | CAP | Fresh | Tosoh | 1.8
Tosoh G7 Auto HPLC | 294 | 5.2 |  | 1.9 | GH2-06 | 5 | 2008 | CAP | Fresh | Tosoh | 1.9
Tosoh G7 Auto HPLC | 240 | 5.2 |  | 1.7 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Tosoh | 1.7
Tosoh G7 Auto HPLC | 240 | 8.6 |  | 1.4 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Tosoh | 1.4
Tosoh G7 Auto HPLC | 240 | 6.2 |  | 1.5 | GH2-03 | 6 | 2009 | CAP | Fresh | Tosoh | 1.5
Tosoh G7 Auto HPLC | 248 | 6.8 | 0.2 | 1.9 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Tosoh | 2.1
Tosoh G7 Auto HPLC | 248 | 9.53 | 0.03 | 1.8 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Tosoh | 1.83
Tosoh G7 Auto HPLC | 248 | 7.67 | 0.27 | 2.3 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Tosoh | 2.57
Tosoh G7 Auto HPLC | 202 | 6.12 | 0.22 | 1.8 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Tosoh | 2.02
Tosoh G7 Auto HPLC | 202 | 9.8 | 0 | 1.4 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Tosoh | 1.4
Tosoh G7 Auto HPLC | 202 | 7.56 | 0.16 | 1.5 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Tosoh | 1.66
Tosoh G7 Auto HPLC | 216 | 5.38 | 0.18 | 2 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Tosoh | 2.18
Tosoh G7 Auto HPLC | 216 | 8.81 | 0.11 | 1.7 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Tosoh | 1.81
Tosoh G7 Auto HPLC | 216 | 6.48 | 0.18 | 1.6 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Tosoh | 1.78
Tosoh G7 Auto HPLC | 170 | 8.73 | 0.23 | 1.7 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Tosoh | 1.93
Tosoh G7 Auto HPLC | 170 | 5.56 | 0.16 | 1.6 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Tosoh | 1.76
Tosoh G7 Auto HPLC | 170 | 6.57 | 0.17 | 2.1 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Tosoh | 2.27
Tosoh G7 Auto HPLC | 181 | 6.62 | 0.32 | 2.2 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Tosoh | 2.52
Tosoh G7 Auto HPLC | 181 | 8.01 | 0.41 | 2 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Tosoh | 2.41
Tosoh G7 Auto HPLC | 181 | 9.56 | 0.36 | 2 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Tosoh | 2.36
Tosoh G7 Auto HPLC | 184 | 5.73 | 0.13 | 2.1 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Tosoh | 2.23
Tosoh G7 Auto HPLC | 184 | 9.76 | 0.36 | 1.7 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Tosoh | 2.06
Tosoh G7 Auto HPLC | 184 | 7.58 | 0.38 | 1.6 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Tosoh | 1.98
Tosoh G7 Auto HPLC | 145 | 5.65 | 0.25 | 2.1 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Tosoh | 2.35
Tosoh G7 Auto HPLC | 145 | 8.61 | 0.31 | 1.8 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Tosoh | 2.11
Tosoh G7 Auto HPLC | 145 | 5.65 | 0 | 2.1 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Tosoh | 2.1
Tosoh G7 Auto HPLC | 116 | 7.38 | 0.27 | 1.7 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Tosoh | 1.97
Tosoh G7 Auto HPLC | 116 | 9.7 | 0.38 | 1.4 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Tosoh | 1.78
Tosoh G7 Auto HPLC | 116 | 6.34 | 0.27 | 1.9 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Tosoh | 2.17
Tosoh G7 Auto HPLC | 112 | 8.41 | 0.36 | 2.1 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Tosoh | 2.46
Tosoh G7 Auto HPLC | 112 | 5.53 | 0.23 | 2.4 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Tosoh | 2.63
Tosoh G7 Auto HPLC | 112 | 6.41 | 0.27 | 2.2 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Tosoh | 2.47
Tosoh G7 Auto HPLC | 88 | 6.77 | 0.19 | 1.9 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Tosoh | 2.09
Tosoh G7 Auto HPLC | 88 | 8.7 | 0.31 | 1.8 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Tosoh | 2.11
Tosoh G7 Auto HPLC | 88 | 5.81 | 0.16 | 2 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Tosoh | 2.16
Tosoh G7 Auto HPLC | 102 | 6.81 | 0.32 | 1.9 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Tosoh | 2.22
Tosoh G7 Auto HPLC | 102 | 7.33 | 0.36 | 1.9 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Tosoh | 2.26
Tosoh G7 Auto HPLC | 102 | 10.06 | 0.41 | 1.9 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Tosoh | 2.31
Tosoh G7 Auto HPLC | 73 | 7.09 | 0.3 | 2.4 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Tosoh | 2.7
Tosoh G7 Auto HPLC | 73 | 10.51 | 0.23 | 2.3 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Tosoh | 2.53
Tosoh G7 Auto HPLC | 73 | 7.09 | 0.27 | 2.6 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Tosoh | 2.87
Tosoh G7 Auto HPLC | 52 | 8.95 | 0.32 | 2.6 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Tosoh | 2.92
Tosoh G7 Auto HPLC | 52 | 5.48 | 0.16 | 2.7 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Tosoh | 2.86
Tosoh G7 Auto HPLC | 50 | 6.06 | 0.22 | 1.6 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Tosoh | 1.82
Tosoh G7 Auto HPLC | 50 | 11.95 | 0.24 | 1.7 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Tosoh | 1.94
Tosoh G7 Auto HPLC | 50 | 9.81 | 0.28 | 1.8 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Tosoh | 2.08
Tosoh G7 Auto HPLC | 50 | 5.2 | 0.16 | 1.6 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Tosoh | 1.76
Tosoh G7 Auto HPLC | 50 | 7.61 | 0.23 | 1.6 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Tosoh | 1.83
Tosoh G7 Auto HPLC | 45-60 | 11.98 | 0.29 | 2 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Tosoh | 2.29
Tosoh G7 Auto HPLC | 45-60 | 6.15 | 0.22 | 2.1 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Tosoh | 2.32
Tosoh G7 Auto HPLC | 45-60 | 5.32 | 0.15 | 2.4 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Tosoh | 2.55
Tosoh G7 Auto HPLC | 45-60 | 8.41 | 0.27 | 1.9 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Tosoh | 2.17
Tosoh G7 Auto HPLC | 45-60 | 9.5 | 0.2 | 2.1 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Tosoh | 2.3
Tosoh G7 Auto HPLC | 36-46 | 5.43 | 0.11 | 1.9 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Tosoh | 2.01
Tosoh G7 Auto HPLC | 36-46 | 9.40 | 0.24 | 2 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Tosoh | 2.24
Tosoh G7 Auto HPLC | 36-46 | 5.45 | 0.14 | 2 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Tosoh | 2.14
Tosoh G7 Auto HPLC | 36-46 | 12.28 | 0.25 | 1.8 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Tosoh | 2.05
Tosoh G7 Auto HPLC | 36-46 | 6.06 | 0.12 | 2 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Tosoh | 2.12
Tosoh G7 Auto HPLC | 31 | 5.40 | 0.13 | 2 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Tosoh | 0.13
Tosoh G7 Auto HPLC | 31 | 10.68 | 0.09 | 2.2 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Tosoh | 2.29
Tosoh G7 Auto HPLC | 31 | 6.34 | 0.14 | 1.7 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Tosoh | 1.84
Tosoh G7 Auto HPLC | 31 | 12.32 | 0.09 | 2 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Tosoh | 2.09
Tosoh G7 Auto HPLC | 31 | 7.67 | 0.16 | 2.1 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Tosoh | 2.26
Tosoh G7 Auto HPLC | 13-21 | 9.24 | 0.13 | 1.8 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Tosoh | 1.93
Tosoh G7 Auto HPLC | 13-21 | 6.16 | 0.15 | 2.6 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Tosoh | 2.75
Tosoh G7 Auto HPLC | 13-21 | 11.82 | 0.11 | 1.8 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Tosoh | 1.91
Tosoh G7 Auto HPLC | 13-21 | 5.16 | 0.14 | 2.8 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Tosoh | 2.94
Tosoh G7 Auto HPLC | 13-21 | 7.71 | 0.13 | 2.1 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Tosoh | 2.23
Tosoh G7 Auto HPLC | 0-10 | 6.57 | 0.16 | 1 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Tosoh | 0.16
Tosoh G7 Auto HPLC | 0-10 | 9.73 | 0.2 | 1.1 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Tosoh | 1.3
Tosoh G7 Auto HPLC | 0-10 | 5.42 | 0.08 | 1.5 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Tosoh | 1.58
Tosoh G7 Auto HPLC | 0-10 |  |  |  | GH5-04 | 8.51 | 2017 | CAP | Fresh | Tosoh | 0
Tosoh G7 Auto HPLC | 0-10 |  |  |  | GH5-05 | 7.25 | 2017 | CAP | Fresh | Tosoh | 0
Tosoh G7 Auto HPLC | 10 | 7.28 | 0.13 | 2.8 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Tosoh | 2.93
Tosoh G7 Auto HPLC | 10 | 5.28 | 0.09 | 1.7 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Tosoh | 1.79
Tosoh G7 Auto HPLC | 10 | 8.52 | 0.1 | 1.1 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Tosoh | 1.2
Tosoh G7 Auto HPLC | 10 | - | - |  | GH5-04 | 9.79 | 2018 | CAP | Fresh | Tosoh | 0
Tosoh G7 Auto HPLC | 10 | - | - |  | GH5-05 | 6.12 | 2018 | CAP | Fresh | Tosoh | 0
Tosoh G8 Auto HPLC | 41 | 5.2 |  | 1.8 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Tosoh | 1.8
Tosoh G8 Auto HPLC | 41 | 8.5 |  | 1.2 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Tosoh | 1.2
Tosoh G8 Auto HPLC | 41 | 6.2 |  | 1.3 | GH2-03 | 6 | 2009 | CAP | Fresh | Tosoh | 1.3
Tosoh G8 Auto HPLC | 84 | 6.8 | 0.2 | 1.7 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Tosoh | 1.9
Tosoh G8 Auto HPLC | 84 | 9.52 | 0.02 | 1.5 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Tosoh | 1.52
Tosoh G8 Auto HPLC | 84 | 7.63 | 0.23 | 1.9 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Tosoh | 2.13
Tosoh G8 Auto HPLC | 116 | 6.1 | 0.2 | 1.9 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Tosoh | 2.1
Tosoh G8 Auto HPLC | 116 | 9.8 | 0 | 1.5 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Tosoh | 1.5
Tosoh G8 Auto HPLC | 116 | 7.56 | 0.16 | 1.5 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Tosoh | 1.66
Tosoh G8 Auto HPLC | 148 | 5.34 | 0.14 | 1.5 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Tosoh | 1.64
Tosoh G8 Auto HPLC | 148 | 8.77 | 0.07 | 1.6 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Tosoh | 1.67
Tosoh G8 Auto HPLC | 148 | 6.45 | 0.15 | 1.4 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Tosoh | 1.55
Tosoh G8 Auto HPLC | 186 | 8.69 | 0.19 | 1.7 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Tosoh | 1.89
Tosoh G8 Auto HPLC | 186 | 5.53 | 0.13 | 1.5 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Tosoh | 1.63
Tosoh G8 Auto HPLC | 186 | 6.56 | 0.16 | 1.4 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Tosoh | 1.56
Tosoh G8 Auto HPLC | 212 | 6.58 | 0.28 | 1.5 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Tosoh | 1.78
Tosoh G8 Auto HPLC | 212 | 7.99 | 0.39 | 1.4 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Tosoh | 1.79
Tosoh G8 Auto HPLC | 212 | 9.54 | 0.34 | 1.4 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Tosoh | 1.74
Tosoh G8 Auto HPLC | 260 | 5.71 | 0.11 | 1.5 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Tosoh | 1.61
Tosoh G8 Auto HPLC | 260 | 9.75 | 0.35 | 1.2 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Tosoh | 1.55
Tosoh G8 Auto HPLC | 260 | 7.58 | 0.38 | 1.3 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Tosoh | 1.68
Tosoh G8 Auto HPLC | 284 | 5.62 | 0.22 | 1.5 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Tosoh | 1.72
Tosoh G8 Auto HPLC | 284 | 8.58 | 0.28 | 1.3 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Tosoh | 1.58
Tosoh G8 Auto HPLC | 284 | 5.64 | -0.01 | 1.8 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Tosoh | 1.79
Tosoh G8 Auto HPLC | 259 | 7.34 | 0.23 | 1.3 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Tosoh | 1.53
Tosoh G8 Auto HPLC | 259 | 9.67 | 0.35 | 1.2 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Tosoh | 1.55
Tosoh G8 Auto HPLC | 259 | 6.31 | 0.24 | 1.4 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Tosoh | 1.64
Tosoh G8 Auto HPLC | 280 | 8.41 | 0.36 | 1.5 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Tosoh | 1.86
Tosoh G8 Auto HPLC | 280 | 5.51 | 0.21 | 1.7 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Tosoh | 1.91
Tosoh G8 Auto HPLC | 280 | 6.4 | 0.26 | 1.7 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Tosoh | 1.96
Tosoh G8 Auto HPLC | 330 | 6.73 | 0.15 | 1.8 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Tosoh | 1.95
Tosoh G8 Auto HPLC | 330 | 8.68 | 0.29 | 1.6 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Tosoh | 1.89
Tosoh G8 Auto HPLC | 330 | 5.78 | 0.13 | 1.8 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Tosoh | 1.93
Tosoh G8 Auto HPLC | 327 | 6.77 | 0.28 | 1.8 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Tosoh | 2.08
Tosoh G8 Auto HPLC | 327 | 7.32 | 0.35 | 1.9 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Tosoh | 2.25
Tosoh G8 Auto HPLC | 327 | 10.05 | 0.4 | 1.8 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Tosoh | 2.2
Tosoh G8 Auto HPLC | 330 | 7.05 | 0.26 | 2.1 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Tosoh | 2.36
Tosoh G8 Auto HPLC | 330 | 10.5 | 0.22 | 1.9 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Tosoh | 2.12
Tosoh G8 Auto HPLC | 330 | 7.06 | 0.24 | 1.9 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Tosoh | 2.14
Tosoh G8 Auto HPLC | 287 | 8.93 | 0.3 | 1.9 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Tosoh | 2.2
Tosoh G8 Auto HPLC | 287 | 5.45 | 0.13 | 1.9 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Tosoh | 2.03
Tosoh G8 Auto HPLC | 285 | 6.03 | 0.19 | 1.7 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Tosoh | 1.89
Tosoh G8 Auto HPLC | 285 | 11.98 | 0.27 | 1.7 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Tosoh | 1.97
Tosoh G8 Auto HPLC | 285 | 9.81 | 0.28 | 1.7 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Tosoh | 1.98
Tosoh G8 Auto HPLC | 285 | 5.17 | 0.13 | 2 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Tosoh | 2.13
Tosoh G8 Auto HPLC | 285 | 7.6 | 0.22 | 1.5 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Tosoh | 1.72
Tosoh G8 Auto HPLC | 298-344 | 11.88 | 0.19 | 1.6 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Tosoh | 1.79
Tosoh G8 Auto HPLC | 298-344 | 6.09 | 0.16 | 1.7 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Tosoh | 1.86
Tosoh G8 Auto HPLC | 298-344 | 5.26 | 0.09 | 1.7 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Tosoh | 1.79
Tosoh G8 Auto HPLC | 298-344 | 8.33 | 0.19 | 1.6 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Tosoh | 1.79
Tosoh G8 Auto HPLC | 298-344 | 9.41 | 0.11 | 1.5 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Tosoh | 1.61
Tosoh G8 Auto HPLC | 308-362 | 5.39 | 0.07 | 1.9 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Tosoh | 1.97
Tosoh G8 Auto HPLC | 308-362 | 9.38 | 0.22 | 1.5 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Tosoh | 1.72
Tosoh G8 Auto HPLC | 308-362 | 5.41 | 0.1 | 1.8 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Tosoh | 1.9
Tosoh G8 Auto HPLC | 308-362 | 12.17 | 0.14 | 1.5 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Tosoh | 1.64
Tosoh G8 Auto HPLC | 308-362 | 6.02 | 0.08 | 1.6 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Tosoh | 1.68
Tosoh G8 Auto HPLC | 313 | 5.34 | 0.07 | 1.6 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Tosoh | 1.67
Tosoh G8 Auto HPLC | 313 | 10.70 | 0.11 | 1.4 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Tosoh | 1.51
Tosoh G8 Auto HPLC | 313 | 6.29 | 0.09 | 1.4 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Tosoh | 1.49
Tosoh G8 Auto HPLC | 313 | 12.28 | 0.05 | 1.4 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Tosoh | 1.45
Tosoh G8 Auto HPLC | 313 | 7.65 | 0.14 | 1.5 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Tosoh | 1.64
Tosoh G8 Auto HPLC | 321-373 | 9.17 | 0.06 | 1.6 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Tosoh | 1.66
Tosoh G8 Auto HPLC | 321-373 | 6.07 | 0.06 | 1.6 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Tosoh | 1.66
Tosoh G8 Auto HPLC | 321-373 | 11.73 | 0.02 | 1.5 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Tosoh | 1.52
Tosoh G8 Auto HPLC | 321-373 | 5.10 | 0.08 | 1.7 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Tosoh | 1.78
Tosoh G8 Auto HPLC | 321-373 | 7.66 | 0.08 | 1.4 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Tosoh | 1.48
Tosoh G8 Auto HPLC | 335-389 | 6.53 | 0.12 | 1.6 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Tosoh | 1.72
Tosoh G8 Auto HPLC | 335-389 | 9.73 | 0.2 | 1.4 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Tosoh | 1.6
Tosoh G8 Auto HPLC | 335-389 | 5.4 | 0.06 | 1.7 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Tosoh | 1.76
Tosoh G8 Auto HPLC | 335-389 | 8.70 | 0.19 | 1.4 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Tosoh | 1.59
Tosoh G8 Auto HPLC | 335-389 | 7.38 | 0.13 | 1.4 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Tosoh | 1.53
Tosoh G8 Auto HPLC | 319 | 7.57 | 0.15 | 1.4 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Tosoh | 1.55
Tosoh G8 Auto HPLC | 319 | 5.38 | 0.18 | 1.7 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Tosoh | 1.88
Tosoh G8 Auto HPLC | 319 | 8.42 | 0.11 | 1.4 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Tosoh | 1.51
Tosoh G8 Auto HPLC | 319 | 10.51 | 0.18 | 1.4 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Tosoh | 1.58
Tosoh G8 Auto HPLC | 319 | 6.02 | 0.15 | 1.5 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Tosoh | 1.65
Tosoh G8 Auto HPLC | 378 | 7.35 | 0.1 | 1.4 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Tosoh | 1.5
Tosoh G8 Auto HPLC | 378 | 6.99 | 0.18 | 1.4 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Tosoh | 1.58
Tosoh G8 Auto HPLC | 378 | 8.72 | 0.07 | 1.3 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Tosoh | 1.37
Tosoh G8 Auto HPLC | 378 | 9.68 | 0.18 | 1.2 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Tosoh | 1.38
Tosoh G8 Auto HPLC | 378 | 5.54 | 0.09 | 1.4 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Tosoh | 1.49
Tosoh G8 Auto HPLC | 373 | 7.21 | 0.06 | 1.4 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Tosoh | 1.46
Tosoh G8 Auto HPLC | 373 | 5.29 | 0.1 | 1.4 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Tosoh | 1.5
Tosoh G8 Auto HPLC | 373 | 8.52 | 0.1 | 1.3 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Tosoh | 1.4
Tosoh G8 Auto HPLC | 373 | 9.92 | 0.13 | 1.1 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Tosoh | 1.23
Tosoh G8 Auto HPLC | 373 | 6.23 | 0.11 | 1.3 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Tosoh | 1.41
Tosoh G8 Auto HPLC | 310 | 6.48 | 0.17 | 1.4 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Tosoh | 1.57
Tosoh G8 Auto HPLC | 310 | 9.27 | 0.16 | 1.2 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Tosoh | 1.36
Tosoh G8 Auto HPLC | 310 | 5.37 | 0.06 | 1.3 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Tosoh | 1.36
Tosoh G8 Auto HPLC | 310 | 8.50 | 0.15 | 1.1 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Tosoh | 1.25
Tosoh G8 Auto HPLC | 310 | 7.76 | 0.14 | 1.2 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Tosoh | 1.34
Tosoh G8 Automated HPLC | 354 | 9.53 | 0.16 | 1.4 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Tosoh | 1.56
Tosoh G8 Automated HPLC | 354 | 5.02 | 0.1 | 1.6 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Tosoh | 1.7
Tosoh G8 Automated HPLC | 354 | 6.25 | 0.21 | 1.4 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Tosoh | 1.61
Tosoh G8 Automated HPLC | 354 | 5.47 | 0.19 | 1.4 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Tosoh | 1.59
Tosoh G8 Automated HPLC | 354 | 8.28 | 0.2 | 1.2 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Tosoh | 1.4
Tosoh G8 Automated HPLC | 347 | 5.60 | 0.14 | 1.4 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Tosoh | 1.54
Tosoh G8 Automated HPLC | 347 | 5.54 | -0.12 | 1.7 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Tosoh | 1.58
Tosoh G8 Automated HPLC | 347 | 9.55 | 0.24 | 1.3 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Tosoh | 1.54
Tosoh G8 Automated HPLC | 347 | 5.36 | 0.08 | 1.4 | GH5-04 | 5.28 | 2019 | CAP | Fresh | Tosoh | 1.48
Tosoh G8 Automated HPLC | 347 | 7.60 | 0.19 | 1.3 | GH5-05 | 7.41 | 2019 | CAP | Fresh | Tosoh | 1.49
Tosoh G8 Automated HPLC | 294 | 6.61 | 0.2 | 1.4 | GH5-06 | 6.41 | 2019 | CAP | Fresh | Tosoh | 1.6
Tosoh G8 Automated HPLC | 294 | 8.93 | 0.33 | 1.3 | GH5-07 | 8.6 | 2019 | CAP | Fresh | Tosoh | 1.63
Tosoh G8 Automated HPLC | 294 | 5.55 | 0.13 | 1.6 | GH5-08 | 5.42 | 2019 | CAP | Fresh | Tosoh | 1.73
Tosoh G8 Automated HPLC | 294 | 7.62 | 0.24 | 1.4 | GH5-09 | 7.38 | 2019 | CAP | Fresh | Tosoh | 1.64
Tosoh G8 Automated HPLC | 294 | 10.07 | 0.32 | 1.4 | GH5-10 | 9.75 | 2019 | CAP | Fresh | Tosoh | 1.72
Tosoh G8 Automated HPLC | 337 | 5.30 | 0.05 | 1.9 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Tosoh | 1.95
Tosoh G8 Automated HPLC | 337 | 6.65 | 0.24 | 1.6 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Tosoh | 1.84
Tosoh G8 Automated HPLC | 337 | 8.46 | 0.25 | 1.4 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Tosoh | 1.65
Tosoh G8 Automated HPLC | 337 | 5.56 | 0.09 | 1.7 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Tosoh | 1.79
Tosoh G8 Automated HPLC | 337 | 9.79 | 0.31 | 1.4 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Tosoh | 1.71
Tosoh G8 Automated HPLC | 308 | 8.2 | 0.23 | 1.5 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Tosoh | 1.73
Tosoh G8 Automated HPLC | 308 | 5.99 | 0.1 | 1.7 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Tosoh | 1.8
Tosoh G8 Automated HPLC | 308 | 5.18 | 0.05 | 2 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Tosoh | 0.05
Tosoh G8 Automated HPLC | 308 | 7.58 | 0.18 | 1.9 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Tosoh | 2.08
Tosoh G8 Automated HPLC | 308 | 9.43 | 0.26 | 1.4 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Tosoh | 1.66
Tosoh G8 Automated HPLC | 271 | 7.99 | 0.18 | 1.6 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Tosoh | 1.78
Tosoh G8 Automated HPLC | 271 | 6.03 | 0.12 | 1.9 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Tosoh | 2.02
Tosoh G8 Automated HPLC | 271 | 8.96 | 0.27 | 1.7 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Tosoh | 1.97
Tosoh G8 Automated HPLC | 271 | 6.04 | 0.18 | 1.8 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Tosoh | 1.98
Tosoh G8 Automated HPLC | 271 | 9.84 | 0.28 | 1.5 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Tosoh | 1.78
Tosoh G8 Automated HPLC | 316 | 5.59 | 0.11 | 1.7 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Tosoh | 1.81
Tosoh G8 Automated HPLC | 316 | 8.46 | 0.2 | 1.4 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Tosoh | 1.6
Tosoh G8 Automated HPLC | 316 | 5.18 | 0.03 | 1.9 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Tosoh | 1.93
Tosoh G8 Automated HPLC | 316 | 10.39 | 0.34 | 1.3 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Tosoh | 1.64
Tosoh G8 Automated HPLC | 316 | 6.11 | 0.12 | 1.6 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Tosoh | 1.72
Tosoh G8 Automated HPLC | 311 | 9.18 | 0.32 | 1.6 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Tosoh | 1.92
Tosoh G8 Automated HPLC | 311 | 7.03 | 0.19 | 1.9 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Tosoh | 2.09
Tosoh G8 Automated HPLC | 311 | 9.65 | 0.26 | 1.5 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Tosoh | 1.76
Tosoh G8 Automated HPLC | 311 | 6.19 | 0.06 | 1.8 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Tosoh | 1.86
Tosoh G8 Automated HPLC | 311 | 7.69 | 0.17 | 1.6 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Tosoh | 1.77
Tosoh G8 Automated HPLC | 266 | 5.95 | 0.15 | 1.6 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Tosoh | 1.75
Tosoh G8 Automated HPLC | 266 | 9.38 | 0.26 | 1.6 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Tosoh | 1.86
Tosoh G8 Automated HPLC | 266 | 5.58 | 0.11 | 1.6 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Tosoh | 1.71
Tosoh G8 Automated HPLC | 266 | 7.37 | 0.12 | 1.5 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Tosoh | 1.62
Tosoh G8 Automated HPLC | 266 | 9.58 | 0.32 | 1.5 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Tosoh | 1.82
Tosoh G8 Automated HPLC | 304 | 5.25 | 0.04 | 1.6 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Tosoh | 1.64
Tosoh G8 Automated HPLC | 304 | 6.84 | 0.16 | 1.4 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Tosoh | 1.56
Tosoh G8 Automated HPLC | 304 | 8.52 |  | 2.3 | GH5-13 |  | 2021 | CAP | Fresh | Tosoh | 2.3
Tosoh G8 Automated HPLC | 304 | 6.09 | 0.11 | 1.4 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Tosoh | 1.51
Tosoh G8 Automated HPLC | 304 | 10.26 | 0.14 | 1.6 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Tosoh | 1.74
Trinity Biotech HPLC | 20 | 7.15 | 0.04 | 2.5 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Trinity Biotech | 2.54
Trinity Biotech HPLC | 20 | 9.23 | -0.09 | 2.8 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Trinity Biotech | 2.71
Trinity Biotech HPLC | 20 | 6.11 | 0.04 | 2.2 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Trinity Biotech | 2.24
Trinity Biotech HPLC | 18 | 7.96 | -0.09 | 1.6 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Trinity Biotech | 1.51
Trinity Biotech HPLC | 18 | 5.32 | 0.02 | 1.5 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Trinity Biotech | 1.52
Trinity Biotech HPLC | 18 | 6.14 | 0 | 1.7 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Trinity Biotech | 1.7
Trinity Biotech HPLC | 16 | 6.49 | -0.09 | 1.8 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Trinity Biotech | 1.71
Trinity Biotech HPLC | 16 | 8.38 | -0.01 | 3.7 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Trinity Biotech | 3.69
Trinity Biotech HPLC | 16 | 5.58 | -0.07 | 2.6 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Trinity Biotech | 2.53
Trinity Biotech HPLC | 20 | 6.39 | -0.1 | 1.9 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Trinity Biotech | 1.8
Trinity Biotech HPLC | 20 | 6.77 | -0.2 | 2.2 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Trinity Biotech | 2
Trinity Biotech HPLC | 20 | 9.51 | -0.14 | 3.1 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Trinity Biotech | 2.96
Trinity Biotech HPLC | 14 | 6.66 | -0.13 | 1.5 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Trinity Biotech | 1.37
Trinity Biotech HPLC | 14 | 10.26 | -0.02 | 1.9 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Trinity Biotech | 1.88
Trinity Biotech HPLC | 14 | 6.68 | -0.14 | 1.9 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Trinity Biotech | 1.76
Trinity Biotech HPLC | 13 | 8.5 | -0.13 | 1.9 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Trinity Biotech | 1.77
Trinity Biotech HPLC | 13 | 5.31 | -0.01 | 2 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Trinity Biotech | 1.99
Trinity Biotech HPLC | 14 | 5.73 | -0.11 | 1.9 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Trinity Biotech | 1.79
Trinity Biotech HPLC | 14 | 11.39 | -0.32 | 1.9 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Trinity Biotech | 1.58
Trinity Biotech HPLC | 14 | 9.39 | -0.14 | 2.7 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Trinity Biotech | 2.56
Trinity Biotech HPLC | 14 | 5.05 | 0.01 | 1.9 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Trinity Biotech | 1.91
Trinity Biotech HPLC | 14 | 7.24 | -0.14 | 2.1 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Trinity Biotech | 1.96
Trinity Biotech HPLC | 12-13 | 11.72 | 0.03 | 2.6 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Trinity Biotech | 2.63
Trinity Biotech HPLC | 12-13 | 5.94 | 0.01 | 2.7 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Trinity Biotech | 2.71
Trinity Biotech HPLC | 12-13 | 5.22 | 0.05 | 2.8 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Trinity Biotech | 2.85
Trinity Biotech HPLC | 12-13 | 8.22 | 0.08 | 2.3 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Trinity Biotech | 2.38
Trinity Biotech HPLC | 12-13 | 9.44 | 0.14 | 2.8 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Trinity Biotech | 2.94
Trinity Biotech HPLC | 9-10 | 5.42 | 0.1 | 2.7 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Trinity Biotech | 2.8
Trinity Biotech HPLC | 9-10 | 9.26 | 0.1 | 2.8 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Trinity Biotech | 2.9
Trinity Biotech HPLC | 9-10 | 5.37 | 0.06 | 2.2 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Trinity Biotech | 2.26
Trinity Biotech HPLC | 9-10 |  |  |  | GH5-04 | 12.03 | 2016 | CAP | Fresh | Trinity Biotech | 0
Trinity Biotech HPLC | 9-10 |  |  |  | GH5-05 | 5.94 | 2016 | CAP | Fresh | Trinity Biotech | 0
Trinity Biotech HPLC (Affinity) | 27 | 6.66 | 0.06 | 3.4 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Trinity Biotech | 3.46
Trinity Biotech HPLC (Affinity) | 27 | 9.46 | -0.04 | 3.9 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Trinity Biotech | 3.86
Trinity Biotech HPLC (Affinity) | 27 | 7.36 | -0.04 | 3.4 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Trinity Biotech | 3.36
Trinity Biotech HPLC (Affinity) | 18 | 5.85 | -0.05 | 3.3 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Trinity Biotech | 3.25
Trinity Biotech HPLC (Affinity) | 18 | 9.69 | -0.11 | 1.6 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Trinity Biotech | 1.49
Trinity Biotech HPLC (Affinity) | 18 | 7.36 | -0.04 | 3.5 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Trinity Biotech | 3.46
Trinity Biotech HPLC (Affinity) | 20 | 5.33 | 0.13 | 2.5 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Trinity Biotech | 2.63
Trinity Biotech HPLC (Affinity) | 20 | 8.63 | -0.07 | 3.1 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Trinity Biotech | 3.03
Trinity Biotech HPLC (Affinity) | 20 | 6.38 | 0.08 | 2.4 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Trinity Biotech | 2.48
Trinity Biotech HPLC (Affinity) | 19 | 8.58 | 0.08 | 2.6 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Trinity Biotech | 2.68
Trinity Biotech HPLC (Affinity) | 19 | 5.57 | 0.17 | 3 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Trinity Biotech | 3.17
Trinity Biotech HPLC (Affinity) | 19 | 6.46 | 0.06 | 1.8 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Trinity Biotech | 1.86
Trinity Biotech HPLC (Affinity) | 20 | 6.41 | 0.11 | 3.8 | GH2-04 | 6.3 | 2011 | CAP | Fresh | Trinity Biotech | 3.91
Trinity Biotech HPLC (Affinity) | 20 | 7.68 | 0.08 | 2.3 | GH2-05 | 7.6 | 2011 | CAP | Fresh | Trinity Biotech | 2.38
Trinity Biotech HPLC (Affinity) | 20 | 9.29 | 0.09 | 2.1 | GH2-06 | 9.2 | 2011 | CAP | Fresh | Trinity Biotech | 2.19
Trinity Biotech HPLC (Affinity) | 31 | 5.76 | 0.16 | 3.1 | GH2-01 | 5.6 | 2012 | CAP | Fresh | Trinity Biotech | 3.26
Trinity Biotech HPLC (Affinity) | 31 | 9.25 | -0.15 | 2.5 | GH2-02 | 9.4 | 2012 | CAP | Fresh | Trinity Biotech | 2.35
Trinity Biotech HPLC (Affinity) | 31 | 7.31 | 0.11 | 2.3 | GH2-03 | 7.2 | 2012 | CAP | Fresh | Trinity Biotech | 2.41
Trinity Biotech HPLC (Affinity) | 30 | 5.53 | 0.13 | 1.8 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Trinity Biotech | 1.93
Trinity Biotech HPLC (Affinity) | 30 | 8.28 | -0.02 | 2.3 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Trinity Biotech | 2.28
Trinity Biotech HPLC (Affinity) | 30 | 5.69 | 0.04 | 1.8 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Trinity Biotech | 1.84
Trinity Biotech Premier | 50 | 6.56 | -0.02 | 2.7 | GH2-04 | 6.58 | 2014 | CAP | Fresh | Trinity Biotech | 2.68
Trinity Biotech Premier | 50 | 8.4 | 0.01 | 2.7 | GH2-05 | 8.39 | 2014 | CAP | Fresh | Trinity Biotech | 2.71
Trinity Biotech Premier | 50 | 5.69 | 0.04 | 2.6 | GH2-06 | 5.65 | 2014 | CAP | Fresh | Trinity Biotech | 2.64
Trinity Biotech Premier | 41 | 6.42 | -0.07 | 2.5 | GH2-01 | 6.49 | 2014 | CAP | Fresh | Trinity Biotech | 2.43
Trinity Biotech Premier | 41 | 6.83 | -0.14 | 2.4 | GH2-02 | 6.97 | 2014 | CAP | Fresh | Trinity Biotech | 2.26
Trinity Biotech Premier | 41 | 9.51 | -0.14 | 2.7 | GH2-03 | 9.65 | 2014 | CAP | Fresh | Trinity Biotech | 2.56
Trinity Biotech Premier | 60 | 6.74 | -0.05 | 2.3 | GH2-01 | 6.79 | 2015 | CAP | Fresh | Trinity Biotech | 2.25
Trinity Biotech Premier | 60 | 10.22 | -0.06 | 2.3 | GH2-02 | 10.28 | 2015 | CAP | Fresh | Trinity Biotech | 2.24
Trinity Biotech Premier | 60 | 6.83 | 0.01 | 2.3 | GH5-03 | 6.82 | 2015 | CAP | Fresh | Trinity Biotech | 2.31
Trinity Biotech Premier | 54 | 8.54 | -0.09 | 1.9 | GH5-04 | 8.63 | 2015 | CAP | Fresh | Trinity Biotech | 1.81
Trinity Biotech Premier | 54 | 5.35 | 0.03 | 2.2 | GH5-05 | 5.32 | 2015 | CAP | Fresh | Trinity Biotech | 2.23
Trinity Biotech Premier | 53 | 5.84 | 0 | 2.9 | GH5-06 | 5.84 | 2015 | CAP | Fresh | Trinity Biotech | 2.9
Trinity Biotech Premier | 53 | 11.73 | 0.02 | 3.3 | GH5-07 | 11.71 | 2015 | CAP | Fresh | Trinity Biotech | 3.32
Trinity Biotech Premier | 53 | 9.61 | 0.08 | 3 | GH5-08 | 9.53 | 2015 | CAP | Fresh | Trinity Biotech | 3.08
Trinity Biotech Premier | 53 | 5.13 | 0.09 | 2.7 | GH5-09 | 5.04 | 2015 | CAP | Fresh | Trinity Biotech | 2.79
Trinity Biotech Premier | 53 | 7.41 | 0.03 | 2.7 | GH5-10 | 7.38 | 2015 | CAP | Fresh | Trinity Biotech | 2.73
Trinity Biotech Premier | 55-65 | 12 | 0.31 | 1.5 | GH5-11 | 11.69 | 2015 | CAP | Fresh | Trinity Biotech | 1.81
Trinity Biotech Premier | 55-65 | 6.06 | 0.13 | 1.9 | GH5-12 | 5.93 | 2015 | CAP | Fresh | Trinity Biotech | 2.03
Trinity Biotech Premier | 55-65 | 5.3 | 0.13 | 1.7 | GH5-13 | 5.17 | 2015 | CAP | Fresh | Trinity Biotech | 1.83
Trinity Biotech Premier | 55-65 | 8.33 | 0.19 | 1.8 | GH5-14 | 8.14 | 2015 | CAP | Fresh | Trinity Biotech | 1.99
Trinity Biotech Premier | 55-65 | 9.67 | 0.37 | 1.7 | GH5-15 | 9.3 | 2015 | CAP | Fresh | Trinity Biotech | 2.07
Trinity Biotech Premier | 54-63 | 5.42 | 0.1 | 2.3 | GH5-01 | 5.32 | 2016 | CAP | Fresh | Trinity Biotech | 2.4
Trinity Biotech Premier | 54-63 | 9.39 | 0.23 | 2.2 | GH5-02 | 9.16 | 2016 | CAP | Fresh | Trinity Biotech | 2.43
Trinity Biotech Premier | 54-63 | 5.45 | 0.14 | 2.3 | GH5-03 | 5.31 | 2016 | CAP | Fresh | Trinity Biotech | 2.44
Trinity Biotech Premier | 54-63 | 12.36 | 0.33 | 2.1 | GH5-04 | 12.03 | 2016 | CAP | Fresh | Trinity Biotech | 2.43
Trinity Biotech Premier | 54-63 | 6.14 | 0.2 | 2.3 | GH5-05 | 5.94 | 2016 | CAP | Fresh | Trinity Biotech | 2.5
Trinity Biotech Premier | 52 | 5.30 | 0.03 | 2.6 | GH5-06 | 5.27 | 2016 | CAP | Fresh | Trinity Biotech | 2.63
Trinity Biotech Premier | 52 | 10.70 | 0.11 | 1.9 | GH5-07 | 10.59 | 2016 | CAP | Fresh | Trinity Biotech | 2.01
Trinity Biotech Premier | 52 | 6.29 | 0.09 | 2.1 | GH5-08 | 6.2 | 2016 | CAP | Fresh | Trinity Biotech | 2.19
Trinity Biotech Premier | 52 | 12.48 | 0.25 | 2 | GH5-09 | 12.23 | 2016 | CAP | Fresh | Trinity Biotech | 2.25
Trinity Biotech Premier | 52 | 7.63 | 0.12 | 2.2 | GH5-10 | 7.51 | 2016 | CAP | Fresh | Trinity Biotech | 2.32
Trinity Biotech Premier | 58-69 | 9.20 | 0.09 | 2.2 | GH5-11 | 9.11 | 2016 | CAP | Fresh | Trinity Biotech | 2.29
Trinity Biotech Premier | 58-69 | 6.12 | 0.11 | 2 | GH5-12 | 6.01 | 2016 | CAP | Fresh | Trinity Biotech | 2.11
Trinity Biotech Premier | 58-69 | 11.99 | 0.28 | 2.1 | GH5-13 | 11.71 | 2016 | CAP | Fresh | Trinity Biotech | 2.38
Trinity Biotech Premier | 58-69 | 5.07 | 0.05 | 2.3 | GH5-14 | 5.02 | 2016 | CAP | Fresh | Trinity Biotech | 2.35
Trinity Biotech Premier | 58-69 | 7.68 | 0.1 | 2.1 | GH5-15 | 7.58 | 2016 | CAP | Fresh | Trinity Biotech | 2.2
Trinity Biotech Premier | 67-75 | 6.36 | -0.05 | 1.9 | GH5-01 | 6.41 | 2017 | CAP | Fresh | Trinity Biotech | 1.85
Trinity Biotech Premier | 67-75 | 9.53 | 0 | 2 | GH5-02 | 9.53 | 2017 | CAP | Fresh | Trinity Biotech | 0
Trinity Biotech Premier | 67-75 | 5.36 | 0.02 | 1.9 | GH5-03 | 5.34 | 2017 | CAP | Fresh | Trinity Biotech | 1.92
Trinity Biotech Premier | 67-75 | 8.50 | -0.01 | 2 | GH5-04 | 8.51 | 2017 | CAP | Fresh | Trinity Biotech | 1.99
Trinity Biotech Premier | 67-75 | 7.39 | 0.14 | 2.1 | GH5-05 | 7.25 | 2017 | CAP | Fresh | Trinity Biotech | 2.24
Trinity Biotech Premier | 85 | 7.35 | 0.2 | 1.7 | GH5-01 | 7.15 | 2018 | CAP | Fresh | Trinity Biotech | 1.9
Trinity Biotech Premier | 85 | 5.41 | 0.22 | 2 | GH5-02 | 5.19 | 2018 | CAP | Fresh | Trinity Biotech | 2.22
Trinity Biotech Premier | 85 | 8.55 | 0.13 | 1.4 | GH5-03 | 8.42 | 2018 | CAP | Fresh | Trinity Biotech | 1.53
Trinity Biotech Premier | 85 | 9.96 | 0.17 | 1.5 | GH5-04 | 9.79 | 2018 | CAP | Fresh | Trinity Biotech | 1.67
Trinity Biotech Premier | 85 | 6.33 | 0.21 | 1.9 | GH5-05 | 6.12 | 2018 | CAP | Fresh | Trinity Biotech | 2.11
Trinity Biotech Premier | 70 | 6.43 | 0.12 | 1.6 | GH5-06 | 6.31 | 2018 | CAP | Fresh | Trinity Biotech | 1.72
Trinity Biotech Premier | 70 | 9.29 | 0.18 | 1.3 | GH5-07 | 9.11 | 2018 | CAP | Fresh | Trinity Biotech | 1.48
Trinity Biotech Premier | 70 | 5.56 | 0.25 | 1.8 | GH5-08 | 5.31 | 2018 | CAP | Fresh | Trinity Biotech | 2.05
Trinity Biotech Premier | 70 | 8.48 | 0.13 | 1.7 | GH5-09 | 8.35 | 2018 | CAP | Fresh | Trinity Biotech | 1.83
Trinity Biotech Premier | 70 | 7.81 | 0.19 | 1.5 | GH5-10 | 7.62 | 2018 | CAP | Fresh | Trinity Biotech | 1.69
Trinity Biotech Premier | 69 | 7.38 | -0.04 | 1.9 | GH5-06 | 7.42 | 2017 | CAP | Fresh | Trinity Biotech | 1.86
Trinity Biotech Premier | 69 | 5.21 | 0.01 | 2 | GH5-07 | 5.2 | 2017 | CAP | Fresh | Trinity Biotech | 2.01
Trinity Biotech Premier | 69 | 8.39 | 0.08 | 1.8 | GH5-08 | 8.31 | 2017 | CAP | Fresh | Trinity Biotech | 1.88
Trinity Biotech Premier | 69 | 10.51 | 0.18 | 1.9 | GH5-09 | 10.33 | 2017 | CAP | Fresh | Trinity Biotech | 2.08
Trinity Biotech Premier | 69 | 5.91 | 0.04 | 1.9 | GH5-10 | 5.87 | 2017 | CAP | Fresh | Trinity Biotech | 1.94
Trinity Biotech Premier | 79 | 7.28 | 0.03 | 2.3 | GH5-11 | 7.25 | 2017 | CAP | Fresh | Trinity Biotech | 2.33
Trinity Biotech Premier | 79 | 6.79 | -0.02 | 2.3 | GH5-12 | 6.81 | 2017 | CAP | Fresh | Trinity Biotech | 2.28
Trinity Biotech Premier | 79 | 8.92 | 0.27 | 2.3 | GH5-13 | 8.65 | 2017 | CAP | Fresh | Trinity Biotech | 2.57
Trinity Biotech Premier | 79 | 9.58 | 0.08 | 2.1 | GH5-14 | 9.50 | 2017 | CAP | Fresh | Trinity Biotech | 2.18
Trinity Biotech Premier | 79 | 5.55 | 0.1 | 2.4 | GH5-15 | 5.45 | 2017 | CAP | Fresh | Trinity Biotech | 2.5
Trinity Biotech Premier Hb9210 | 11 | 7.17 | 0.06 | 2.5 | GH2-01 | 7.11 | 2013 | CAP | Fresh | Trinity Biotech | 2.56
Trinity Biotech Premier Hb9210 | 11 | 9.26 | -0.06 | 2.4 | GH2-02 | 9.32 | 2013 | CAP | Fresh | Trinity Biotech | 2.34
Trinity Biotech Premier Hb9210 | 11 | 6.17 | 0.1 | 2.5 | GH2-03 | 6.07 | 2013 | CAP | Fresh | Trinity Biotech | 2.6
Trinity Biotech Premier Hb9210 | 29 | 7.96 | -0.09 | 2.5 | GH2-04 | 8.05 | 2013 | CAP | Fresh | Trinity Biotech | 2.41
Trinity Biotech Premier Hb9210 | 29 | 5.36 | 0.06 | 2.9 | GH2-05 | 5.3 | 2013 | CAP | Fresh | Trinity Biotech | 2.96
Trinity Biotech Premier Hb9210 | 29 | 6.18 | 0.04 | 2.5 | GH2-06 | 6.14 | 2013 | CAP | Fresh | Trinity Biotech | 2.54
Trinity Biotech Premier Hb9210 HPLC | 87 | 9.45 | 0.08 | 1.8 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Trinity Biotech | 1.88
Trinity Biotech Premier Hb9210 HPLC | 87 | 5.12 | 0.2 | 1.8 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Trinity Biotech | 2
Trinity Biotech Premier Hb9210 HPLC | 87 | 6.11 | 0.07 | 1.8 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Trinity Biotech | 1.87
Trinity Biotech Premier Hb9210 HPLC | 87 | 5.38 | 0.1 | 1.6 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Trinity Biotech | 1.7
Trinity Biotech Premier Hb9210 HPLC | 87 | 8.20 | 0.12 | 1.7 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Trinity Biotech | 1.82
Trinity Biotech Premier Hb9210 HPLC | 83 | 8.79 | -0.07 | 1.7 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Trinity Biotech | 1.63
Trinity Biotech Premier Hb9210 HPLC | 83 | 6.81 | -0.03 | 1.8 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Trinity Biotech | 1.77
Trinity Biotech Premier Hb9210 HPLC | 83 | 9.36 | -0.03 | 1.8 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Trinity Biotech | 1.77
Trinity Biotech Premier Hb9210 HPLC | 83 | 6.14 | 0.01 | 1.8 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Trinity Biotech | 1.81
Trinity Biotech Premier Hb9210 HPLC | 83 | 7.55 | 0.03 | 1.7 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Trinity Biotech | 1.73
Trinity Biotech Premier Hb9210 HPLC | 89 | 5.43 | -0.03 | 2.2 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Trinity Biotech | 2.17
Trinity Biotech Premier Hb9210 HPLC | 89 | 5.74 | 0.08 | 2.8 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Trinity Biotech | 2.88
Trinity Biotech Premier Hb9210 HPLC | 89 | 9.43 | 0.12 | 2 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Trinity Biotech | 2.12
Trinity Biotech Premier Hb9210 HPLC | 89 | 5.35 | 0.07 | 2.2 | GH5-04 | 5.28 | 2019 | CAP | Fresh | Trinity Biotech | 2.27
Trinity Biotech Premier Hb9210 HPLC | 89 | 7.45 | 0.04 | 2.2 | GH5-05 | 7.41 | 2019 | CAP | Fresh | Trinity Biotech | 2.24
Trinity Biotech Premier Hb9210 HPLC | 80 | 6.39 | -0.02 | 1.9 | GH5-06 | 6.41 | 2019 | CAP | Fresh | Trinity Biotech | 1.88
Trinity Biotech Premier Hb9210 HPLC | 80 | 8.73 | 0.13 | 1.7 | GH5-07 | 8.6 | 2019 | CAP | Fresh | Trinity Biotech | 1.83
Trinity Biotech Premier Hb9210 HPLC | 80 | 5.51 | 0.09 | 2.1 | GH5-08 | 5.42 | 2019 | CAP | Fresh | Trinity Biotech | 2.19
Trinity Biotech Premier Hb9210 HPLC | 80 | 7.50 | 0.12 | 1.8 | GH5-09 | 7.38 | 2019 | CAP | Fresh | Trinity Biotech | 1.92
Trinity Biotech Premier Hb9210 HPLC | 80 | 9.93 | 0.18 | 1.9 | GH5-10 | 9.75 | 2019 | CAP | Fresh | Trinity Biotech | 2.08
Trinity Biotech Premier Hb9210 HPLC | 86 | 5.37 | 0.12 | 2.1 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Trinity Biotech | 2.22
Trinity Biotech Premier Hb9210 HPLC | 86 | 6.46 | 0.05 | 1.8 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Trinity Biotech | 1.85
Trinity Biotech Premier Hb9210 HPLC | 86 | 8.41 | 0.2 | 1.6 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Trinity Biotech | 1.8
Trinity Biotech Premier Hb9210 HPLC | 86 | 5.61 | 0.14 | 2 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Trinity Biotech | 2.14
Trinity Biotech Premier Hb9210 HPLC | 86 | 9.73 | 0.25 | 1.7 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Trinity Biotech | 1.95
Trinity Biotech Premier Hb9210 HPLC | 80 | 8.11 | 0.14 | 2.1 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Trinity Biotech | 2.24
Trinity Biotech Premier Hb9210 HPLC | 80 | 6.06 | 0.17 | 2.3 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Trinity Biotech | 2.47
Trinity Biotech Premier Hb9210 HPLC | 80 | 5.28 | 0.15 | 2.2 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Trinity Biotech | 2.35
Trinity Biotech Premier Hb9210 HPLC | 80 | 7.53 | 0.13 | 2.5 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Trinity Biotech | 2.63
Trinity Biotech Premier Hb9210 HPLC | 80 | 9.31 | 0.14 | 2.3 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Trinity Biotech | 2.44
Trinity Biotech Premier Hb9210 HPLC | 71 | 7.94 | 0.13 | 2.2 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Trinity Biotech | 2.33
Trinity Biotech Premier Hb9210 HPLC | 71 | 6.03 | 0.12 | 2.2 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Trinity Biotech | 2.32
Trinity Biotech Premier Hb9210 HPLC | 71 | 8.84 | 0.15 | 2.2 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Trinity Biotech | 2.35
Trinity Biotech Premier Hb9210 HPLC | 71 | 5.97 | 0.11 | 2.5 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Trinity Biotech | 2.61
Trinity Biotech Premier Hb9210 HPLC | 71 | 9.81 | 0.25 | 2.1 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Trinity Biotech | 2.35
Trinity Biotech Premier Hb9210 HPLC | 79 | 5.47 | -0.01 | 1.8 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Trinity Biotech | 1.79
Trinity Biotech Premier Hb9210 HPLC | 79 | 8.2 | -0.06 | 1.8 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Trinity Biotech | 1.74
Trinity Biotech Premier Hb9210 HPLC | 79 | 5.2 | 0.05 | 2 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Trinity Biotech | 0.05
Trinity Biotech Premier Hb9210 HPLC | 79 | 9.94 | -0.11 | 1.7 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Trinity Biotech | 1.59
Trinity Biotech Premier Hb9210 HPLC | 79 | 5.95 | -0.04 | 2 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Trinity Biotech | -0.04
Trinity Biotech Premier Hb9210 HPLC | 73 | 5.78 | -0.02 | 2 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Trinity Biotech | -0.02
Trinity Biotech Premier Hb9210 HPLC | 73 | 9.26 | 0.14 | 2 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Trinity Biotech | 0.14
Trinity Biotech Premier Hb9210 HPLC | 73 | 5.54 | 0.07 | 1.9 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Trinity Biotech | 1.97
Trinity Biotech Premier Hb9210 HPLC | 73 | 7.38 | 0.13 | 2 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Trinity Biotech | 0.13
Trinity Biotech Premier Hb9210 HPLC | 73 | 9.37 | 0.11 | 2 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Trinity Biotech | 0.11
Trinity Biotech Premier Hb9210 HPLC | 81 | 5.31 | 0.1 | 2.2 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Trinity Biotech | 2.3
Trinity Biotech Premier Hb9210 HPLC | 81 | 6.74 | 0.06 | 2 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Trinity Biotech | 0.06
Trinity Biotech Premier Hb9210 HPLC | 81 | 8.5 |  | 2.4 | GH5-13 |  | 2021 | CAP | Fresh | Trinity Biotech | 2.4
Trinity Biotech Premier Hb9210 HPLC | 81 | 6.04 | 0.06 | 2.1 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Trinity Biotech | 2.16
Trinity Biotech Premier Hb9210 HPLC | 81 | 10.3 | 0.18 | 1.9 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Trinity Biotech | 2.08
Vitros 5.1 FS Chem System | 26 | 11 |  | 4.5 | GH2-01 | 10.7 | 2006 | CAP | Fresh | Vitros | 4.5
Vitros 5.1 FS Chem System | 26 | 5.7 |  | 3.7 | GH2-02 | 5.30 | 2006 | CAP | Fresh | Vitros | 3.7
Vitros 5.1 FS Chem System | 26 | 8.2 |  | 3.7 | GH2-03 | 8.40 | 2006 | CAP | Fresh | Vitros | 3.7
Vitros 5.1 FS Chem System | 39 | 6.7 |  | 2.7 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Vitros | 2.7
Vitros 5.1 FS Chem System | 39 | 6.7 |  | 3.9 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Vitros | 3.9
Vitros 5.1 FS Chem System | 39 | 11.4 |  | 4.5 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Vitros | 4.5
Vitros 5.1 FS Chem System | 39 | 6.7 |  | 2.7 | GH2-04 | 7.1 | 2006 | CAP | Fresh | Vitros | 2.7
Vitros 5.1 FS Chem System | 39 | 6.7 |  | 3.9 | GH2-05 | 7.0 | 2006 | CAP | Fresh | Vitros | 3.9
Vitros 5.1 FS Chem System | 39 | 11.4 |  | 4.5 | GH2-06 | 11.7 | 2006 | CAP | Fresh | Vitros | 4.5
Vitros 5.1 FS Chem System | 66 | 5.7 |  | 3.6 | GH2-01 | 5.4 | 2007 | CAP | Fresh | Vitros | 3.6
Vitros 5.1 FS Chem System | 66 | 12.3 |  | 4.4 | GH2-02 | 11.6 | 2007 | CAP | Fresh | Vitros | 4.4
Vitros 5.1 FS Chem System | 66 | 7.7 |  | 4 | GH2-03 | 7.6 | 2007 | CAP | Fresh | Vitros | 4
Vitros 5.1 FS Chem System | 83 | 8.9 |  | 4.9 | GH2-04 | 9.2 | 2007 | CAP | Fresh | Vitros | 4.9
Vitros 5.1 FS Chem System | 83 | 11.2 |  | 5.3 | GH2-05 | 11.1 | 2007 | CAP | Fresh | Vitros | 5.3
Vitros 5.1 FS Chem System | 83 | 6.1 |  | 3.8 | GH2-06 | 6.2 | 2007 | CAP | Fresh | Vitros | 3.8
Vitros 5.1 FS Chem System | 98 | 12.9 |  | 5.9 | GH2-01 | 13 | 2008 | CAP | Fresh | Vitros | 5.9
Vitros 5.1 FS Chem System | 98 | 5.3 |  | 2.9 | GH2-02 | 5.3 | 2008 | CAP | Fresh | Vitros | 2.9
Vitros 5.1 FS Chem System | 98 | 8.6 |  | 3.7 | GH2-03 | 8.8 | 2008 | CAP | Fresh | Vitros | 3.7
Vitros 5.1 FS Chem System | 117 | 11.5 |  | 4.5 | GH2-04 | 11.9 | 2008 | CAP | Fresh | Vitros | 4.5
Vitros 5.1 FS Chem System | 117 | 6.6 |  | 2.5 | GH2-05 | 6.8 | 2008 | CAP | Fresh | Vitros | 2.5
Vitros 5.1 FS Chem System | 117 | 5.2 |  | 3.6 | GH2-06 | 5 | 2008 | CAP | Fresh | Vitros | 3.6
Vitros 5.1 FS Chem System | 111 | 5.1 |  | 3.4 | GH2-01 | 5.1 | 2009 | CAP | Fresh | Vitros | 3.4
Vitros 5.1 FS Chem System | 111 | 8.3 |  | 3.8 | GH2-02 | 8.4 | 2009 | CAP | Fresh | Vitros | 3.8
Vitros 5.1 FS Chem System | 111 | 5.9 |  | 3.9 | GH2-03 | 6 | 2009 | CAP | Fresh | Vitros | 3.9
Vitros 5.1 FS Chem System | 124 | 6.8 | 0.2 | 4.1 | GH2-04 | 6.6 | 2009 | CAP | Fresh | Vitros | 4.3
Vitros 5.1 FS Chem System | 124 | 9.58 | 0.08 | 3.7 | GH2-05 | 9.5 | 2009 | CAP | Fresh | Vitros | 3.78
Vitros 5.1 FS Chem System | 124 | 7.55 | 0.15 | 4.2 | GH2-06 | 7.4 | 2009 | CAP | Fresh | Vitros | 4.35
Vitros 5.1 FS Chem System | 212 | 5.28 | -0.12 | 2 | GH2-04 | 5.4 | 2012 | CAP | Fresh | Vitros | 1.88
Vitros 5.1 FS Chem System | 212 | 8.1 | -0.2 | 2.4 | GH2-05 | 8.3 | 2012 | CAP | Fresh | Vitros | 2.2
Vitros 5.1 FS Chem System | 212 | 5.41 | -0.24 | 2.6 | GH2-06 | 5.65 | 2012 | CAP | Fresh | Vitros | 2.36
Vitros 5.1 FS Chem System | 132 | 6.27 | 0.37 | 3.4 | GH2-01 | 5.9 | 2010 | CAP | Fresh | Vitros | 3.77
Vitros 5.1 FS Chem System | 132 | 10.2 | 0.4 | 3.7 | GH2-02 | 9.8 | 2010 | CAP | Fresh | Vitros | 4.1
Vitros 5.1 FS Chem System | 132 | 7.85 | 0.45 | 3.7 | GH2-03 | 7.4 | 2010 | CAP | Fresh | Vitros | 4.15
Vitros 5.1 FS Chem System | 154 | 5.52 | 0.32 | 4.6 | GH2-04 | 5.2 | 2010 | CAP | Fresh | Vitros | 4.92
Vitros 5.1 FS Chem System | 154 | 8.92 | 0.22 | 3.4 | GH2-05 | 8.7 | 2010 | CAP | Fresh | Vitros | 3.62
Vitros 5.1 FS Chem System | 154 | 6.57 | 0.27 | 3.7 | GH2-06 | 6.3 | 2010 | CAP | Fresh | Vitros | 3.97
Vitros 5.1 FS Chem System | 155 | 8.71 | 0.21 | 3.5 | GH2-01 | 8.5 | 2011 | CAP | Fresh | Vitros | 3.71
Vitros 5.1 FS Chem System | 155 | 5.48 | 0.08 | 3.6 | GH2-02 | 5.4 | 2011 | CAP | Fresh | Vitros | 3.68
Vitros 5.1 FS Chem System | 155 | 6.56 | 0.16 | 3.3 | GH2-03 | 6.4 | 2011 | CAP | Fresh | Vitros | 3.46
Vitros 5.1 FS/4600/5600 Chemistry Systems | 188 | 9.38 | 0.01 | 2.9 | GH5-11 | 9.37 | 2018 | CAP | Fresh | Vitros | 2.91
Vitros 5.1 FS/4600/5600 Chemistry Systems | 188 | 4.88 | -0.04 | 1.7 | GH5-12 | 4.92 | 2018 | CAP | Fresh | Vitros | 1.66
Vitros 5.1 FS/4600/5600 Chemistry Systems | 188 | 5.90 | -0.14 | 2.4 | GH5-13 | 6.04 | 2018 | CAP | Fresh | Vitros | 2.26
Vitros 5.1 FS/4600/5600 Chemistry Systems | 188 | 5.19 | -0.09 | 2.1 | GH5-14 | 5.28 | 2018 | CAP | Fresh | Vitros | 2.01
Vitros 5.1 FS/4600/5600 Chemistry Systems | 188 | 8.04 | -0.04 | 2.8 | GH5-15 | 8.08 | 2018 | CAP | Fresh | Vitros | 2.76
Vitros 5.1 FS/4600/5600 Chemistry Systems | 193 | 5.41 | -0.05 | 2.5 | GH5-01 | 5.46 | 2019 | CAP | Fresh | Vitros | 2.45
Vitros 5.1 FS/4600/5600 Chemistry Systems | 193 | 5.59 | -0.07 | 2.7 | GH5-02 (HbAS) | 5.66 | 2019 | CAP | Fresh | Vitros | 2.63
Vitros 5.1 FS/4600/5600 Chemistry Systems | 193 | 9.36 | 0.05 | 2.7 | GH5-03 | 9.31 | 2019 | CAP | Fresh | Vitros | 2.75
Vitros 5.1 FS/4600/5600 Chemistry Systems | 193 | 5.17 | -0.11 | 2.4 | GH5-04 | 5.28 | 2019 | CAP | Fresh | Vitros | 2.29
Vitros 5.1 FS/4600/5600 Chemistry Systems | 193 | 7.31 | -0.1 | 2.9 | GH5-05 | 7.41 | 2019 | CAP | Fresh | Vitros | 2.8
Vitros 5.1 FS/4600/5600 Chemistry Systems | 163 | 6.32 | -0.09 | 2.7 | GH5-06 | 6.41 | 2019 | CAP | Fresh | Vitros | 2.61
Vitros 5.1 FS/4600/5600 Chemistry Systems | 163 | 8.71 | 0.11 | 2.7 | GH5-07 | 8.6 | 2019 | CAP | Fresh | Vitros | 2.81
Vitros 5.1 FS/4600/5600 Chemistry Systems | 163 | 5.31 | -0.11 | 2.6 | GH5-08 | 5.42 | 2019 | CAP | Fresh | Vitros | 2.49
Vitros 5.1 FS/4600/5600 Chemistry Systems | 163 | 7.34 | -0.04 | 2.7 | GH5-09 | 7.38 | 2019 | CAP | Fresh | Vitros | 2.66
Vitros 5.1 FS/4600/5600 Chemistry Systems | 163 | 9.94 | 0.19 | 2.9 | GH5-10 | 9.75 | 2019 | CAP | Fresh | Vitros | 3.09
Vitros 5.1 FS/4600/5600 Chemistry Systems | 194 | 5.30 | 0.05 | 3 | GH5-11 | 5.25 | 2019 | CAP | Fresh | Vitros | 3.05
Vitros 5.1 FS/4600/5600 Chemistry Systems | 194 | 6.54 | 0.13 | 3.4 | GH5-12 | 6.41 | 2019 | CAP | Fresh | Vitros | 3.53
Vitros 5.1 FS/4600/5600 Chemistry Systems | 194 | 8.25 | 0.04 | 2.8 | GH5-13 | 8.21 | 2019 | CAP | Fresh | Vitros | 2.84
Vitros 5.1 FS/4600/5600 Chemistry Systems | 194 | 5.52 | 0.05 | 3.3 | GH5-14 | 5.47 | 2019 | CAP | Fresh | Vitros | 3.35
Vitros 5.1 FS/4600/5600 Chemistry Systems | 194 | 9.46 | -0.02 | 3 | GH5-15 | 9.48 | 2019 | CAP | Fresh | Vitros | 2.98
Vitros 5.1 FS/4600/5600 Chemistry Systems | 184 | 8.1 | 0.13 | 2.9 | GH5-01 | 7.97 | 2020 | CAP | Fresh | Vitros | 3.03
Vitros 5.1 FS/4600/5600 Chemistry Systems | 184 | 5.94 | 0.05 | 2.5 | GH5-02 | 5.89 | 2020 | CAP | Fresh | Vitros | 2.55
Vitros 5.1 FS/4600/5600 Chemistry Systems | 184 | 5.07 | -0.06 | 2.1 | GH5-03 | 5.13 | 2020 | CAP | Fresh | Vitros | 2.04
Vitros 5.1 FS/4600/5600 Chemistry Systems | 184 | 7.48 | 0.08 | 3.1 | GH5-04 | 7.4 | 2020 | CAP | Fresh | Vitros | 3.18
Vitros 5.1 FS/4600/5600 Chemistry Systems | 184 | 9.02 | -0.15 | 3.3 | GH5-05 | 9.17 | 2020 | CAP | Fresh | Vitros | 3.15
Vitros 5.1 FS/4600/5600 Chemistry Systems | 180 | 7.74 | -0.07 | 2.8 | GH5-06 | 7.81 | 2020 | CAP | Fresh | Vitros | 2.73
Vitros 5.1 FS/4600/5600 Chemistry Systems | 180 | 5.88 | -0.03 | 2.4 | GH5-07 | 5.91 | 2020 | CAP | Fresh | Vitros | 2.37
Vitros 5.1 FS/4600/5600 Chemistry Systems | 180 | 8.45 | -0.24 | 3.1 | GH5-08 | 8.69 | 2020 | CAP | Fresh | Vitros | 2.86
Vitros 5.1 FS/4600/5600 Chemistry Systems | 180 | 5.82 | -0.04 | 2.4 | GH5-09 | 5.86 | 2020 | CAP | Fresh | Vitros | 2.36
Vitros 5.1 FS/4600/5600 Chemistry Systems | 180 | 9.16 | -0.4 | 2.9 | GH5-10 | 9.56 | 2020 | CAP | Fresh | Vitros | 2.5
Vitros 5.1 FS/4600/5600/XT 7600 Chem. Systems | 200 | 5.28 | -0.2 | 2.3 | GH5-11 | 5.48 | 2020 | CAP | Fresh | Vitros | 2.1
Vitros 5.1 FS/4600/5600/XT 7600 Chem. Systems | 200 | 8.2 | -0.06 | 3 | GH5-12 | 8.26 | 2020 | CAP | Fresh | Vitros | -0.06
Vitros 5.1 FS/4600/5600/XT 7600 Chem. Systems | 200 | 4.94 | -0.21 | 2.2 | GH5-13 | 5.15 | 2020 | CAP | Fresh | Vitros | 1.99
Vitros 5.1 FS/4600/5600/XT 7600 Chem. Systems | 200 | 9.77 | -0.28 | 2.9 | GH5-14 | 10.05 | 2020 | CAP | Fresh | Vitros | 2.62
Vitros 5.1 FS/4600/5600/XT 7600 Chem. Systems | 200 | 5.81 | -0.18 | 2.5 | GH5-15 | 5.99 | 2020 | CAP | Fresh | Vitros | 2.32
Vitros 5.1 FS/4600/5600/XT 7600 Chem. Systems | 183 | 5.75 | -0.05 | 2.5 | GH5-06 | 5.8 | 2021 | CAP | Fresh | Vitros | 2.45
Vitros 5.1 FS/4600/5600/XT 7600 Chem. Systems | 183 | 9.04 | -0.08 | 2.8 | GH5-07 | 9.12 | 2021 | CAP | Fresh | Vitros | 2.72
Vitros 5.1 FS/4600/5600/XT 7600 Chem. Systems | 183 | 5.41 | -0.06 | 2.3 | GH5-08 | 5.47 | 2021 | CAP | Fresh | Vitros | 2.24
Vitros 5.1 FS/4600/5600/XT 7600 Chem. Systems | 183 | 7.33 | 0.08 | 2.6 | GH5-09 | 7.25 | 2021 | CAP | Fresh | Vitros | 2.68
Vitros 5.1 FS/4600/5600/XT 7600 Chem. Systems | 183 | 9.26 | 0 | 2.9 | GH5-10 | 9.26 | 2021 | CAP | Fresh | Vitros | 2.9
Vitros 5.1 FS/4600/5600/XT 7600 Chem. Systems | 199 | 5.12 | -0.09 | 1.9 | GH5-11 | 5.21 | 2021 | CAP | Fresh | Vitros | 1.81
Vitros 5.1 FS/4600/5600/XT 7600 Chem. Systems | 199 | 6.67 | ‐0.01 | 2.4 | GH5-12 | 6.68 | 2021 | CAP | Fresh | Vitros | 2.4
Vitros 5.1 FS/4600/5600/XT 7600 Chem. Systems | 199 | 8.22 |  | 2.6 | GH5-13 |  | 2021 | CAP | Fresh | Vitros | 2.6
Vitros 5.1 FS/4600/5600/XT 7600 Chem. Systems | 199 | 5.94 | ‐0.04 | 2.1 | GH5-14 | 5.98 | 2021 | CAP | Fresh | Vitros | 2.1
Vitros 5.1 FS/4600/5600/XT 7600 Chem. Systems | 199 | 10.15 | 0.03 | 3.2 | GH5-15 | 10.12 | 2021 | CAP | Fresh | Vitros | 3.23
Vitros 5.1 FS/4600/5600/XT7600 Chem. Systems | 204 | 9.08 | 0.22 | 3.1 | GH5-01 | 8.86 | 2021 | CAP | Fresh | Vitros | 3.32
Vitros 5.1 FS/4600/5600/XT7600 Chem. Systems | 204 | 6.86 | 0.02 | 2.6 | GH5-02 | 6.84 | 2021 | CAP | Fresh | Vitros | 2.62
Vitros 5.1 FS/4600/5600/XT7600 Chem. Systems | 204 | 9.4 | 0.01 | 3.4 | GH5-03 | 9.39 | 2021 | CAP | Fresh | Vitros | 3.41
Vitros 5.1 FS/4600/5600/XT7600 Chem. Systems | 204 | 6.14 | 0.01 | 2.5 | GH5-04 | 6.13 | 2021 | CAP | Fresh | Vitros | 2.51
Vitros 5.1 FS/4600/5600/XT7600 Chem. Systems | 204 | 7.59 | 0.07 | 2.8 | GH5-05 | 7.52 | 2021 | CAP | Fresh | Vitros | 2.87